| 3GPP TR 38.870 V18.2.0 (2024-06) |  |
| --- | --- |
| Technical Report |  |
| 3rd Generation Partnership Project; Technical Specification Group Radio Access Network; Enhanced Over-the-Air (OTA) test methods for NR FR1 Total Radiated Power (TRP) and Total Radiated Sensitivity (TRS) (Release 18) |  |
| The present document has been developed within the 3rd Generation Partnership Project (3GPP TM) and may be further elaborated for the purposes of 3GPP. The present document has not been subject to any approval process by the 3GPP Organizational Partners and shall not be implemented. This Specification is provided for future development work within 3GPP only. The Organizational Partners accept no liability for any use of this Specification. Specifications and Reports for implementation of the 3GPP TM system should be obtained via the 3GPP Organizational Partners' Publications Offices. |  |

3GPP Postal address 3GPP support office address 650 Route des Lucioles - Sophia Antipolis Valbonne - FRANCE Tel.: +33 4 92 94 42 00 Fax: +33 4 93 65 47 16 Internet https://www.3gpp.org

Copyright Notification No part may be reproduced except as authorized by written permission. The copyright and the foregoing restriction extend to reproduction in all media. © 2024, 3GPP Organizational Partners (ARIB, ATIS, CCSA, ETSI, TSDSI, TTA, TTC). All rights reserved. UMTS™ is a Trade Mark of ETSI registered for the benefit of its members 3GPP™ is a Trade Mark of ETSI registered for the benefit of its Members and of the 3GPP Organizational Partners LTE™ is a Trade Mark of ETSI registered for the benefit of its Members and of the 3GPP Organizational Partners GSM® and the GSM logo are registered and owned by the GSM Association

Contents

Foreword	10

1	Scope	12

2	References	12

3	Definitions of terms, symbols and abbreviations	13

3.1	Terms	13

3.2	Symbols	14

3.3	Abbreviations	14

4	General	14

4.1	Device types	14

4.2	Testing configuration	15

4.2.1	UE use scenarios for TRP TRS test	15

4.2.2	UE mechanical mode description	15

4.3	Testing bands	15

4.3.1	General	15

4.3.2	Operating bands	15

4.3.3	Test parameters for each band	16

4.3.4	EN-DC band combinations	22

4.3.5	CA band combinations	24

4.3.6	Alternative parameters related to channel bandwidth and bands	25

4.3.6.1	General	25

4.3.6.2	Test parameters	25

4.3.6.3	TRS requirements for additional channel bandwidths	26

5	Performance metrics	26

5.1	Definition of the Total Radiated Power (TRP)	26

5.1.1	Definition of the Total Radiated Power (TRP) for AC	26

5.1.1.1	Applicability of TRP metrics	26

5.1.1.2	TRP for UEs supporting single-port transmission	26

5.1.1.3	TRP for UEs supporting non-coherent UL MIMO	27

5.1.1.4	TRP for UEs supporting coherent UL MIMO	28

5.1.2	Definition of the Total Radiated Power (TRP) for RC method	28

5.2	Definition of Total Radiated Sensitivity (TRS)	29

5.2.1	Definition of the Total Radiated Sensitivity (TRS) for AC	29

5.2.2	Definition of the Total Radiated Sensitivity (TRS) for RC method	30

6	UE positioning guidelines	30

6.1	Free space	30

6.2	Hand phantom only (Browsing mode)	31

6.2.1	Wide Grip Hand	31

6.2.2	PDA Grip Hand	32

6.3	Head and Hand phantom (Talk Mode)	33

6.3.1	General	33

6.3.2	Wide Grip Hand and Head	34

6.3.3	PDA Grip Hand and Head	34

6.4	Head phantom only	34

6.5	Forearm phantom	34

6.5.1	Forearm Phantom placement in the chamber	34

6.5.2	Wrist-Worn RedCap Device mounted on the Forearm Phantom	35

7	Anechoic Chamber method (Reference method)	36

7.1	General	36

7.2	Test setup	36

7.3	Calibration procedure	37

7.4	TRP Test procedure	38

7.4.1	General	38

7.4.2	TRP test procedure for NR 1Tx configuration	38

7.4.2.1	UE configuration	38

7.4.2.2	Test procedure	38

7.4.3	TRP test procedure for NR 2Tx configuration	39

7.4.3.1	UE configuration	39

7.4.3.2	TxD TRP Test procedure	39

7.4.3.3	Single-layer UL-MIMO TRP Test procedure	39

7.4.4	TRP test procedure for NR DL CA configuration	40

7.4.4.1	UE configuration	40

7.4.4.2	Test procedure	40

7.5	TRS Test procedure	40

7.5.1	General	40

7.5.2	TRS test procedure for NR 1Tx configuration	41

7.5.2.1	UE configuration	41

7.5.2.2	Test Procedure	41

7.5.3	TRS test procedure for NR 2Tx configuration	41

7.5.3.1	UE configuration	41

7.5.3.2	Test procedure	42

7.5.4	TRS test procedure for NR DL CA configuration	42

7.5.4.1	UE configuration	42

7.5.4.2	Test procedure	42

7.6	Ripple Test for Quiet Zone	42

7.6.1	General	42

7.6.2	Ripple test procedure	43

7.7	Minimum Range Length	46

8	Reverberation Chamber test methodology	48

8.1	General	48

8.2	Test setup	48

8.3	Chamber Characterization	49

8.3.1	S-parameters and power transfer functions	49

8.3.2	Chamber loading for coherence bandwidth	49

8.3.2.1	Coherence bandwidth calculation	49

8.3.3	Chamber spatial uniformity	50

8.4	Calibration procedure	52

8.4.1	S-parameters measurement	52

8.4.2	Calculation of the chamber reference transfer function	53

8.5	TRP Test procedure	53

8.5.1	TRP for SA and EN-DC	53

8.5.1.1	Test conditions	53

8.5.1.2	UE configurations	53

8.5.1.3	Test procedure	54

8.6	TRS Test procedure	54

8.6.1	TRS for SA and EN-DC	54

8.6.1.1	Test conditions	54

8.6.1.2	UE configurations	54

8.6.1.3	Test procedure	54

8.7	Test Volume	55

9	Testing time reduction methodologies	55

9.1	General	55

9.2	Measurement grids for Anechoic Chamber method	55

9.3	Other solutions	67

Annex A: UE coordinate system	67

A.1	Reference coordinate system	67

Annex B: Estimation of Measurement uncertainty	69

B.1	General	69

B.2	MU contribution descriptions for Anechoic Chamber method	70

B.2.1	Mismatch uncertainty	70

B.2.1.1	Mismatch uncertainty between measurement receiver / communication tester and the measurement antenna	70

B.2.1.1.1	Mismatch uncertainty through the connector between two elements	70

B.2.1.1.2	Mismatch uncertainty due to the interaction of several elements	71

B.2.1.1.3	Total combined mismatch uncertainty	72

B.2.1.2	Mismatch uncertainty of the RF relay	72

B.2.1.2.1	First part: RF Relay switched on the co-polarized signal	73

B.2.1.2.1.1	The mismatch through the connector between two elements	73

B.2.1.2.1.2	Mismatch due to the interaction between two elements or more	73

B.2.1.2.2	Second part: RF relay switched on the cross-polarized signal	74

B.2.1.2.2.1	The mismatch through the connector between two elements	74

B.2.1.2.2.2	Mismatch due to the interaction between two elements or more	74

B.2.1.2.3	Total combined mismatch uncertainty	75

B.2.2	Insertion loss	75

B.2.2.1	Insertion loss of the measurement antenna cable	75

B.2.2.2	Insertion loss of the measurement antenna attenuator (if used)	75

B.2.2.3	Insertion loss of the RF relays (if used)	75

B.2.2.4	Insertion loss: calibration antenna feed cable	75

B.2.2.5	Insertion loss: calibration antenna attenuator (if used)	75

B.2.3	Influence of the antenna cable	76

B.2.3.1	Measurement antenna cable	76

B.2.3.2	Calibration antenna cable	76

B.2.4	Measurement receiver: uncertainty of the absolute level	76

B.2.5	Communication tester: uncertainty of the absolute level	76

B.2.6	Sensitivity measurement: output level step resolution	76

B.2.7	Measurement distance	76

B.2.7.1	Offset of phase centre from axis(es) of rotation	77

B.2.7.1.1	Offset of DUT phase centre from axis(es) of rotation	77

B.2.7.1.2	Offset of calibration antenna phase centre from axis(es) of rotation	77

B.2.7.2	Mutual coupling	77

B.2.7.3	Phase curvature	77

B.2.8	Quality of quiet zone	78

B.2.9	DUT Tx-power drift	78

B.2.10	DUT sensitivity drift	79

B.2.11	Uncertainty related to the use of phantoms	79

B.2.11.1	Uncertainty from using different types of SAM phantom	79

B.2.11.2	Simulated tissue liquid uncertainty	79

B.2.11.3	Uncertainty of dielectric properties and shape of the hand phantom	79

B.2.11.4	Uncertainty from using different types of Laptop Ground Plane phantom	81

B.2.11.5	Positioning Uncertainty from using Phantoms	81

B.2.11.6	Uncertainty of dielectric properties and shape of the forearm phantom	81

B.2.12	Coarse sampling grid	82

B.2.13	Random uncertainty	82

B.2.14	Frequency response	83

B.2.15	Uncertainty of network analyser	83

B.2.16	Uncertainty of the gain/efficiency of the calibration antenna	83

B.3	MU contribution descriptions for Alternative method	83

B.3.1	Additional Power Loss in EUT Chassis	83

B.3.3	Quality of Spatial Uniformity	84

B.3.4	Sensitivity measurement: output level step resolution	84

B.4	MU Assessment for TRP	84

B.4.1	MU Assessment for TRP in Anechoic Chamber	84

B.4.2	MU Assessment for TRP in Reverberation Chamber	88

B.5	MU Assessment for TRS	90

B.5.1	MU Assessment for TRS in Anechoic Chamber	90

B.5.2	MU Assessment for TRS in Reverberation Chamber	95

Annex C: Environmental requirements	99

C.1	Scope	99

C.2	Ambient temperature	99

C.3	Operating voltage	99

Annex D: Phantom Definition	100

D.1	Head Phantom	100

D.2	Hand Phantom	100

D.2.1	PDA Grip Hand	100

D.2.2	Wide Grip Hand	100

D.3	Forearm Phantom	101

Annex E: Harmonization outcome of Alternative method and Reference method	102

E.1	General	102

E.2	Rel-18 Harmonization activity	102

E.2.1	Framework and workplan	102

E.2.2	Measurement results	103

E.2.3	Pass/fail limit	107

E.2.4	Conclusions	108

Annex F: Lab alignment outcome of Reference method	109

F.1	General	109

F.2	Rel-17 lab alignment campaign (Browsing mode)	109

F.2.1	Framework and workplan	109

F.2.2	Measurement results	111

F.2.3	Pass/fail limit	112

F.2.4	Conclusions	112

F.3	Rel-18 lab alignment (Talk mode)	112

F.3.1	Framework and workplan	112

F.3.2	Measurement results	114

F.3.3	Pass/fail limit	114

F.3.4	Conclusions	115

Annex G (informative): Change history	116

# Foreword

This Technical Report has been produced by the 3rd Generation Partnership Project (3GPP).

The contents of the present document are subject to continuing work within the TSG and may change following formal TSG approval. Should the TSG modify the contents of the present document, it will be re-released by the TSG with an identifying change of release date and an increase in version number as follows:

Version x.y.z

where:

x	the first digit:

1	presented to TSG for information;

2	presented to TSG for approval;

3	or greater indicates TSG approved document under change control.

y	the second digit is incremented for all changes of substance, i.e. technical enhancements, corrections, updates, etc.

z	the third digit is incremented when editorial only changes have been incorporated in the document.

In the present document, modal verbs have the following meanings:

shall		indicates a mandatory requirement to do something

shall not	indicates an interdiction (prohibition) to do something

The constructions "shall" and "shall not" are confined to the context of normative provisions, and do not appear in Technical Reports.

The constructions "must" and "must not" are not used as substitutes for "shall" and "shall not". Their use is avoided insofar as possible, and they are not used in a normative context except in a direct citation from an external, referenced, non-3GPP document, or so as to maintain continuity of style when extending or modifying the provisions of such a referenced document.

should		indicates a recommendation to do something

should not	indicates a recommendation not to do something

may		indicates permission to do something

need not	indicates permission not to do something

The construction "may not" is ambiguous and is not used in normative elements. The unambiguous constructions "might not" or "shall not" are used instead, depending upon the meaning intended.

can		indicates that something is possible

cannot		indicates that something is impossible

The constructions "can" and "cannot" are not substitutes for "may" and "need not".

will		indicates that something is certain or expected to happen as a result of action taken by an agency the behaviour of which is outside the scope of the present document

will not		indicates that something is certain or expected not to happen as a result of action taken by an agency the behaviour of which is outside the scope of the present document

might	indicates a likelihood that something will happen as a result of action taken by some agency the behaviour of which is outside the scope of the present document

might not	indicates a likelihood that something will not happen as a result of action taken by some agency the behaviour of which is outside the scope of the present document

In addition:

is	(or any other verb in the indicative mood) indicates a statement of fact

is not	(or any other negative verb in the indicative mood) indicates a statement of fact

The constructions "is" and "is not" do not indicate requirements.

# 1	Scope

The present document is a technical report for enhanced Over-the-Air (OTA) test methods for NR FR1 Total Radiated Power (TRP) and Total Radiated Sensitivity (TRS), for NR standalone (SA) and NR non-standalone (NSA) operation mode.

In Rel-17, the TRP TRS OTA test method was just developed for a limited scope. This TR targets to define a full set of OTA test methods to cover different UE features for TRP and TRS measurement.

# 2	References

The following documents contain provisions which, through reference in this text, constitute provisions of the present document.

-	References are either specific (identified by date of publication, edition number, version number, etc.) or non-specific.

-	For a specific reference, subsequent revisions do not apply.

-	For a non-specific reference, the latest version applies. In the case of a reference to a 3GPP document (including a GSM document), a non-specific reference implicitly refers to the latest version of that document in the same Release as the present document.

[1]	3GPP TR 21.905: "Vocabulary for 3GPP Specifications".

[2]	3GPP TR 38.827: “Study on radiated metrics and test methodology for the verification of multi-antenna reception performance of NR User Equipment (UE)”.

[3]	3GPP TS 38.101-1: “NR; User Equipment (UE) radio transmission and reception; Part 1: Range 1 Standalone”.

[4]	3GPP TS 38.101-3: “NR; User Equipment (UE) radio transmission and reception; Part 3: Range 1 and Range 2 Interworking operation with other radios”.

[5]	3GPP TS 38.521-1: “NR; User Equipment (UE) conformance specification; Radio transmission and reception; Part 1: Range 1 Standalone”.

[6]	3GPP TS 38.521-3: “NR; User Equipment (UE) conformance specification; Radio transmission and reception; Part 3: Range 1 and Range 2 Interworking operation with other radios”.

[7]	3GPP TS 38.508-1: “5GS; User Equipment (UE) conformance specification; Part 1: Common test environment “.

[8]	3GPP TR 25.914: “Measurements of radio performances for UMTS terminals in speech mode”.

[9]	IEEE Std 149: “IEEE Standard Test Procedures for Antennas”, IEEE.

[10]	JCGM 100:2008: “Evaluation of measurement data — Guide to the expression of uncertainty in measurement”.

[11]	ETSI TR 102 273-1-1: “Electromagnetic compatibility and Radio spectrum Matters (ERM); Improvement on Radiated Methods of Measurement (using test site) and evaluation of the corresponding measurement uncertainties; Part 1: Uncertainties in the measurement of mobile radio equipment characteristics; Sub-part 1: Introduction”.

[12]	ETSI TR 100 028-2: “ElectroMagnetic Compatibility and Radio Spectrum Matters (ERM); Uncertainties in the measurement of mobile radio equipment characteristics; Part 2”.

[13]	ETSI TR 102 273-1-2: “Electromagnetic compatibility and Radio spectrum Matters (ERM); Improvement on Radiated Methods of Measurement (using test site) and evaluation of the corresponding measurement uncertainties; Part 1: Uncertainties in the measurement of mobile radio equipment characteristics; Sub-part 2: Examples and annexes”.

[14]	CTIA Certification™: “CTIA Certification Test Plan for Wireless Device Over-the-Air Performance, CTIA 01.71 Device Setup and Positioning Guidelines”, latest active version available at: https://ctiacertification.org/test-plans/

[15]	Foegelle, M.D., “The Surface Standard Deviation Method for TRP Measurement Uncertainty”, 25th Proceedings of the Antenna Measurement Techniques Association (AMTA 2003), A03-027

[16]	3GPP TR 37.902: “Measurements of User Equipment (UE) radio performances for LTE/UMTS terminals; Total Radiated Power (TRP) and Total Radiated Sensitivity (TRS) test “.

[17]		3GPP TS 37.544: “Universal Terrestrial Radio Access (UTRA) and Evolved Universal Terrestrial Radio Access (E-UTRA); User Equipment (UE) Over The Air (OTA) performance; Conformance testing “.

[18]	3GPP TR 37.941: “Radio Frequency (RF) conformance testing background for radiated Base Station (BS) requirements”

[19]	3GPP TR 38.810: “NR; Study on test methods”

[20]	3GPP TR 38.903, “NR; Derivation of test tolerances and measurement uncertainty for User Equipment (UE) conformance test cases”

[21]	“Reverberation Chamber Metrology for Wireless Internet of Things Devices”, Anouk Hubrechsen, Kate A. Remley and Sara Catteau, IEEE Microwave Magazine, February 2022, pp.75-85

[22]	“Proximity and antenna orientation effects for large-form-factor devices in a reverberation chamber” Willem T. C. Burger, Kate A. Remley, Christopher L. Holloway, John M. Ladbury, 2013 IEEE International Symposium on Electromagnetic Compatibility, pp.671-676

[23]	“A Significance Test for Reverberation-Chamber Measurement Uncertainty in Total Radiated Power of Wireless Devices”, Kate A. Remley, Chih-Ming Jack Wang, Dylan F. Williams, Johannes J. aan den Toorn and Christopher L. Holloway, IEEE TRANSACTIONS ON ELECTROMAGNETIC COMPATIBILITY, VOL. 58, NO. 1, FEBRUARY 2016, pp.207-219

[24]	3GPP TR 38.834: “Measurements of User Equipment (UE) Over-the-Air (OTA) performance for NR FR1; Total Radiated Power (TRP) and Total Radiated Sensitivity (TRS) test methodology (Release 17)”

[25]	CTIA Certification™: “CTIA Certification Test Plan for Wireless Device Over-the-Air Performance, CTIA 01.72: Near-Field Phantoms”, latest active version available at: https://ctiacertification.org/test-plans/

[26]	CTIA Certification™: “CTIA Certification Test Plan for Wireless Device Over-the-Air Performance, CTIA 01.70: Measurement Uncertainty”, latest active version available at: https://ctiacertification.org/test-plans/

# 3	Definitions of terms, symbols and abbreviations

## 3.1	Terms

For the purposes of the present document, the terms given in 3GPP TR 21.905 [1] and the following apply. A term defined in the present document takes precedence over the definition of the same term, if any, in 3GPP TR 21.905 [1].

Browsing mode usage: This mode corresponds to “data” mode, the device is tested via hand-only phantoms.

Primary mechanical mode: The mode that is most often used for a specific user scenario. Every terminal has at least one primary mechanical mode, if multiple modes are supported, different primary mechanical modes may be applicable for different user scenarios, e.g., different primary mechanical modes for Browsing mode usage and Talk mode usage for the same UE.

Talk mode usage: This mode corresponds to “talk” mode, the device is tested via head&hand phantoms.

Wrist-worn mode usage: This mode corresponds to wearable device for wrist-worn mode, the device is tested via forearm phantoms.

## 3.2	Symbols

For the purposes of the present document, the following symbols apply:

## 3.3	Abbreviations

For the purposes of the present document, the abbreviations given in 3GPP TR 21.905 [1] and the following apply. An abbreviation defined in the present document takes precedence over the definition of the same abbreviation, if any, in 3GPP TR 21.905 [1].

AC	Anechoic Chamber

BHHL	Beside Head and Hand Left Side (Head and Hand Phantom)

BHHR	Beside Head and Hand Right Side (Head and Hand Phantom)

CA	Carrier Aggregation

DUT	Device Under Test

E-UTRA	Evolved UTRA

EIRP 	Effective Isotropic Radiated Power

EUT	Equipment Under Test

FR1	Frequency Range 1

FS	Free Space

HL	Hand Left (Hand Phantom Only)

HR	Hand Right (Hand Phantom Only)

MPR	Allowed maximum power reduction

NR	New Radio

NSA	Non-Standalone, a mode of operation where operation of an other radio is assisted with an other radio

OTA	Over The Air

QZ	Quiet Zone

RedCap	Reduced Capability

SA	Standalone

SS	System Simulator

TAA	Time-Averaging Algorithm

TAS	Tx Antenna Switching

TRP	Total Radiated Power

TRS	Total Radiated Sensitivity

TxD	Tx Diversity

UE	User Equipment

UL MIMO	Uplink Multiple Antenna transmission

# 4	General

## 4.1	Device types

The following device types are within the scope of enhanced FR1 TRP TRS WI:

-	Smartphone

-	Considering UEs with antenna configurations of 1Tx, 2Tx, 2 Rx and 4 Rx

-	Wearable Redcap UE

-	Tablet

-	Laptop embedded equipment (LEE)

-	Laptop mounted equipment (LME)

## 4.2	Testing configuration

### 4.2.1	UE use scenarios for TRP TRS test

The following use scenarios are considered for TRP TRS test:

-	Talk mode using head & hand phantom for narrow phones between 56 mm and 72 mm and for wide phones with a width >72 mm and <92 mm.

-	Browsing mode using hand phantom for narrow and wide phones

-	Using forearm phantom for wrist-worn devices

-	Free Space is used for devices not used in above-mentioned scenarios, other phantoms are not precluded for wearable devices

For smartphones, both browsing mode and talk mode shall be covered.

For wrist-worn Redcap devices, forearm phantom is the first priority. FFS other Redcap form factor devices.

For other device types, free space (FS) testing configuration is the first priority.

### 4.2.2	UE mechanical mode description

The mechanical modes of a device under test (DUT) are declared by the manufacturer. A DUT shall have at least one mechanical mode. If only one mode is supported, then this is defined as the primary. If multiple modes are supported, the manufacturer can declare different primary mechanical modes applicable for different user scenarios, e.g., different primary mechanical mode for Browsing mode usage and Talk mode usage for the same UE.

## 4.3	Testing bands

<Editor’s note: parameters for EN-DC and CA bands can be further added. >

### 4.3.1	General

The frequency ranges in which NR can operate are identified as described in Table 4.3.1-1.

Table 4.3.1-1: Definition of frequency ranges

| Frequency range designation | Corresponding frequency range |
| --- | --- |
| FR1 | 410 MHz – 7125 MHz |
| FR2 | 24250 MHz – 52600 MHz |

The present technical report covers FR1 operating bands and FR1 non-standalone (NSA) operation mode with E-UTRA.

### 4.3.2	Operating bands

Operating bands for NR FR1 are defined in Table 5.2-1 and Carrier Aggregation (CA) are defined in Clause 5.5A in TS 38.101-1 [3]. The operating bands for EN-DC are defined in Clause 5.5B in TS 38.101-3 [4].

### 4.3.3	Test parameters for each band

The detailed test parameters for each band are defined in Table 4.3.3-1 and Table 4.3.3-2.

Table 4.3.3-1: NR FR1 TRP measurement parameters

| NR Band | CBW [MHz] | SCS (kHz) | UL modulation | Range | UL Carrier centre [ARFCN] | UL Carrier Center (MHz) | DL Carrier centre [ARFCN] | DL Carrier Center (MHz) | UL RB Allocation (LCRB @ RBstart) | DL configuration |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| n1 | 15 | 15 | DFT-s-OFDM QPSK | Low | 385500 | 1927.5 | 423500 | 2117.5 | 36@18 | N/A |
|  |  |  |  | Mid | 390000 | 1950 | 428000 | 2140 |  |  |
|  |  |  |  | High | 394500 | 1972.5 | 432500 | 2162.5 |  |  |
| n2 | 15 | 15 | DFT-s-OFDM QPSK | Low | 371500 | 1857.5 | 387500 | 1937.5 | 36@18 | N/A |
|  |  |  |  | Mid | 376000 | 1880 | 392000 | 1960 |  |  |
|  |  |  |  | High | 380500 | 1902.5 | 396500 | 1982.5 |  |  |
| n3 | 20 | 15 | DFT-s-OFDM QPSK | Low | 344000 | 1720 | 363000 | 1815 | 50@25 | N/A |
|  |  |  |  | Mid | 349500 | 1747.5 | 368500 | 1842.5 |  |  |
|  |  |  |  | High | 355000 | 1775 | 374000 | 1870 |  |  |
| n5 | 15 | 15 | DFT-s-OFDM QPSK | Low | 166300 | 831.5 | 175300 | 876.5 | 36@18 | N/A |
|  |  |  |  | Mid | 167300 | 836.5 | 176300 | 881.5 |  |  |
|  |  |  |  | High | 168300 | 841.5 | 177300 | 886.5 |  |  |
| n7 | 15 | 15 | DFT-s-OFDM QPSK | Low | 501500 | 2507.5 | 525500 | 2627.5 | 36@18 | N/A |
|  |  |  |  | Mid | 507000 | 2535 | 531000 | 2655 |  |  |
|  |  |  |  | High | 512500 | 2562.5 | 536500 | 2682.5 |  |  |
| n8 | 15 | 15 | DFT-s-OFDM QPSK | Low | 177500 | 887.5 | 186500 | 932.5 | 36@18 | N/A |
|  |  |  |  | Mid | 179500 | 897.5 | 188500 | 942.5 |  |  |
|  |  |  |  | High | 181500 | 907.5 | 190500 | 952.5 |  |  |
| n12 | 10 | 15 | DFT-s-OFDM QPSK | Low | 140800 | 704 | 146800 | 734 | 25@12 | N/A |
|  |  |  |  | Mid | 141500 | 707.5 | 147500 | 737.5 |  |  |
|  |  |  |  | High | 142200 | 711 | 148200 | 741 |  |  |
| n14 | 10 | 15 | DFT-s-OFDM QPSK | Low | 158600 | 793 | 152600 | 763 | 25@12 | N/A |
|  |  |  |  | Mid |  |  |  |  |  |  |
|  |  |  |  | High |  |  |  |  |  |  |
| n20 | 15 | 15 | DFT-s-OFDM QPSK | Low | 167900 | 839.5 | 159700 | 798.5 | 36@18 | N/A |
|  |  |  |  | Mid | 169400 | 847 | 161200 | 806 |  |  |
|  |  |  |  | High | 170900 | 854.5 | 162700 | 813.5 |  |  |
| n25 | 15 | 15 | DFT-s-OFDM QPSK | Low | 371500 | 1857.5 | 387500 | 1937.5 | 36@18 | N/A |
|  |  |  |  | Mid | 376500 | 1882.5 | 392500 | 1962.5 |  |  |
|  |  |  |  | High | 381500 | 1907.5 | 397500 | 1987.5 |  |  |
| n26 | 10 | 15 | DFT-s-OFDM QPSK | Low | 163800 | 819 | 172800 | 864 | 25@12 | N/A |
|  |  |  |  | Mid | 166300 | 831.5 | 175300 | 876.5 |  |  |
|  |  |  |  | High | 168800 | 844 | 177800 | 889 |  |  |
| n28 | 20 | 15 | DFT-s-OFDM QPSK | Low | 142600 | 713 | 153600 | 768 | 50@25 | N/A |
|  |  |  |  | Mid | 145600 | 728 | 156600 | 783 |  |  |
|  |  |  |  | High | 147600 | 738 | 158600 | 793 |  |  |
| n30 | 10 | 15 | DFT-s-OFDM QPSK | Low | 462000 | 2310 | 471000 | 2355 | 25@12 | N/A |
|  |  |  |  | Mid |  |  |  |  |  |  |
|  |  |  |  | High |  |  |  |  |  |  |
| n34 | 10 | 15 | DFT-s-OFDM QPSK | Low | 403000 | 2015 | 403000 | 2015 | 25@12 | N/A |
|  |  |  |  | Mid | 403500 | 2017.5 | 403500 | 2017.5 |  |  |
|  |  |  |  | High | 404000 | 2020 | 404000 | 2020 |  |  |
| n38 | 15 | 15 | DFT-s-OFDM QPSK | Low | 515500 | 2577.5 | 515500 | 2577.5 | 36@18 | N/A |
|  |  |  |  | Mid | 519000 | 2595 | 519000 | 2595 |  |  |
|  |  |  |  | High | 522500 | 2612.5 | 522500 | 2612.5 |  |  |
| n39 | 20 | 15 | DFT-s-OFDM QPSK | Low | 378000 | 1890 | 378000 | 1890 | 50@25 | N/A |
|  |  |  |  | Mid | 380000 | 1900 | 380000 | 1900 |  |  |
|  |  |  |  | High | 382000 | 1910 | 382000 | 1910 |  |  |
| n40 | 30 | 15 | DFT-s-OFDM QPSK | Low | 463000 | 2315 | 463000 | 2315 | TBD | N/A |
|  |  |  |  | Mid | 470000 | 2350 | 470000 | 2350 |  |  |
|  |  |  |  | High | 477000 | 2385 | 477000 | 2385 |  |  |
| n41 | 100 | 30 | DFT-s-OFDM QPSK | Low | 509202 | 2546.01 | 509202 | 2546.01 | 135@67 | N/A |
|  |  |  |  | Mid | 518598 | 2592.99 | 518598 | 2592.99 |  |  |
|  |  |  |  | High | 528000 | 2640 | 528000 | 2640 |  |  |
| n48 | 20 | 15 | DFT-s-OFDM QPSK | Low | 637334 | 3560.01 | 637334 | 3560.01 | 50@25 | N/A |
|  |  |  |  | Mid | 641666 | 3624.99 | 641666 | 3624.99 |  |  |
|  |  |  |  | High | 646000 | 3690 | 646000 | 3690 |  |  |
| n50 | 20 | 15 | DFT-s-OFDM QPSK | Low | 288400 | 1442 | 288400 | 1442 | 50@25 | N/A |
|  |  |  |  | Mid | 294900 | 1474.5 | 294900 | 1474.5 |  |  |
|  |  |  |  | High | 301400 | 1507 | 301400 | 1507 |  |  |
| n51 | 5 | 15 | DFT-s-OFDM QPSK | Low | 285900 | 1429.5 | 285900 | 1429.5 | 12@6 | N/A |
|  |  |  |  | Mid |  |  |  |  |  |  |
|  |  |  |  | High |  |  |  |  |  |  |
| n53 | 10 | 15 | DFT-s-OFDM QPSK | Low | 497700 | 2488.5 | 497700 | 2488.5 | 25@12 | N/A |
|  |  |  |  | Mid | 497860 | 2489.3 | 497860 | 2489.3 |  |  |
|  |  |  |  | High | 498000 | 2490 | 498000 | 2490 |  |  |
| n65 | 15 | 15 | DFT-s-OFDM QPSK | Low | 423500 | 2117.5 | 423500 | 2117.5 | 36@18 | N/A |
|  |  |  |  | Mid | 431000 | 2155 | 431000 | 2155 |  |  |
|  |  |  |  | High | 438500 | 2192.5 | 438500 | 2192.5 |  |  |
| n66 | 20 (20+20) | 15 | DFT-s-OFDM QPSK | Low | 344000 | 1720 | 424000 | 2120 | 50@25 | N/A |
|  |  |  |  | Mid | 349000 | 1745 | 429000 | 2145 |  |  |
|  |  |  |  | High | 354000 | 1770 | 434000 | 2170 |  |  |
| n70 | 15 (15+15) | 15 | DFT-s-OFDM QPSK | Low | 340500 | 1702.5 | 400500 | 2002.5 | 36@18 | N/A |
|  |  |  |  | Mid |  |  |  |  |  |  |
|  |  |  |  | High |  |  |  |  |  |  |
| n71 | 10 | 15 | DFT-s-OFDM QPSK | Low | 133600 | 668 | 124400 | 622 | 25@12 | N/A |
|  |  |  |  | Mid | 136100 | 680.5 | 126900 | 634.5 |  |  |
|  |  |  |  | High | 138600 | 693 | 129400 | 647 |  |  |
| n74 | 15 | 15 | DFT-s-OFDM QPSK | Low | 286900 | 1434.5 | 296500 | 1482.5 | 36@18 | N/A |
|  |  |  |  | Mid | 289700 | 1448.5 | 299300 | 1496.5 |  |  |
|  |  |  |  | High | 292500 | 1462.5 | 302100 | 1510.5 |  |  |
| n77 | 100 | 30 | DFT-s-OFDM QPSK | Low | 623334 | 3350.01 | 623334 | 3350.01 | 135@67 | N/A |
|  |  |  |  | Mid | 650000 | 3750 | 650000 | 3750 |  |  |
|  |  |  |  | High | 676666 | 4149.99 | 676666 | 4149.99 |  |  |
| n78 | 100 | 30 | DFT-s-OFDM QPSK | Low | 623334 | 3350.01 | 623334 | 3350.01 | 135@67 | N/A |
|  |  |  |  | Mid | 636666 | 3549.99 | 636666 | 3549.99 |  |  |
|  |  |  |  | High | 650000 | 3750 | 650000 | 3750 |  |  |
| n79 | 100 | 30 | DFT-s-OFDM QPSK | Low | 696668 | 4450.02 | 696668 | 4450.02 | 135@67 | N/A |
|  |  |  |  | Mid | 713334 | 4700.01 | 713334 | 4700.01 |  |  |
|  |  |  |  | High | 730000 | 4950 | 730000 | 4950 |  |  |
| n80 | 20 | 15 | DFT-s-OFDM QPSK | Low | 344000 | 1720 | N/A | N/A | 50@25 | N/A |
|  |  |  |  | Mid | 349500 | 1747.5 | N/A | N/A |  |  |
|  |  |  |  | High | 355000 | 1775 | N/A | N/A |  |  |
| n81 | 15 | 15 | DFT-s-OFDM QPSK | Low | 177500 | 887.5 | N/A | N/A | 36@18 | N/A |
|  |  |  |  | Mid | 179500 | 897.5 | N/A | N/A |  |  |
|  |  |  |  | High | 181500 | 907.5 | N/A | N/A |  |  |
| n82 | 15 | 15 | DFT-s-OFDM QPSK | Low | 167900 | 839.5 | N/A | N/A | 36@18 | N/A |
|  |  |  |  | Mid | 169400 | 847 | N/A | N/A |  |  |
|  |  |  |  | High | 170900 | 854.5 | N/A | N/A |  |  |
| n83 | 15 | 15 | DFT-s-OFDM QPSK | Low | 142100 | 710.5 | N/A | N/A | 36@18 | N/A |
|  |  |  |  | Mid | 145100 | 725.5 | N/A | N/A |  |  |
|  |  |  |  | High | 148100 | 740.5 | N/A | N/A |  |  |
| n84 | 15 | 15 | DFT-s-OFDM QPSK | Low | 385500 | 1927.5 | N/A | N/A | 36@18 | N/A |
|  |  |  |  | Mid | 390000 | 1950 | N/A | N/A |  |  |
|  |  |  |  | High | 394500 | 1972.5 | N/A | N/A |  |  |
| n86 | 20 | 15 | DFT-s-OFDM QPSK | Low | 344000 | 1720 | N/A | N/A | 50@25 | N/A |
|  |  |  |  | Mid | 349000 | 1745 | N/A | N/A |  |  |
|  |  |  |  | High | 354000 | 1770 | N/A | N/A |  |  |
| N95 | 10 | 15 | DFT-s-OFDM QPSK | Low | 403000 | 2015 | N/A | N/A | 50@25 | N/A |
|  |  |  |  | Mid | 403500 | 2017.5 | N/A | N/A |  |  |
|  |  |  |  | High | 404000 | 2020 | N/A | N/A |  |  |

Table 4.3.3-2: NR FR1 TRS measurement parameters

| NR Band | CBW (MHz) | SCS (kHz) | DL modulation | UL modulation | Range | UL Carrier centre [ARFCN] | UL Carrier Center (MHz) | DL Carrier centre [ARFCN] | DL Carrier Center (MHz) | UL RB Allocation (LCRB @ RBstart) | DL Configuration (FULL RB, LCRB @ RBstart) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| n1 | 15 | 15 | CP-OFDM QPSK | DFT-s-OFDM QPSK | Low | 385500 | 1927.5 | 423500 | 2117.5 | 75@4 | 79@0 |
|  |  |  |  |  | Mid | 390000 | 1950 | 428000 | 2140 |  |  |
|  |  |  |  |  | High | 394500 | 1972.5 | 432500 | 2162.5 |  |  |
| n2 | 15 | 15 | CP-OFDM QPSK | DFT-s-OFDM QPSK | Low | 371500 | 1857.5 | 387500 | 1937.5 | 50@29 | 79@0 |
|  |  |  |  |  | Mid | 376000 | 1880 | 392000 | 1960 |  |  |
|  |  |  |  |  | High | 380500 | 1902.5 | 396500 | 1982.5 |  |  |
| n3 | 20 | 15 | CP-OFDM QPSK | DFT-s-OFDM QPSK | Low | 344000 | 1720 | 363000 | 1815 | 50@56 | 106@0 |
|  |  |  |  |  | Mid | 349500 | 1747.5 | 368500 | 1842.5 |  |  |
|  |  |  |  |  | High | 355000 | 1775 | 374000 | 1870 |  |  |
| n5 | 15 | 15 | CP-OFDM QPSK | DFT-s-OFDM QPSK | Low | 166300 | 831.5 | 175300 | 876.5 | 25@54 | 79@0 |
|  |  |  |  |  | Mid | 167300 | 836.5 | 176300 | 881.5 |  |  |
|  |  |  |  |  | High | 168300 | 841.5 | 177300 | 886.5 |  |  |
| n7 | 15 | 15 | CP-OFDM QPSK | DFT-s-OFDM QPSK | Low | 501500 | 2507.5 | 525500 | 2627.5 | 75@4 | 79@0 |
|  |  |  |  |  | Mid | 507000 | 2535 | 531000 | 2655 |  |  |
|  |  |  |  |  | High | 512500 | 2562.5 | 536500 | 2682.5 |  |  |
| n8 | 15 | 15 | CP-OFDM QPSK | DFT-s-OFDM QPSK | Low | 177500 | 887.5 | 186500 | 932.5 | 25@54 | 79@0 |
|  |  |  |  |  | Mid | 179500 | 897.5 | 188500 | 942.5 |  |  |
|  |  |  |  |  | High | 181500 | 907.5 | 190500 | 952.5 |  |  |
| n12 | 10 | 15 | CP-OFDM QPSK | DFT-s-OFDM QPSK | Low | 140800 | 704 | 146800 | 734 | 20@32 | 52@0 |
|  |  |  |  |  | Mid | 141500 | 707.5 | 147500 | 737.5 |  |  |
|  |  |  |  |  | High | 142200 | 711 | 148200 | 741 |  |  |
| n14 | 10 | 15 | CP-OFDM QPSK | DFT-s-OFDM QPSK | Low | 158600 | 793 | 152600 | 763 | 20@32 | 52@0 |
|  |  |  |  |  | Mid |  |  |  |  |  |  |
|  |  |  |  |  | High |  |  |  |  |  |  |
| n20 | 15 | 15 | CP-OFDM QPSK | DFT-s-OFDM QPSK | Low | 167900 | 839.5 | 159700 | 798.5 | 20@11 | 79@0 |
|  |  |  |  |  | Mid | 169400 | 847 | 161200 | 806 |  |  |
|  |  |  |  |  | High | 170900 | 854.5 | 162700 | 813.5 |  |  |
| n25 | 15 | 15 | CP-OFDM QPSK | DFT-s-OFDM QPSK | Low | 371500 | 1857.5 | 387500 | 1937.5 | 50@29 | 79@0 |
|  |  |  |  |  | Mid | 376500 | 1882.5 | 392500 | 1962.5 |  |  |
|  |  |  |  |  | High | 381500 | 1907.5 | 397500 | 1987.5 |  |  |
| n26 | 10 | 15 | CP-OFDM QPSK | DFT-s-OFDM QPSK | Low | 163800 | 819 | 172800 | 864 | 25@27 | 52@0 |
|  |  |  |  |  | Mid | 166300 | 831.5 | 175300 | 876.5 |  |  |
|  |  |  |  |  | High | 168800 | 844 | 177800 | 889 |  |  |
| n28 | 20 | 15 | CP-OFDM QPSK | DFT-s-OFDM QPSK | Low | 142600 | 713 | 153600 | 768 | 25@81 | 106@0 |
|  |  |  |  |  | Mid | 145600 | 728 | 156600 | 783 |  |  |
|  |  |  |  |  | High | 147600 | 738 | 158600 | 793 |  |  |
| n30 | 10 | 15 | CP-OFDM QPSK | DFT-s-OFDM QPSK | Low | 462000 | 2310 | 471000 | 2355 | 20@32 | 52@0 |
|  |  |  |  |  | Mid |  |  |  |  |  |  |
|  |  |  |  |  | High |  |  |  |  |  |  |
| n34 | 10 | 15 | CP-OFDM QPSK | DFT-s-OFDM QPSK | Low | 403000 | 2015 | 403000 | 2015 | 50@0 | 52@0 |
|  |  |  |  |  | Mid | 403500 | 2017.5 | 403500 | 2017.5 |  |  |
|  |  |  |  |  | High | 404000 | 2020 | 404000 | 2020 |  |  |
| n38 | 15 | 15 | CP-OFDM QPSK | DFT-s-OFDM QPSK | Low | 515500 | 2577.5 | 515500 | 2577.5 | 75@0 | 79@0 |
|  |  |  |  |  | Mid | 519000 | 2595 | 519000 | 2595 |  |  |
|  |  |  |  |  | High | 522500 | 2612.5 | 522500 | 2612.5 |  |  |
| n39 | 20 | 15 | CP-OFDM QPSK | DFT-s-OFDM QPSK | Low | 378000 | 1890 | 378000 | 1890 | 100@0 | 106@0 |
|  |  |  |  |  | Mid | 380000 | 1900 | 380000 | 1900 |  |  |
|  |  |  |  |  | High | 382000 | 1910 | 382000 | 1910 |  |  |
| n40 | 30 | 15 | CP-OFDM QPSK | DFT-s-OFDM QPSK | Low | 463000 | 2315 | 463000 | 2315 | 160@0 | 160@0 |
|  |  |  |  |  | Mid | 470000 | 2350 | 470000 | 2350 |  |  |
|  |  |  |  |  | High | 477000 | 2385 | 477000 | 2385 |  |  |
| n41 | 100 | 30 | CP-OFDM QPSK | DFT-s-OFDM QPSK | Low | 509202 | 2546.01 | 509202 | 2546.01 | 270@0 | 273@0 |
|  |  |  |  |  | Mid | 518598 | 2592.99 | 518598 | 2592.99 |  |  |
|  |  |  |  |  | High | 528000 | 2640 | 528000 | 2640 |  |  |
| n48 | 20 | 15 | CP-OFDM QPSK | DFT-s-OFDM QPSK | Low | 637334 | 3560.01 | 637334 | 3560.01 | 100@0 | 106@0 |
|  |  |  |  |  | Mid | 641666 | 3624.99 | 641666 | 3624.99 |  |  |
|  |  |  |  |  | High | 646000 | 3690 | 646000 | 3690 |  |  |
| n50 | 20 | 15 | CP-OFDM QPSK | DFT-s-OFDM QPSK | Low | 288400 | 1442 | 288400 | 1442 | 100@0 | 106@0 |
|  |  |  |  |  | Mid | 294900 | 1474.5 | 294900 | 1474.5 |  |  |
|  |  |  |  |  | High | 301400 | 1507 | 301400 | 1507 |  |  |
| n51 | 5 | 15 | CP-OFDM QPSK | DFT-s-OFDM QPSK | Low | 285900 | 1429.5 | 285900 | 1429.5 | 25@0 | 25@0 |
|  |  |  |  |  | Mid |  |  |  |  |  |  |
|  |  |  |  |  | High |  |  |  |  |  |  |
| n53 | 10 | 15 | CP-OFDM QPSK | DFT-s-OFDM QPSK | Low | 497700 | 2488.5 | 497700 | 2488.5 | 50@0 | 52@0 |
|  |  |  |  |  | Mid | 497860 | 2489.3 | 497860 | 2489.3 |  |  |
|  |  |  |  |  | High | 498000 | 2490 | 498000 | 2490 |  |  |
| n65 | 15 | 15 | CP-OFDM QPSK | DFT-s-OFDM QPSK | Low | 423500 | 2117.5 | 423500 | 2117.5 | 75@4 | 79@0 |
|  |  |  |  |  | Mid | 431000 | 2155 | 431000 | 2155 |  |  |
|  |  |  |  |  | High | 438500 | 2192.5 | 438500 | 2192.5 |  |  |
| n66 | 20 (20+20) | 15 | CP-OFDM QPSK | DFT-s-OFDM QPSK | Low | 344000 | 1720 | 424000 | 2120 | 100@6 | 106@0 |
|  |  |  |  |  | Mid | 349000 | 1745 | 429000 | 2145 |  |  |
|  |  |  |  |  | High | 354000 | 1770 | 434000 | 2170 |  |  |
| n70 | 15 | 15 | CP-OFDM QPSK | DFT-s-OFDM QPSK | Low | 340500 | 1702.5 | 400500 | 2002.5 | 75@4 | 79@0 |
|  |  |  |  |  | Mid |  |  |  |  |  |  |
|  |  |  |  |  | High |  |  |  |  |  |  |
| n71 | 10 | 15 | CP-OFDM QPSK | DFT-s-OFDM QPSK | Low | 133600 | 668 | 124400 | 622 | 25@0 | 52@0 |
|  |  |  |  |  | Mid | 136100 | 680.5 | 126900 | 634.5 |  |  |
|  |  |  |  |  | High | 138600 | 693 | 129400 | 647 |  |  |
| n74 | 15 | 15 | CP-OFDM QPSK | DFT-s-OFDM QPSK | Low | 286900 | 1434.5 | 296500 | 1482.5 | 25@54 | 79@0 |
|  |  |  |  |  | Mid | 289700 | 1448.5 | 299300 | 1496.5 |  |  |
|  |  |  |  |  | High | 292500 | 1462.5 | 302100 | 1510.5 |  |  |
| n75 SDL | 15 | 15 | CP-OFDM QPSK | DFT-s-OFDM QPSK | Low | N/A | N/A | 287900 | 1439.5 | NA | 79@0 |
|  |  |  |  |  | Mid | N/A | N/A | 294900 | 1474.5 |  |  |
|  |  |  |  |  | High | N/A | N/A | 301900 | 1509.5 |  |  |
| n76 SDL | 5 | 15 | CP-OFDM QPSK | DFT-s-OFDM QPSK | Low | N/A | N/A | 285900 | 1429.5 | NA | 25@0 |
|  |  |  |  |  | Mid |  |  |  |  |  |  |
|  |  |  |  |  | High |  |  |  |  |  |  |
| n77 | 100 | 30 | CP-OFDM QPSK | DFT-s-OFDM QPSK | Low | 623334 | 3350.01 | 623334 | 3350.01 | 270@0 | 273@0 |
|  |  |  |  |  | Mid | 650000 | 3750 | 650000 | 3750 |  |  |
|  |  |  |  |  | High | 676666 | 4149.99 | 676666 | 4149.99 |  |  |
| n78 | 100 | 30 | CP-OFDM QPSK | DFT-s-OFDM QPSK | Low | 623334 | 3350.01 | 623334 | 3350.01 | 270@0 | 273@0 |
|  |  |  |  |  | Mid | 636666 | 3549.99 | 636666 | 3549.99 |  |  |
|  |  |  |  |  | High | 650000 | 3750 | 650000 | 3750 |  |  |
| n79 | 100 | 30 | CP-OFDM QPSK | DFT-s-OFDM QPSK | Low | 696668 | 4450.02 | 696668 | 4450.02 | 270@0 | 273@0 |
|  |  |  |  |  | Mid | 713334 | 4700.01 | 713334 | 4700.01 |  |  |
|  |  |  |  |  | High | 730000 | 4950 | 730000 | 4950 |  |  |

The detailed testing parameters for each band for RedCap UE is defined in Table 4.3.3-3 and Table 4.3.3-4.

Table 4.3.3-3: NR FR1 TRP measurement parameters for RedCap UE

| NR Band | CBW [MHz] | SCS (kHz) | UL modulation | Range | UL Carrier centre [ARFCN] | UL Carrier Center (MHz) | DL Carrier centre [ARFCN] | DL Carrier Center (MHz) | UL RB Allocation (LCRB @ RBstart) | DL configuration |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| n1 | 15 | 15 | DFT-s-OFDM QPSK | Low | 385500 | 1927.5 | 423500 | 2117.5 | 36@18 | N/A |
|  |  |  |  | Mid | 390000 | 1950 | 428000 | 2140 |  |  |
|  |  |  |  | High | 394500 | 1972.5 | 432500 | 2162.5 |  |  |
| n2 | 15 | 15 | DFT-s-OFDM QPSK | Low | 371500 | 1857.5 | 387500 | 1937.5 | 36@18 | N/A |
|  |  |  |  | Mid | 376000 | 1880 | 392000 | 1960 |  |  |
|  |  |  |  | High | 380500 | 1902.5 | 396500 | 1982.5 |  |  |
| n3 | 15 | 15 | DFT-s-OFDM QPSK | Low | 343500 | 1717.5 | 362500 | 1812.5 | 36@18 | N/A |
|  |  |  |  | Mid | 349500 | 1747.5 | 368500 | 1842.5 |  |  |
|  |  |  |  | High | 355500 | 1777.5 | 374500 | 1872.5 |  |  |
| n5 | 15 | 15 | DFT-s-OFDM QPSK | Low | 166300 | 831.5 | 175300 | 876.5 | 36@18 | N/A |
|  |  |  |  | Mid | 167300 | 836.5 | 176300 | 881.5 |  |  |
|  |  |  |  | High | 168300 | 841.5 | 177300 | 886.5 |  |  |
| n7 | 15 | 15 | DFT-s-OFDM QPSK | Low | 501500 | 2507.5 | 525500 | 2627.5 | 36@18 | N/A |
|  |  |  |  | Mid | 507000 | 2535 | 531000 | 2655 |  |  |
|  |  |  |  | High | 512500 | 2562.5 | 536500 | 2682.5 |  |  |
| n8 | 15 | 15 | DFT-s-OFDM QPSK | Low | 177500 | 887.5 | 186500 | 932.5 | 36@18 | N/A |
|  |  |  |  | Mid | 179500 | 897.5 | 188500 | 942.5 |  |  |
|  |  |  |  | High | 181500 | 907.5 | 190500 | 952.5 |  |  |
| n12 | 10 | 15 | DFT-s-OFDM QPSK | Low | 140800 | 704 | 146800 | 734 | 25@12 | N/A |
|  |  |  |  | Mid | 141500 | 707.5 | 147500 | 737.5 |  |  |
|  |  |  |  | High | 142200 | 711 | 148200 | 741 |  |  |
| n14 | 10 | 15 | DFT-s-OFDM QPSK | Low | 158600 | 793 | 152600 | 763 | 25@12 | N/A |
|  |  |  |  | Mid |  |  |  |  |  |  |
|  |  |  |  | High |  |  |  |  |  |  |
| n20 | 15 | 15 | DFT-s-OFDM QPSK | Low | 167900 | 839.5 | 159700 | 798.5 | 36@18 | N/A |
|  |  |  |  | Mid | 169400 | 847 | 161200 | 806 |  |  |
|  |  |  |  | High | 170900 | 854.5 | 162700 | 813.5 |  |  |
| n25 | 15 | 15 | DFT-s-OFDM QPSK | Low | 371500 | 1857.5 | 387500 | 1937.5 | 36@18 | N/A |
|  |  |  |  | Mid | 376500 | 1882.5 | 392500 | 1962.5 |  |  |
|  |  |  |  | High | 381500 | 1907.5 | 397500 | 1987.5 |  |  |
| n26 | 15 | 15 | DFT-s-OFDM QPSK | Low | 164300 | 821.5 | 173300 | 866.5 | 36@18 | N/A |
|  |  |  |  | Mid | 166300 | 831.5 | 175300 | 876.5 |  |  |
|  |  |  |  | High | 168300 | 841.5 | 177300 | 886.5 |  |  |
| n28 | 15 | 15 | DFT-s-OFDM QPSK | Low | 142100 | 710.5 | 153100 | 765.5 | 36@18 | N/A |
|  |  |  |  | Mid | 145100 | 725.5 | 156100 | 780.5 |  |  |
|  |  |  |  | High | 148100 | 740.5 | 159100 | 795.5 |  |  |
| n30 | 10 | 15 | DFT-s-OFDM QPSK | Low | 462000 | 2310 | 471000 | 2355 | 25@12 | N/A |
|  |  |  |  | Mid |  |  |  |  |  |  |
|  |  |  |  | High |  |  |  |  |  |  |
| n34 | 10 | 15 | DFT-s-OFDM QPSK | Low | 403000 | 2015 | 403000 | 2015 | 25@12 | N/A |
|  |  |  |  | Mid | 403500 | 2017.5 | 403500 | 2017.5 |  |  |
|  |  |  |  | High | 404000 | 2020 | 404000 | 2020 |  |  |
| n38 | 15 | 15 | DFT-s-OFDM QPSK | Low | 515500 | 2577.5 | 515500 | 2577.5 | 36@18 | N/A |
|  |  |  |  | Mid | 519000 | 2595 | 519000 | 2595 |  |  |
|  |  |  |  | High | 522500 | 2612.5 | 522500 | 2612.5 |  |  |
| n39 | 15 | 15 | DFT-s-OFDM QPSK | Low | 377500 | 1887.5 | 377500 | 1887.5 | 36@18 | N/A |
|  |  |  |  | Mid | 380000 | 1900 | 380000 | 1900 |  |  |
|  |  |  |  | High | 382500 | 1912.5 | 382500 | 1912.5 |  |  |
| n40 | 15 | 15 | DFT-s-OFDM QPSK | Low | 461500 | 2307.5 | 461500 | 2307.5 | 36@18 | N/A |
|  |  |  |  | Mid | 470000 | 2350 | 470000 | 2350 |  |  |
|  |  |  |  | High | 478500 | 2392.5 | 478500 | 2392.5 |  |  |
| n41 | 15 | 15 | DFT-s-OFDM QPSK | Low | 500700 | 2503.5 | 500700 | 2503.5 | 36@18 | N/A |
|  |  |  |  | Mid | 518601 | 2593.005 | 518601 | 2593.005 |  |  |
|  |  |  |  | High | 536499 | 2682.495 | 536499 | 2682.495 |  |  |
| n48 | 15 | 15 | DFT-s-OFDM QPSK | Low | 637168 | 3557.52 | 637168 | 3557.52 | 36@18 | N/A |
|  |  |  |  | Mid | 641666 | 3624.99 | 641666 | 3624.99 |  |  |
|  |  |  |  | High | 646166 | 3692.49 | 646166 | 3692.49 |  |  |
| n50 | 15 | 15 | DFT-s-OFDM QPSK | Low | 287900 | 1439.5 | 287900 | 1439.5 | 36@18 | N/A |
|  |  |  |  | Mid | 294900 | 1474.5 | 294900 | 1474.5 |  |  |
|  |  |  |  | High | 301900 | 1509.5 | 301900 | 1509.5 |  |  |
| n51 | 5 | 15 | DFT-s-OFDM QPSK | Low | 285900 | 1429.5 | 285900 | 1429.5 | 12@6 | N/A |
|  |  |  |  | Mid |  |  |  |  |  |  |
|  |  |  |  | High |  |  |  |  |  |  |
| n53 | 10 | 15 | DFT-s-OFDM QPSK | Low | 497700 | 2488.5 | 497700 | 2488.5 | 25@12 | N/A |
|  |  |  |  | Mid | 497860 | 2489.3 | 497860 | 2489.3 |  |  |
|  |  |  |  | High | 498000 | 2490 | 498000 | 2490 |  |  |
| n65 | 15 | 15 | DFT-s-OFDM QPSK | Low | 423500 | 2117.5 | 423500 | 2117.5 | 36@18 | N/A |
|  |  |  |  | Mid | 431000 | 2155 | 431000 | 2155 |  |  |
|  |  |  |  | High | 438500 | 2192.5 | 438500 | 2192.5 |  |  |
| n71 | 10 | 15 | DFT-s-OFDM QPSK | Low | 133600 | 668 | 124400 | 622 | 25@12 | N/A |
|  |  |  |  | Mid | 136100 | 680.5 | 126900 | 634.5 |  |  |
|  |  |  |  | High | 138600 | 693 | 129400 | 647 |  |  |
| n74 | 15 | 15 | DFT-s-OFDM QPSK | Low | 286900 | 1434.5 | 296500 | 1482.5 | 36@18 | N/A |
|  |  |  |  | Mid | 289700 | 1448.5 | 299300 | 1496.5 |  |  |
|  |  |  |  | High | 292500 | 1462.5 | 302100 | 1510.5 |  |  |
| n77 | 15 | 15 | DFT-s-OFDM QPSK | Low | 620500 | 3307.5 | 620500 | 3307.5 | 36@18 | N/A |
|  |  |  |  | Mid | 650000 | 3750 | 650000 | 3750 |  |  |
|  |  |  |  | High | 679500 | 4192.5 | 679500 | 4192.5 |  |  |
| n78 | 15 | 15 | DFT-s-OFDM QPSK | Low | 620500 | 3307.5 | 620500 | 3307.5 | 36@18 | N/A |
|  |  |  |  | Mid | 636666 | 3549.99 | 636666 | 3549.99 |  |  |
|  |  |  |  | High | 652832 | 3792.48 | 652832 | 3792.48 |  |  |
| n79 | 20 | 15 | DFT-s-OFDM QPSK | Low | 694000 | 4410 | 694000 | 4410 | 50@25 | N/A |
|  |  |  |  | Mid | 713333 | 4699.995 | 713333 | 4699.995 |  |  |
|  |  |  |  | High | 732667 | 4990.005 | 732667 | 4990.005 |  |  |

Table 4.3.3-4: NR FR1 TRS measurement parameters for RedCap UE

| NR Band | CBW (MHz) | SCS (kHz) | DL modulation | UL modulation | Range | UL Carrier centre [ARFCN] | UL Carrier Center (MHz) | DL Carrier centre [ARFCN] | DL Carrier Center (MHz) | UL RB Allocation (LCRB @ RBstart) | DL Configuration (FULL RB, LCRB @ RBstart) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| n1 | 15 | 15 | CP-OFDM QPSK | DFT-s-OFDM QPSK | Low | 385500 | 1927.5 | 423500 | 2117.5 | 75@4 | 79@0 |
|  |  |  |  |  | Mid | 390000 | 1950 | 428000 | 2140 |  |  |
|  |  |  |  |  | High | 394500 | 1972.5 | 432500 | 2162.5 |  |  |
| n2 | 15 | 15 | CP-OFDM QPSK | DFT-s-OFDM QPSK | Low | 371500 | 1857.5 | 387500 | 1937.5 | 50@29 | 79@0 |
|  |  |  |  |  | Mid | 376000 | 1880 | 392000 | 1960 |  |  |
|  |  |  |  |  | High | 380500 | 1902.5 | 396500 | 1982.5 |  |  |
| n3 | 15 | 15 | CP-OFDM QPSK | DFT-s-OFDM QPSK | Low | 343500 | 1717.5 | 362500 | 1812.5 | 50@29 | 79@0 |
|  |  |  |  |  | Mid | 349500 | 1747.5 | 368500 | 1842.5 |  |  |
|  |  |  |  |  | High | 355500 | 1777.5 | 374500 | 1872.5 |  |  |
| n5 | 15 | 15 | CP-OFDM QPSK | DFT-s-OFDM QPSK | Low | 166300 | 831.5 | 175300 | 876.5 | 25@54 | 79@0 |
|  |  |  |  |  | Mid | 167300 | 836.5 | 176300 | 881.5 |  |  |
|  |  |  |  |  | High | 168300 | 841.5 | 177300 | 886.5 |  |  |
| n7 | 15 | 15 | CP-OFDM QPSK | DFT-s-OFDM QPSK | Low | 501500 | 2507.5 | 525500 | 2627.5 | 75@4 | 79@0 |
|  |  |  |  |  | Mid | 507000 | 2535 | 531000 | 2655 |  |  |
|  |  |  |  |  | High | 512500 | 2562.5 | 536500 | 2682.5 |  |  |
| n8 | 15 | 15 | CP-OFDM QPSK | DFT-s-OFDM QPSK | Low | 177500 | 887.5 | 186500 | 932.5 | 25@54 | 79@0 |
|  |  |  |  |  | Mid | 179500 | 897.5 | 188500 | 942.5 |  |  |
|  |  |  |  |  | High | 181500 | 907.5 | 190500 | 952.5 |  |  |
| n12 | 10 | 15 | CP-OFDM QPSK | DFT-s-OFDM QPSK | Low | 140800 | 704 | 146800 | 734 | 20@32 | 52@0 |
|  |  |  |  |  | Mid | 141500 | 707.5 | 147500 | 737.5 |  |  |
|  |  |  |  |  | High | 142200 | 711 | 148200 | 741 |  |  |
| n14 | 10 | 15 | CP-OFDM QPSK | DFT-s-OFDM QPSK | Low | 158600 | 793 | 152600 | 763 | 20@32 | 52@0 |
|  |  |  |  |  | Mid |  |  |  |  |  |  |
|  |  |  |  |  | High |  |  |  |  |  |  |
| n20 | 15 | 15 | CP-OFDM QPSK | DFT-s-OFDM QPSK | Low | 167900 | 839.5 | 159700 | 798.5 | 20@11 | 79@0 |
|  |  |  |  |  | Mid | 169400 | 847 | 161200 | 806 |  |  |
|  |  |  |  |  | High | 170900 | 854.5 | 162700 | 813.5 |  |  |
| n25 | 15 | 15 | CP-OFDM QPSK | DFT-s-OFDM QPSK | Low | 371500 | 1857.5 | 387500 | 1937.5 | 50@29 | 79@0 |
|  |  |  |  |  | Mid | 376500 | 1882.5 | 392500 | 1962.5 |  |  |
|  |  |  |  |  | High | 381500 | 1907.5 | 397500 | 1987.5 |  |  |
| n26 | 15 | 15 | CP-OFDM QPSK | DFT-s-OFDM QPSK | Low | 164300 | 821.5 | 173300 | 866.5 | 25@54 | 79@0 |
|  |  |  |  |  | Mid | 166300 | 831.5 | 175300 | 876.5 |  |  |
|  |  |  |  |  | High | 168300 | 841.5 | 177300 | 886.5 |  |  |
| n28 | 15 | 15 | CP-OFDM QPSK | DFT-s-OFDM QPSK | Low | 142100 | 710.5 | 153100 | 765.5 | 25@54 | 79@0 |
|  |  |  |  |  | Mid | 145100 | 725.5 | 156100 | 780.5 |  |  |
|  |  |  |  |  | High | 148100 | 740.5 | 159100 | 795.5 |  |  |
| n30 | 10 | 15 | CP-OFDM QPSK | DFT-s-OFDM QPSK | Low | 462000 | 2310 | 471000 | 2355 | 20@32 | 52@0 |
|  |  |  |  |  | Mid |  |  |  |  |  |  |
|  |  |  |  |  | High |  |  |  |  |  |  |
| n34 | 10 | 15 | CP-OFDM QPSK | DFT-s-OFDM QPSK | Low | 403000 | 2015 | 403000 | 2015 | 50@0 | 52@0 |
|  |  |  |  |  | Mid | 403500 | 2017.5 | 403500 | 2017.5 |  |  |
|  |  |  |  |  | High | 404000 | 2020 | 404000 | 2020 |  |  |
| n38 | 15 | 15 | CP-OFDM QPSK | DFT-s-OFDM QPSK | Low | 515500 | 2577.5 | 515500 | 2577.5 | 75@0 | 79@0 |
|  |  |  |  |  | Mid | 519000 | 2595 | 519000 | 2595 |  |  |
|  |  |  |  |  | High | 522500 | 2612.5 | 522500 | 2612.5 |  |  |
| n39 | 15 | 15 | CP-OFDM QPSK | DFT-s-OFDM QPSK | Low | 377500 | 1887.5 | 377500 | 1887.5 | 75@0 | 79@0 |
|  |  |  |  |  | Mid | 380000 | 1900 | 380000 | 1900 |  |  |
|  |  |  |  |  | High | 382500 | 1912.5 | 382500 | 1912.5 |  |  |
| n40 | 15 | 15 | CP-OFDM QPSK | DFT-s-OFDM QPSK | Low | 461500 | 2307.5 | 461500 | 2307.5 | 75@0 | 79@0 |
|  |  |  |  |  | Mid | 470000 | 2350 | 470000 | 2350 |  |  |
|  |  |  |  |  | High | 478500 | 2392.5 | 478500 | 2392.5 |  |  |
| n41 | 15 | 15 | CP-OFDM QPSK | DFT-s-OFDM QPSK | Low | 500700 | 2503.5 | 500700 | 2503.5 | 75@0 | 79@0 |
|  |  |  |  |  | Mid | 518601 | 2593.005 | 518601 | 2593.005 |  |  |
|  |  |  |  |  | High | 536499 | 2682.495 | 536499 | 2682.495 |  |  |
| n48 | 15 | 15 | CP-OFDM QPSK | DFT-s-OFDM QPSK | Low | 637168 | 3557.52 | 637168 | 3557.52 | 75@0 | 79@0 |
|  |  |  |  |  | Mid | 641666 | 3624.99 | 641666 | 3624.99 |  |  |
|  |  |  |  |  | High | 646166 | 3692.49 | 646166 | 3692.49 |  |  |
| n50 | 15 | 15 | CP-OFDM QPSK | DFT-s-OFDM QPSK | Low | 287900 | 1439.5 | 287900 | 1439.5 | 75@0 | 79@0 |
|  |  |  |  |  | Mid | 294900 | 1474.5 | 294900 | 1474.5 |  |  |
|  |  |  |  |  | High | 301900 | 1509.5 | 301900 | 1509.5 |  |  |
| n51 | 5 | 15 | CP-OFDM QPSK | DFT-s-OFDM QPSK | Low | 285900 | 1429.5 | 285900 | 1429.5 | 25@0 | 25@0 |
|  |  |  |  |  | Mid |  |  |  |  |  |  |
|  |  |  |  |  | High |  |  |  |  |  |  |
| n53 | 10 | 15 | CP-OFDM QPSK | DFT-s-OFDM QPSK | Low | 497700 | 2488.5 | 497700 | 2488.5 | 50@0 | 52@0 |
|  |  |  |  |  | Mid | 497860 | 2489.3 | 497860 | 2489.3 |  |  |
|  |  |  |  |  | High | 498000 | 2490 | 498000 | 2490 |  |  |
| n65 | 15 | 15 | CP-OFDM QPSK | DFT-s-OFDM QPSK | Low | 423500 | 2117.5 | 423500 | 2117.5 | 75@4 | 79@0 |
|  |  |  |  |  | Mid | 431000 | 2155 | 431000 | 2155 |  |  |
|  |  |  |  |  | High | 438500 | 2192.5 | 438500 | 2192.5 |  |  |
| n70 | 15 | 15 | CP-OFDM QPSK | DFT-s-OFDM QPSK | Low | 340500 | 1702.5 | 400500 | 2002.5 | 75@4 | 79@0 |
|  |  |  |  |  | Mid |  |  |  |  |  |  |
|  |  |  |  |  | High |  |  |  |  |  |  |
| n71 | 10 | 15 | CP-OFDM QPSK | DFT-s-OFDM QPSK | Low | 133600 | 668 | 124400 | 622 | 25@0 | 52@0 |
|  |  |  |  |  | Mid | 136100 | 680.5 | 126900 | 634.5 |  |  |
|  |  |  |  |  | High | 138600 | 693 | 129400 | 647 |  |  |
| n74 | 15 | 15 | CP-OFDM QPSK | DFT-s-OFDM QPSK | Low | 286900 | 1434.5 | 296500 | 1482.5 | 25@54 | 79@0 |
|  |  |  |  |  | Mid | 289700 | 1448.5 | 299300 | 1496.5 |  |  |
|  |  |  |  |  | High | 292500 | 1462.5 | 302100 | 1510.5 |  |  |
| n77 | 15 | 15 | CP-OFDM QPSK | DFT-s-OFDM QPSK | Low | 620500 | 3307.5 | 620500 | 3307.5 | 75@0 | 79@0 |
|  |  |  |  |  | Mid | 650000 | 3750 | 650000 | 3750 |  |  |
|  |  |  |  |  | High | 679500 | 4192.5 | 679500 | 4192.5 |  |  |
| n78 | 15 | 15 | CP-OFDM QPSK | DFT-s-OFDM QPSK | Low | 620500 | 3307.5 | 620500 | 3307.5 | 75@0 | 79@0 |
|  |  |  |  |  | Mid | 636666 | 3549.99 | 636666 | 3549.99 |  |  |
|  |  |  |  |  | High | 652832 | 3792.48 | 652832 | 3792.48 |  |  |
| n79 | 20 | 15 | CP-OFDM QPSK | DFT-s-OFDM QPSK | Low | 694000 | 4410 | 694000 | 4410 | 100@0 | 106@0 |
|  |  |  |  |  | Mid | 713333 | 4699.995 | 713333 | 4699.995 |  |  |
|  |  |  |  |  | High | 732667 | 4990.005 | 732667 | 4990.005 |  |  |

### 4.3.4	EN-DC band combinations

<Editor’s note: Example EN-DC combinations can be further added. >

Principle of EN-DC band combinations selection for FR1 TRP TRS OTA testing:

1)	Focus on the performance of the NR carrier and do not consider multiple permutations between different LTE bands and NR band under test, i.e., for each NR band, only select one EN-DC band combination.

2)	For UE supporting multiple EN-DC band combinations for the same NR band, consider only those EN-DC configurations which have no MSD impact on either LTE or NR, i.e., the selected EN-DC combination should be no MSD issue identified in TS 38.101-3 Section 7.3B.2.3 (Inter-band EN-DC within FR1).

Table 4.3.4-1: Measurement parameters for example inter-band EN-DC band combinations (two bands)

| EN-DC configuration | E-UTRA configurations | NR configurations |
| --- | --- | --- |
| DC_3A_n28A | Note1 | Note2 |
| DC_2A_n41A | Note1 | Note2 |
| DC_1A_n78A | Note1 | Note2 |
| DC_1A_n79A | Note1 | Note2 |
| Note 1: As per TS 37.544 [14], Clause 5.3 and 5.4 (Measurement frequencies for E-UTRA FDD and TDD). Note 2: As per Table 4.3.3-1 and Table 4.3.3-2 in this specification. The measurement parameters for NR Low Mid High ranges correspond to E-UTRA Low Mid High ranges respectively. |  |  |

With the above basic principle and EN-DC example band combination, the selection logic for testing is defined by the decision tree below.

Figure 4.3.4-1: Decision tree to select the EN-DC band combination for TRP/TRS testing

### 4.3.5	CA band combinations

The principle of CA combination selection for FR1 TRP TRS OTA testing:

1.	Limit CA combinations to the bands n1, n3, n5, n7, n8, n28, n77 and n79.

2.	The selected CA combinations should not have MSD issue identified in TS 38.101-1 Section 7.3A.

The CA combinations for inter-band DL CA (two bands) with single UL are given in Table 4.3.5-1.

Table 4.3.5-1: Measurement parameters for example inter-band DL CA band combinations (two bands)

| CA configuration | Uplink configurations | NR configurations |
| --- | --- | --- |
| CA_n1A-n79A | n1 | Note 1 |
| CA_n3A-n28A | n28 | Note 1 |
| CA_n3A-n79A | n3 | Note 1 |
| CA_n7A-n28A | n28 | Note 1 |
| CA_n7A-n77A | n7 | Note 1 |
| CA_n28A-n79A | n28 | Note 1 |
| Note 1: As per Table 4.3.3-1 and Table 4.3.3-2 in this specification, NR Low Mid High ranges in one component carrier correspond to those in the other component carrier. |  |  |

The CA band combinations provided in this clause do not imply correspondence to any TRP TRS requirement.

### 4.3.6	Alternative parameters related to channel bandwidth and bands

#### 4.3.6.1	General

This clause defines the alternative parameters related to channel bandwidths and bands and is applicable to bands n28, n41, n77, and n78 with the following applicability:

-	For a given band, Power Class, and test mode, if a UE fulfils the TRS requirements according to the test parameters defined in Clause 4.3.3, then it is considered to have fulfilled the TRS requirements according to the parameters defined in Clause 4.3.6, where applicable; or vice versa

-	For a given band, Power Class, and test mode, if a UE fulfils the TRS requirements according to the test parameters defined in Clause 4.3.6, then it is considered to have fulfilled the TRS requirements according to the parameters defined in Clause 4.3.3, where applicable; or vice versa.

#### 4.3.6.2	Test parameters

Table 4.3.6.2-1 below defines the alternative NR FR1 TRS measurement parameters.

Table 4.3.6.2-1: Alternative NR FR1 TRS measurement parameters

| NR Band | CBW (MHz) | SCS (kHz) | DL modulation | UL modulation | Range | UL Carrier centre [ARFCN] | UL Carrier Center (MHz) | DL Carrier centre [ARFCN] | DL Carrier Center (MHz) | UL RB Allocation (LCRB @ RBstart) | DL Configuration (FULL RB, LCRB @ RBstart) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| n28 | 10 | 15 | CP-OFDM QPSK | DFT-s-OFDM QPSK | Low | 141600 | 708 | 152600 | 763 | 25@27 | 52@0 |
|  |  |  |  |  | Mid | 145100 | 725.5 | 156100 | 780.5 |  |  |
|  |  |  |  |  | High | 148600 | 743 | 159600 | 798 |  |  |
| n41 | 20 | 30 | CP-OFDM QPSK | DFT-s-OFDM QPSK | Low | 501204 | 2506.02 | 501204 | 2506.02 | 50@0 | 106@0 |
|  |  |  |  |  | Mid | 518598 | 2592.99 | 518598 | 2592.99 |  |  |
|  |  |  |  |  | High | 535998 | 2679.99 | 535998 | 2679.99 |  |  |
| n77 | 20 | 30 | CP-OFDM QPSK | DFT-s-OFDM QPSK | Low | 620668 | 3310.02 | 620668 | 3310.02 | 50@0 | 106@0 |
|  |  |  |  |  | Mid | 650000 | 3750 | 650000 | 3750 |  |  |
|  |  |  |  |  | High | 679332 | 4189.98 | 679332 | 4189.98 |  |  |
| n78 | 20 | 30 | CP-OFDM QPSK | DFT-s-OFDM QPSK | Low | 620668 | 3310.02 | 620668 | 3310.02 | 50@0 | 106@0 |
|  |  |  |  |  | Mid | 636666 | 3549.99 | 636666 | 3549.99 |  |  |
|  |  |  |  |  | High | 652666 | 3789.99 | 652666 | 3789.99 |  |  |

#### 4.3.6.3	TRS requirements for additional channel bandwidths

The below framework shall be used for deriving TRS requirements for additional channel bandwidths.

-	The core specification shall only list the requirement based on the band parameters defined in Clasue 4.3.3

-	Scaled requirements for additional channel bandwidths is limited to bands n28 (additional channel bandwidth 10 MHz), n41, n77, n78 (additional channel bandwidth 20 MHz).

The ratio of the channel bandwidths shall be used to derive and scale requirements on alternate CBW (20 MHz) for n41, n77 and n78 using the below formula.

Average_TRS20MHz  (dBm) = Average_TRS100MHz (dBm) - 7 dB

The ratio of channel bandwidths shall be used to derive and scale requirements on alternate CBW (10 MHz) for n28, using the below formula.

Average_TRS10MHz  (dBm) = Average_TRS20MHz (dBm) - 3 dB

# 5	Performance metrics

## 5.1	Definition of the Total Radiated Power (TRP)

### 5.1.1	Definition of the Total Radiated Power (TRP) for AC

#### 5.1.1.1	Applicability of TRP metrics

The TRP metrics are defined in Clauses 5.1.1.2 through 5.1.1.3, with the applicability provided in Table 5.1.1.1-1 below.

Table 5.1.1.1-1: Applicability of TRP metrics

| Clause | TRP definition | Applicability |
| --- | --- | --- |
| 5.1.1.2 | TRP=14πθ=0πϕ=02πEIRPθθ,ϕ+EIRPϕθ,ϕsinθdϕdθ | TRP for UEs supporting single-port transmission |
| 5.1.1.3 | TRP= 14πθ=0πϕ=02πEIRPθ,TPMI_2θ,ϕ+EIRPϕ,TPMI_2θ,ϕsinθdϕdθ | TRP for UEs supporting non-coherent UL MIMO |
| 5.1.1.4 | TRPenv= 14πθ=0πϕ=02πEIRPmaxθ,ϕsinθdϕdθ | Both TRPenv and TRPavg are applicable for UEs supporting coherent UL MIMO |
|  | TRPavg=14πθ=0πϕ=02πEIRPavgθ,ϕsinθdϕdθ |  |

#### 5.1.1.2	TRP for UEs supporting single-port transmission

Transmitter power measurements shall be performed using the Total Radiated Power (TRP) as the measurement metric. This clause defines the definition of TRP value of NR FR1 DUT for Anechoic Chamber (AC) method.

The TRP with Anechoic Chamber method is defined as:

TRP= 14πθ=0πϕ=02πEIRPθθ,ϕ+EIRPϕθ,ϕsinθdϕdθ		(5.1)

Where the effective isotropic radiated power (EIRP) is defined as

EIRPθ,ϕ=PTGTθ,ϕ              	  (5.2)

Where PTGT is the product of the power delivered to the antenna and the antenna’s power gain.

Where EIRPθ and EIRPϕ are the EIRP in the corresponding θ and ϕ polarizations.

The summation form based on the sin weights of TRP with Anechoic Chamber method is defined as:

TRP≈ π2NMn=0N-1m=0M-1EIRPθθn,ϕm+EIRPϕθn,ϕmsinθn      	      (5.3)

Where N and M are the number of sampling intervals for θ and ϕ. θn and ϕm are the measurement angles.

The summation form based on the Clenshaw-Curtis quadrature integral approximation of TRP with Anechoic Chamber method is defined as:

TRP≈ 12Mn=0Nm=0M-1EIRPθθn,ϕm+EIRPϕθn,ϕmWθn            	(5.4)

Where the value of W(θn) can be calculated as follows:

W(θn)=ciN1-j=1int(N2)bj4j2-1cos⁡(2jθn)	(5.5)

with

bj=    1,  2j=N       2,  otherwise

and

ci=   1,  i=0 or N   2,  otherwise

The applicability of TRP quadratures, frequency ranges, and measurement grids is tabulated in Table 5.1.1-1.

Table 5.1.1-1: Applicability for TRP measurement grids

| Frequency Range | Quadrature | [°] | N | M | Min. Number of Grid Points |
| --- | --- | --- | --- | --- | --- |
| < 3GHz | sin() | 15 | 12 | 24 | 264 |
|  |  | 30 | 6 | 12 | 60 |
|  | Clenshaw-Curtis | 15 | 12 | 24 | 266 |
|  |  | 30 | 6 | 12 | 62 |
| > 3GHz | sin() | 15 | 12 | 24 | 264 |
|  |  | 30 | 6 | 12 | 60 |
|  | Clenshaw-Curtis | 15 | 12 | 24 | 266 |
|  |  | 30 | 6 | 12 | 62 |

#### 5.1.1.3	TRP for UEs supporting non-coherent UL MIMO

For devices capable of non-coherent UL MIMO and intended to be verified under the fixed TPMI test condition, the TRP metric defined in 5.1.1.2 is enhanced as follows:

Where the effective isotropic radiated power (EIRP) is measured by configuring the UE with the transmit precoding matrix index (TPMI) set to 2:

EIRPTPMI_2θ,ϕ=PTGT,TPMI_2θ,ϕ              	  (5.7)

By substituting equation 5.1 in Clause 5.1.1.2 for 5.8 in this Clause, all of the remaining definitions in Clause 5.1.1.2 apply for TRP for non-coherent UL MIMO.

#### 5.1.1.4	TRP for UEs supporting coherent UL MIMO

For devices capable of coherent UL MIMO, the TRP metric defined in 5.1.1.2 is enhanced as follows:

TRPenv= 14πθ=0πϕ=02πEIRPmaxθ,ϕsinθdϕdθ		(5.8)
	TRPavg=14πθ=0πϕ=02πEIRPavgθ,ϕsinθdϕdθ		(5.9)

Where both TRPenv and TRPavg are applicable for UEs supporting coherent UL MIMO.

In Equation 5.8, the effective isotropic radiated power (EIRP) is defined as the outcome of selecting the best EIRP from a sweep of all EIRP per grid point measured by configuring the UE with the transmit precoding matrix indices (TPMI) which are applicable according to the UE reported UL MIMO capability:

EIRPmaxθ,ϕ=maxn∈2,3,4,5EIRPθ,TPMI_nθ,ϕ+EIRPϕ,TPMI_nθ,ϕ              	  (5.10)

<< Editor’s note: MU associated with TRPenv to be investigated based on measurements >>

In Equation 5.9, EIRPavg is defined according to Equations 5.10, 5.11, and 5.12 below.  Equation 5.11 is the baseline definition of EIRPavg.

EIRPavgθ,ϕ=14n=25EIRPθ,TPMI_nθ,ϕ+EIRPϕ,TPMI_nθ,ϕ	(5.11)

EIRPavgθ,ϕ=12n=23EIRPθ,TPMI_nθ,ϕ+EIRPϕ,TPMI_nθ,ϕ	(5.12)

EIRPavgθ,ϕ=12n=45EIRPθ,TPMI_nθ,ϕ+EIRPϕ,TPMI_nθ,ϕ	(5.13)

<< Editor’s note: the practical equivalence between the metrics defined by equations 5.12 and 5.13 with the baseline in equation 5.11 and the MU associated with the metrics defined in equations 5.11, 5.12, and 5.13 to be investigated based on measurements >>

The effective isotropic radiated power (EIRP) of the n-th TPMI is measured by configuring the UE with the corresponding transmit precoding matrix index set to n, where n is selected out of the set of TPMI indices according to equations 5.11, 5.12, and 5.13:

EIRPTPMI_nθ,ϕ=PTGT,TPMI_nθ,ϕ              	  (5.14)

By substituting equation 5.1 in Clause 5.1.1.2 for 5.8 and 5.9 in this Clause, all of the remaining definitions in Clause 5.1.1.2 apply for TRP for coherent UL MIMO.

### 5.1.2	Definition of the Total Radiated Power (TRP) for RC method

TRP in the reverberation chamber is a measurement of transmitter performance in an isotropic Rayleigh fading environment that is based on sampling the radiated power of the UE for a discrete number of field combinations. The average value of these statistically distributed samples is proportional to the TRP and by calibrating the average power transfer function, an absolute value of the TRP can be obtained.  The TRP with Reverberation Chamber method is defined as:

TRP=  PavgC1-R Pref

Where Pref is the reference power transfer function for the fixed measurement antenna, R is the reflection coefficient for the fixed measurement antenna and C is the path loss in the cables connecting the measurement receiver to fixed measurement antenna. These parameters are calculated from the calibration measurement and are further discussed in calibration section. Pavg is the average power measured by the fixed measurement antenna and can be calculated using the following expression:

Pavg=  m=1MS21,.m2M

Where S21,m is sample number m of the complex transfer function measured with the fixed measurement antenna and M is the total number of samples measured.

## 5.2	Definition of Total Radiated Sensitivity (TRS)

### 5.2.1	Definition of the Total Radiated Sensitivity (TRS) for AC

Receiver sensitivity measurements shall be performed using data throughput as the measurement metric. The DUT’s receiver sensitivity corresponds to the minimum downlink signal power required to provide a data throughput rate greater than or equal to 95% of the maximum throughput of the reference measurement channel (RMC).

This definition will be used to calculate the Total Radiated Sensitivity (TRS) value of NR FR1 DUT.

The TRS with Anechoic Chamber method is defined as:

TRS=4πθ=0πϕ=02π1EISθθ,ϕ+1EISϕθ,ϕsinθdϕdθ     	   (5.15)

Where the effective isotropic sensitivity (EIS) is defined as the minimum power level at which the throughput exceeds or equal to 95% of the maximum throughput of the specified RMC, at each sampling point.

Where EISθ and EISϕ are the EIS in the corresponding θ and ϕ polarizations.

The summation form based on the sin weights of TRS with Anechoic Chamber method defined as:

TRS≈2NMπn=0N-1m=0M-11EISθθn,ϕm+1EISϕθn,ϕmsinθn                 	 (5.16)

Where N and M are the number of sampling intervals for θ and ϕ. θn and ϕm are the measurement angles.

The summation form based on the Clenshaw-Curtis quadrature integral approximation of TRS with Anechoic Chamber method is defined as:

TRS≈2Mn=0Nm=0M-11EISθθn,ϕm+1EISϕθn,ϕmWθn                  	(5.17)

Where the value of W(θn) follows Equation 5.5.

The applicability of TRS quadratures, frequency ranges, and measurement grids is tabulated in Table 5.2.1-1.

Table 5.2.1-1: Applicability for TRS measurement grids

| Frequency Range | Quadrature | [°] | N | M | Min. Number of Grid Points |
| --- | --- | --- | --- | --- | --- |
| < 3GHz | sin() | 30 | 6 | 12 | 60 |
|  | Clenshaw-Curtis | 30 | 6 | 12 | 62 |
|  |  | 45 | 4 | 8 | 26 |
| > 3GHz | sin() | 30 | 6 | 12 | 60 |
|  | Clenshaw-Curtis | 30 | 6 | 12 | 62 |
|  |  | 45 | 4 | 8 | 26 (Note 1) |
| Note 1: When the back pole at = 180° cannot be measured due to obstruction and/or blocking, extrapolation is used to estimate EIS at = 180° for measurement grids with =45° by either a) using at least two points within 15° of the pole or b) averaging the last cut (i.e. = 135°) |  |  |  |  |  |

### 5.2.2	Definition of the Total Radiated Sensitivity (TRS) for RC method

The calculation of TRS is based on searching for the lowest power received by the UE for a discrete number of field combinations in the chamber. The downlink power received by the UE at each discrete field combination that provides a  throughput which is better than the specified target  throughput level shall be averaged with other such measurements using different field combinations. By calibrating the average power transfer function, an absolute value of the TRS can be obtained when the linear values of all downlink power levels have been averaged.

The TRS with Reverberation Chamber method is defined as:

TRS=(Pref 1-R C)1M m=1M1PBSSm-1

Where Pref is the reference power transfer function for the fixed measurement antenna, R is the reflection coefficient for the fixed measurement antenna and C is the path loss in the cables connecting the measurement receiver to fixed measurement antenna. The parameters are calculated from the calibration measurement and are further discussed in calibration section. PBSSm is the output power from the base station emulator when it is adjusted to give the specified throughput from the DUT for mode-stirring sample m.

# 6	UE positioning guidelines

## 6.1	Free space

For Free space configuration, the centre of the reference coordinate system shall be aligned with the geometric centre of the DUT in order to minimize the offset between antenna arrays integrated at any position of the UE and the centre of the quiet zone.

Table 6.1-1: UE positioning for Free space

| Test condition | DUT orientation | Diagram |
| --- | --- | --- |
| Free space DUT | α = 0º; β = 0º; γ = 0º |  |

## 6.2	Hand phantom only (Browsing mode)

The positioning specified in this clause is used for the test cases for Browsing Mode with Hand Phantom. The characteristics of the Hand Phantom are specified in Annex D. Browsing mode is used to simulate user cases where the DUT is held in hand, but not pressed against ear e.g. web browsing and navigation. For hand phantom only, the DUT shall be mounted in a suitable hand phantom and oriented such that the DUT’s main display is tilted 45 degrees from vertical:

-	Wide Grip Hand for UE with Width >72mm and ≤92mm

-	PDA Grip Hand for UE with Width ≥56mm and ≤72mm

### 6.2.1	Wide Grip Hand

This positioning guideline is suitable for DUTs with width >72mm and ≤92mm.

The positioning guideline for mounting a DUT in the Wide Grip Hand Phantom defined in CTIA Certification OTA Test Plan 01.71 section 2.2.7 [14], is used for FR1 TRP TRS testing for UE with Width >72mm and ≤92mm in this technical report.

Figure 6.2.1-1: Positioning guidance for Wide Grip Hand (© 2001 – 2022 CTIA Certification. Reproduced with permission.), defined in the CTIA Certification OTA Test Plan

### 6.2.2	PDA Grip Hand

This clause defines the positioning guideline for mounting a DUT in the PDA Grip Hand Phantom. This positioning guideline is suitable for DUTs with width ≥56mm and ≤72mm.

To help achieve a consistent positioning, the DUT is aligned to a PDA palm spacer. No alignment tool is required. The PDA spacer features side and bottom walls to ensure consistent alignment of DUTs of various sizes.

1.	Place the DUT on the PDA spacer between the fingers and align the DUT to the side wall of the PDA.

2.	If the DUT is shorter than 135 mm, then align the top of the DUT with the top of the PDA spacer. Otherwise, align the bottom of the DUT with the bottom wall of the PDA spacer.

Figure 6.2.2-1: Right-handed PDA Grip hand phantom with a spacer

NOTE: Use left-handed (mirror-imaged) spacers with left-handed phantoms.

## 6.3	Head and Hand phantom (Talk Mode)

### 6.3.1	General

The positioning specified in this clause is used for the test cases for Talk Mode using Head & Hand Phantom. The characteristics of the Hand Phantom and Head Phantom are specified in Annex D.

Talk mode is used to simulate user cases where the DUT is placed into a hand phantom, which is holding the DUT against the SAM head phantom, presenting a realistic voice call operation of the DUT. Same as Browsing mode, the DUT for talk mode shall also be mounted in a suitable hand phantom:

-	Wide Grip Hand for UE with Width >72mm and ≤92mm

-	PDA Grip Hand for UE with Width ≥56mm and ≤72mm

In this section, the procedure provides the guideline on how to place the DUT+hand against the head phantom. The detailed DUT positioning into hand phantom for talk mode is defined in section 6.3.2 and 6.3.3.

For talk mode, the DUT is attached to the head phantom in “cheek” position. The DUT performance is measured on both left and right side of the head. Three points as shown in Figure 6.3.1-1 define the reference plane: centre of the right ear piece (RE), centre of the left ear piece (LE) and centre of mouth (M).

Definition of the ‘Cheek’ position:

1.	Align the ear piece of the phone (see Figure 6.3.1-1) at the line RE-LE. Then, position the DUT beside the phantom so that the vertical line (see Figure 6.3.1-3) is parallel to the reference plane in Figure 6.3.1-2 and is aligned with the line M-RE on the reference plane (see Figure 6.3.1-3).

2.	Position the DUT so that the ear piece of the DUT touches the ear piece of the phantom head on the line RE-LE. Tilt the DUT chassis towards the cheek of the phantom having the vertical line aligned with the reference plane until any point on the front side of the DUT is in contact with the cheek or until the contact with the ear is lost.

NOTE:	A holder fixture made of e.g. plastic may be used to position the handset against the phantom.

Figure 6.3.1-1: Reference plane on head phantom, front view

Figure 6.3.1-2: Reference plane on head phantom, side view

Figure 6.3.1-3: Reference lines at a mobile handset.

In addition, 6° tilt angle from the cheek is being used instead of having direct contact between the cheek and DUT. A mask may be used to help configuration of cheek + 6° tilt angle. The mask is a 32 mm wide conformal strip, created by sweeping the surface of the head phantom through a 6º rotation about the ear. Direct DUT contact against the mask thus establishes the required 6º spacing away from the cheek, regardless of DUT form factor. The material for the head phantom mask spacer shall have dielectric constant of less than 1.3 and a loss tangent of less than 0.003. Material additions can be used to help fixing of the mask spacer onto the head phantom.

In summary, for head + hand phantom, keeping the DUT in the hand phantom in the position defined in clause 6.2, while place the DUT and the hand phantom against the head phantom in such way that the DUT is in 6° tilt angle as described above.

### 6.3.2	Wide Grip Hand and Head

This procedure is suitable for talk mode use with DUTs of width >72mm and ≤92mm. The positioning of the DUT in the Wide Grip hand for talk mode is identical to that for browsing mode described in section 6.2.1.

### 6.3.3	PDA Grip Hand and Head

This procedure is suitable for talk mode use with DUTs of width ≥56mm and ≤72mm. The positioning of the DUT in the PDA Grip hand for talk mode is identical to that for browsing mode described in section 6.2.2.

## 6.4	Head phantom only

The positioning specified in this clause is used for the test cases for test cases using Head Phantom only. The characteristics of the Head Phantom is specified in Annex D.1.

Positioning of the DUT against the head only is analogous to the positioning of the DUT for the head+hand (talk mode) configuration of section 6.3.1 with the exception that direct contact with the cheek of the head phantom is used instead of the 6º tilt angle from the cheek. The same coordinate system and reference points previously defined on the head phantom and DUT in section 6.3.1 are used.

## 6.5	Forearm phantom

### 6.5.1	Forearm Phantom placement in the chamber

The Forearm Phantom shall be mounted in the chamber coordinate system as shown in Figure 6.5.1-1.

Figure 6.5.1-1: Positioning guidance for Forearm Phantom in the chamber (© 2001 – 2022 CTIA Certification. Reproduced with permission.), defined in the CTIA Certification OTA Test Plan

### 6.5.2	Wrist-Worn RedCap Device mounted on the Forearm Phantom

This positioning guideline is suitable for wrist-worn RedCap devices.

The positioning guideline defined in CTIA Certification OTA Test Plan 01.71 section 2.3 [14], is used for FR1 TRP TRS testing for wrist-worn RedCap devices in this technical report.

The Wrist-worn Device should be fully aligned under some specific virtual Plane to make sure the Wrist-worn Device is properly mounted on the Forearm Phantom.  Plane J cuts through the surface of the forearm phantom and passes through the target test position and is perpendicular to the Y-axis. Plane J is the X-Z plane. Plane K cuts through the forearm phantom at the target test position and is perpendicular to the Z-axis.  Plane K is the X-Y plane. Plane A and Plane B are virtual planes on wrist-worn device. Plane A cuts through the center of wrist band and Plane B cuts through the center of the display.

Plane B shall be fully aligned with Plane J when the device is mounted on the forearm phantom.

Figure 6.5.2-1: Positioning guidance for Wrist-Worn Device mounted on the Forearm Phantom (© 2001 – 2022 CTIA Certification. Reproduced with permission.), defined in the CTIA Certification OTA Test Plan

DUTs shall be mounted with sufficiently snug band tightness so as to prevent the DUT from slipping off under the force of gravity when the phantom is inverted.

Similar to handheld UEs, there are also two Orientations for Wrist-worn RedCap Devices representing the Left and Right Wrist, as shown in Figure 6.5.2-2.

Figure 6.5.2-2: Left and Right positioning for Wrist-Worn Device on Forearm Phantom (© 2001 – 2022 CTIA Certification. Reproduced with permission.), defined in the CTIA Certification OTA Test Plan

# 7	Anechoic Chamber method (Reference method)

## 7.1	General

This clause defines the test method with Anechoic Chamber system, which is agreed as a reference method for FR1 TRP and TRS testing.

## 7.2	Test setup

For FR1 TRP TRS testing, both Single-antenna and multiple-antennas anechoic chambers can be applied. In Figure 7.2-1, an example TRP TRS test system with combined axes system is presented.

Figure 7.2-1: Example of a FR1 TRP TRS OTA test system with combined axis

In Figure 7.2-2, an example TRP TRS test system with distributed axes system is presented.

Figure 7.2-2: Example of a FR1 TRP TRS OTA test system with distributed axis

## 7.3	Calibration procedure

The relative power values of the measurement points will be transformed to absolute radiated power values (in dBm) by performing a range path loss calibration measurement. The system needs to be calibrated by using a reference calibration antenna with known gain values. In the range path loss calibration measurement, the reference antenna is measured in the same place as the DUT, i.e. the center of the QZ, and the attenuation of the complete transmission path (Ltotal) from the DUT to the measurement receiver/BS simulator is calibrated out.

Figure 7.3-1: Example FR1 TRP TRS calibration setup

The calibration measurement is repeated for each measurement path (two orthogonal polarizations and each signal path). The range path loss calibration measurement is performed in a two-step process including total path loss measurement and cable calibration.

Step 1: Cable calibration: the measurement of path loss LDE, by connecting the cable from D to E to the two ports of VNA, and measure the cable path loss.

Step 2: Total path loss measurement: the measurement of total path loss LBC;

1.	Place the reference calibration antenna (e.g. reference dipole) in the center of the test zone aligned with θ polarization of the measurement antenna, connected to a VNA port E, with the other VNA port C connected to the input of the Switch box – in Figure 7.3-1.

2.	Configure the proper output power of VNA.

3.	Measure the response LCE of each path from each θ polarization of the measurement antenna to the reference antenna in the center of QZ.

4.	Repeat the steps 1 to 3 with the reference antenna aligned with the ϕ polarization of the measurement antenna.

Then, the Ltotal=LCE-LDE+Gcal, Where LDE is cable loss from D to E. Gcal is the gain or efficiency of the calibration antenna at the frequency of interest. In TRP and TRS measurements point C is connected to the calibrated input/output port of measurement receiver.

This range path loss calibration procedure is common to both SA and EN-DC measurements.

## 7.4	TRP Test procedure

<Editor’s note: includes test procedure for both SA and EN-DC. Consider NR 1Tx, 2Tx, RedCap, and CA,>

### 7.4.1	General

For TRP and TRS testing in SA or EN-DC mode, measurements should be only performed at NR carrier. The LTE link antenna in EN-DC mode is used to provide a stable LTE link to the DUT without precise path loss or polarization control.

The TRP of the DUT is measured by sampling the radiated transmit power of the DUT with three-dimensional scan at various locations surrounding the device. The measurement is performed with the sampling step for theta () and phi () axes defined in Table B.2.12-1, Annex B.2.12. For some test system that can not measure 180º EIRP, then the extrapolation approach can be adopted when generating the 3D antenna pattern. All of the measured power values will be integrated to TRP, as defined in Clause 5.1.1.2.

For TRP measurement, the evaluations shall be performed at maximum transmit power.

### 7.4.2	TRP test procedure for NR 1Tx configuration

#### 7.4.2.1	UE configuration

For devices containing multiple Tx antennas, the Tx Antenna Switching (TAS) function should be OFF, and the TRP should be measured for each Tx antenna individually. The antenna with better TRP is identified as the primary antenna, and the corresponding TRP result will be used to determine the pass/fail compliance. Otherwise, the primary antenna should be selected based on manufacturer declaration. To ensure the TAS OFF testing, the manufacture should provide either software/guidance to lab to control which Tx antenna is used, or the pre-configured DUT locked at primary antenna.

For Standalone, the NR System Simulator (SS) and DUT shall be configured per TS 38.521-1 [5], section 6.2.1 (UE maximum output power) using the default settings specified in TS 38.521-1 [5] and TS 38.508-1 [7] as applicable. The measurement should be carried out based on the detailed test parameters for each band, as defined in Clause 4.3.

For EN-DC, the SS and DUT shall be configured per TS 38.521-3 [6], Section 6.2B.1 (UE Maximum Output Power for EN-DC) using the default settings specified in TS 38.521-3 [6] and TS 38.508-1 [7] as applicable. The measurement should be carried out based on the detailed test parameters for each band, as defined in Clause 4.3. The UL output power of LTE carrier should be set as a constant power of 10dBm, while measuring NR at maximum output power, i.e., with fixed p-MaxEUTRA-r15=10 dBm, and p-NR-FR1 not configured.

#### 7.4.2.2	Test procedure

<Editor’s note: in general, RedCap test procedure is the same except for forearm positioning.>

For UE configured with 1Tx for NR carrier, SA or EN-DC mode, the measurement procedure includes the following steps:

1)	Place the DUT inside the QZ following the positioning guideline defined in Clause 6.

2)	Connect the SS with the DUT through the link antenna following steps 1 and 2 in section 6.2.1.4.2 of TS 38.521-1 [5] and ensure the DUT transmits with its maximum power.

3)	Measure the power at each measurement point, and calculate EIRPθ,ϕ by adding the composite loss of the entire transmission path.

The TRP value is calculated using the TRP integration approaches outlined in Clause 5.1.1.2.

### 7.4.3	TRP test procedure for NR 2Tx configuration

<Editor’s note: this clause defines test procedure for single-layer UL MIMO, TxD >

#### 7.4.3.1	UE configuration

In general, the UE 2Tx configuration can be categorized into two test cases, i.e., TxD and UL-MIMO.

For TxD test case, the TRP should be measured with all the Tx antenna ON. For SA, the NR System Simulator (SS) and DUT shall be configured per TS 38.521-1 [5], section 6.2G.1 (UE maximum output power for Tx Diversity) using the default settings specified in TS 38.521-1 [5] and TS 38.508-1 [7] as applicable. The measurement should be carried out based on the detailed test parameters for each band, as defined in Clause 4.3 of this TR.

FFS phase issue between 2Tx antennas under TxD configuration, e.g., with additional UE-specific configuration based on UE declaration.

For single-layer UL-MIMO, the baseline TRP measurement should be performed with all the Tx antenna ON. The detailed TPMI configuration for each UE type is as follows:

-	For non-coherent UE support fullpowerMode1, fixed TPMI index =2 as baseline configuration

-	For non-coherent UE does not support fullpowerMode1, single-layer UL-MIMO TRP testing is not required.

-	For coherent UE, two or four TPMI index from TPMI index =2~5 based on UE declaration.

#### 7.4.3.2	TxD TRP Test procedure

<Editor’s note: the study to mitigate or consider the phase between antennas unresolved issues is FFS >

For UE configured with 2Tx for TxD mode, the measurement procedure of TRP includes the following steps:

1)	Place the DUT inside the QZ following the positioning guideline defined in Clause 6.

2)	Connect the SS with the DUT through the link antenna following steps 1 and 2 in section 6.2G.1.4.2 of TS 38.521-1 [5] and ensure the DUT transmits with its maximum power.

3)	Measure the power at each measurement point, and calculate EIRPθ,ϕ by adding the composite loss of the entire transmission path.

The TRP value is calculated using the TRP integration approaches outlined in Clause 5.1.1.2.

#### 7.4.3.3	Single-layer UL-MIMO TRP Test procedure

For non-coherent UE support fullpowerMode1, the test procedure for single-layer UL-MIMO TRP Testing is as follows:

1)	Place the DUT inside the QZ following the positioning guideline defined in Clause 6.

2)	Connect the SS with the DUT through the link antenna following step 4 for ULFPTx in section 6.2D.1.4.2 of TS 38.521-1 [5] with the exception of UE configuration defined in clause 7.4.3.1 and ensure the DUT transmits with its maximum power.

3)	Measure the power at each measurement point, and calculate EIRPθ,ϕ by adding the composite loss of the entire transmission path.

The TRP value is calculated using the TRP integration approaches outlined in Clause 5.1.1.3.

For coherent UE UL-MIMO TRP testing, the common test procedure is as follows:

1)	Place the DUT inside the QZ following the positioning guideline defined in Clause 6.

2)	Connect the SS with the DUT through the link antenna following steps for ULFPTx in section 6.2D.1.4 of TS 38.521-1 [5] with the exception of UE configuration defined in clause 7.4.3.1 and ensure the DUT transmits with its maximum power.

3)	Set the SS to transmit TPMIi with [i=2].

4)	Measure the power and calculate EIRPTPMIi,θ,ϕ by adding the composite loss of the entire transmission path.

5)	Repeat steps 3) and 4) for the remaining or subset of TPMIi with [i={3,4,5}].

6)	Repeat steps 3) to 5) for each measurement grid point.

Note: Based on different test system implementation, step 5 and step 6 may be switched.

The TRP value is calculated using the TRP integration approaches outlined in Clause 5.1.1.4.

### 7.4.4	TRP test procedure for NR DL CA configuration

#### 7.4.4.1	UE configuration

For UE radiated conformance testing P-MPRc shall be 0 dB.

FR1 TRP and TRS radiated conformance testing shall be performed with the UE consistently operating at maximum power level, e.g., Time-Averaged Algorithm (TAA) and other power back-off functions should be disabled. The above functions OFF should be based on manufacturer declaration, if declared, then the manufacturer is required to provide a mechanism for the test lab to enable/disable the function.

The NR SS should send continuous uplink power control “up” commands to the DUT to ensure the DUT’s transmitter is at maximum output power during the CA TRP and TRS test.

For devices containing multiple Tx antennas, the Tx Antenna Switching (TAS), the guidelines specified in Clause 7.4.2 of TR 38.870 shall be used. However, devices supporting dual Tx in CA mode, without TAS, it can be tested as normal, namely no special handling of the Tx antennas is needed.

For CA, the SS and DUT shall be configured per TS 38.521-1 [5], Section 6.2A.1 (UE Maximum Output Power for CA) using the default settings specified in TS 38.521-1 [5] and TS 38.508-1 [7] as applicable. The measurement should be carried out based on the detailed test parameters for each band, as defined in Table 4.3.3-1 in TR 38.870.

#### 7.4.4.2	Test procedure

For UE configured with inter-band DL CA (two bands) with only single uplink CC, the measurement procedure includes the following steps:

1)	Place the DUT inside the QZ following the positioning guideline defined in Clause 6.

2)	Connect the SS with the DUT through the link antenna following steps 1 and 2 in section 6.2A.1.1.4.2 of TS 38.521-1 [5] and ensure the DUT transmits with its maximum power.

3)	Measure the power for the UL CC at each measurement point, and calculate EIRPθ,ϕ by adding the composite loss of the entire transmission path.

The TRP value is calculated using the TRP integration approaches outlined in Clause 5.1.1.2.

## 7.5	TRS Test procedure

<Editor’s note: includes test procedure for both SA and EN-DC. Consider NR 1Tx, 2Tx, RedCap, and CA.>

### 7.5.1	General

For SA and EN-DC, the TRS of the DUT is measured by sampling effective isotropic sensitivity (EIS) of the DUT with three-dimensional scan at various locations surrounding the device. The measurement is performed with the sampling step for theta () and phi () axes defined in Table B.2.12-1, Annex B.2.12, for TRS measurement.

EIS, or receiver sensitivity measurements, is defined as the minimum downlink signal power received at the UE antenna input required to provide a data throughput rate greater than or equal to 95% of the maximum throughput of the reference measurement channel (RMC) (the maximum throughput is per Appendix A of TS 38.521-1 [5]).

For TRS measurement, the evaluations shall be performed at maximum transmit power.

For TRS measurement, no specific setting is needed for Rx antennas. By default, the maximum number of Rx antennas supported at each band should be enabled during the TRS test.

### 7.5.2	TRS test procedure for NR 1Tx configuration

#### 7.5.2.1	UE configuration

<Editor’s note: RedCap configuration can be further added, mainly the band parameters and number of antennas>

For Standalone, the NR System Simulator (SS) and DUT shall be configured per section 7.3.2 (Reference sensitivity power level) of TS 38.521-1 [5] using the defaults specified in TS 38.521-1 [5] and TS 38.508-1 [7] as applicable. The measurement should be carried out based on the detailed test parameters for each band, as defined in Clause 4.

For EN-DC, the EN-DC SS and DUT shall be configured per section 7.3B.2 (Reference Sensitivity for EN-DC) of TS 38.521-3 [6], using the defaults specified in TS 38.521-3 [6] and TS 38.508-1 [7], as applicable. The measurement should be carried out based on the detailed test parameters for each band, as defined in Clause 4. The UL power configuration for LTE and NR is 50%-50% power splitting, i.e.,

-	For PC3, p-MaxEUTRA-r15=20 dBm, and p-NR-FR1= 20dBm;

-	or PC2, p-MaxEUTRA-r15=23 dBm, and p-NR-FR1= 23dBm.

Different from handheld UE, a Redcap UE is required to be equipped with a minimum of single Rx antenna and maximum of two Rx antenna.

#### 7.5.2.2	Test Procedure

<Editor’s note: in general, RedCap test procedure is the same except for forearm positioning.>

For UE configured with 1Tx for NR carrier, SA or EN-DC mode, the measurement procedure includes the following steps:

1)	Place the DUT inside the QZ following the positioning guideline defined in Clause 6.

2)	Connect the SS with the DUT through the measurement antenna.

3)	Follow steps 1 through 4 in section 7.3.2.4.2 of TS 38.521-1 [5], with the following exception: determine each EIS, i.e., by adjusting the downlink signal level until the minimum power level at which the throughput exceeds or equal to 95% of the maximum throughput of the specified RMC, at each sampling point. The downlink power step size shall be no more than 0.5 dB when the RF power level is near the NR sensitivity level.

The TRS value is calculated using the equation outlined in Clause 5.2.2.

### 7.5.3	TRS test procedure for NR 2Tx configuration

#### 7.5.3.1	UE configuration

For TRS testing of UE with TxD, the UE configuration is same as Clause 7.5.2.1, with an exception configuration that all Tx antenna shall be ON and consistently operating at maximum power level condition.

For TRS testing of UE with single-layer UL-MIMO (coherent UE and non-coherent UE), the UE configuration is same as Clause 7.5.2.1, with exception that all Tx antenna shall be ON, the UE shall be configured by fixed TPMI index = 2 and consistently operating at maximum power level condition.

#### 7.5.3.2	Test procedure

For UE TRS testing under TxD or single-layer UL-MIMO configuration, the general test procedure defined in Clause 7.5.2.2 applies.

### 7.5.4	TRS test procedure for NR DL CA configuration

<Editor’s note: In the current release of the specification, RAN4 has not introduced CA TRS requirements in Rel-18. >

#### 7.5.4.1	UE configuration

The CA SS and DUT shall be configured per section 7.3A.1 (Reference sensitivity power level for 2DL CA without exception) of TS 38.521-1 [5], using the defaults specified in TS 38.521-1 [5] and TS 38.508-1 [7], as applicable. The measurement should be carried out based on the detailed test parameters for each band, as defined in Table 4.3.3-2, in Clause 4.3.3.

#### 7.5.4.2	Test procedure

For UE configured with inter-band DL CA (two bands) with only single uplink CC, the measurement procedure includes the following steps:

1)	Place the DUT inside the QZ following the positioning guideline defined in Clause 6.

2)	Connect the SS with the DUT through the link antenna following steps 1 and 2 in section 7.3A.1.4.2 of TS 38.521-1 [5] and ensure the DUT transmits with its maximum power.

3)	Follow steps 1 through 4 in section 7.3A.1.4.2 of TS 38.521-1 [5], with the following exception: determine each EIS per CC, i.e., by adjusting the downlink signal level until the minimum power level at which the throughput exceeds or equal to 95% of the maximum throughput of the specified RMC, at each sampling point. The downlink power step size shall be no more than 0.5 dB when the RF power level is near the NR sensitivity level.

The TRS value per CC is calculated using the equation outlined in Clause 5.2.2.

## 7.6	Ripple Test for Quiet Zone

### 7.6.1	General

The ripple test procedure is defined in this clause. Frequencies to be used for ripple test:

Table 7.6.1-1: Frequencies for FR1 ripple test

| NR FR1 Bands | Range | Test frequency (MHz) |
| --- | --- | --- |
| n71 | Low | 617MHz |
| n12, n17, n29, n14, n28 |  | 722MHz |
| n5, n8, n18, n20 |  | 836.5MHz |
| n50, n51, n74 | Mid | 1575.42MHz |
| n3, n2, n25, n39 |  | 1880MHz |
| n1, n34, n65 |  | 2132.5MHz |
| n7, n30, n41, n40, n38, n90 |  | 2450MHz |
| n77, n78 | High | 3600MHz |
| n79 |  | [4700MHz] |

### 7.6.2	Ripple test procedure

Unwanted reflections and support structure blockage cause a volumetric ripple to the field magnitude measured by or created by the measurement antenna as shown in Figure 7.6.2-1, affecting every possible test point within a desired test volume. By rotating an omnidirectional antenna through the test volume as illustrated by the red line, this volumetric ripple may be probed to obtain an estimate of the measurement uncertainty due to this volumetric error. Note, however, that the volumetric ripple caused by multipath reflections is related to the wavelength and the relative path lengths of the direct and reflected paths to the measurement antenna, and is not inherently a function of the test volume. As illustrated in Figure 7.6.2-1(left), when the wavelength is relatively large compared to the test volume, it may not be possible to probe the entire range of errors that may actually exist within the test volume. Doing so would require probing an area larger than the test volume in order to accurately estimate the error within the test volume. Even when the test volume is several wavelengths in size, a given evaluation of the ripple may not completely caliper the worst-case error conditions within the test volume (Figure 7.6.2-1 (right)).

Figure 7.6.2-1: Volumetric ripple and 20cm Phi axis cut

Conversely, field non-uniformities in the test volume caused by amplitude taper of the measurement antenna and shadowing of portions of the test volume by support structure are geometric in nature and relate directly to the size of the test volume and the related factors of the measurement antenna and support structure. This test procedure attempts to capture the impact of these effects to within the practical limits of the available test volume and test equipment expected in the lab.

The quiet zone ripple test covers two cylindrical test volumes, one for handsets and smaller devices, with or without head and/or hand phantoms, and a larger volume for testing up to notebook PC sized devices. The smaller cylinder is 30 cm in diameter, concentric to the phi axis, and 30 cm tall along the phi axis, centered on the intersection of the theta and phi axes.  The larger cylinder is 50 cm in diameter, concentric to the phi axis, with the base of the cylinder coincident with that of the smaller cylinder and a height of 36 cm along the phi axis. The test consists of a set of individual ripple tests about the phi- and theta-axes utilizing both electric dipoles and magnetic loop dipoles to generate uniform omnidirectional dipole-like patterns about the axis of rotation. Data is measured on the co-polarized measurement antenna element for each corresponding test.

For the phi-axis ripple test, each reference antenna is oriented with its axis parallel to the phi axis at a total of three positions, offset 15 cm perpendicular to the phi axis with 0 cm and ±15 cm offsets parallel to the phi axis. At each position, the phi axis is rotated 360° to record the ripple. Each position is labeled by its radial and axial offset from the center position, (R, Z). See Figure 7.6.2-2 for additional information.

Figure 7.6.2-2: Phi-axis test geometry

For the notebook sized test volume, three additional test positions are added to the phi-axis test in order to cover the larger test cylinder.  These positions are offset 25 cm perpendicular to the phi axis with -15, 0, and +21 cm offsets parallel to the phi axis, as shown in Figure 7.6.2-3.

Figure 7.6.2-3: Phi-axis test geometry for the notebook sized test volume

For the theta-axis ripple test, each reference antenna is oriented with its axis parallel to the theta axis at a total of eight positions on the surface of the cylinder defined above. Defining Cartesian coordinates (X, Y, Z) with the Z-axis along the phi-axis, the Y-axis along the theta-axis, and the measurement antenna moving in the XZ plane relative to the reference antenna used for the ripple test, the test positions are given as (±15 cm, 0, ±15 cm) and (0, ±15 cm, ±15 cm), relative to the center of the test volume as shown in Figure 7.6.2-4.  At each position, the theta axis is rotated over as much of 360° as supported by the system to record the ripple (e.g. ±165°).  For systems that can only move theta in the range of 0-180° or less, the phi axis may be rotated 180° and a second cut measured to meet or exceed the ±165° range.  Each position is labeled by its (X, Y, Z) offset from the center position.

Figure 7.6.2-4: Theta-axis test geometry

For the notebook sized test volume, eight additional test positions are added to the theta-axis test in order to cover the larger test cylinder.  These positions are given as (±25 cm, 0, -15 cm), (±25 cm, 0, +21 cm), (0, ±25 cm, 0, -15 cm), and (0, ±25 cm, +21 cm), as shown in Figure 7.6.2-5.

Figure 7.6.2-5: Theta-axis test geometry for the notebook sized test volume

For each polarization and band, repeat the following steps:

For the phi-axis ripple test:

1. Place the measurement antenna and any associated theta-axis positioner at theta = 90° such that the measurement antenna is boresight with the center of the quiet zone. The measurement antenna should be at the same separation distance to be used for actual pattern measurements. This distance must be at least the minimum measurement distance away from the center of the quiet zone as defined in clause 7.7.  Select the polarization of the measurement antenna to correspond to the polarization (theta or phi) to be tested.

2. Mount the reference antenna to the phi-axis positioner using a low permittivity dielectric support. Use the sleeve dipole for the theta polarization and the loop for the phi polarization. At each of the specified offset positions, ensure that the axis of the reference antenna is parallel to the phi axis of rotation.

3. Attach a signal source to a coaxial cable feeding the measurement antenna and set the frequency to the appropriate channel. Set the amplitude to a level appropriate for the measurement receiver. Connect a measurement receiver to the reference antenna. The received signal during the ripple test measurement should be at least 40 dB above the noise floor or noise errors greater than 0.1 dB will result. Ensure that all coaxial cables are dressed to minimize effects upon the measurement results.

4. Rotate the reference antenna about the phi axis and record the signal received by the measurement antenna at resolution sufficient to ensure smoothly varying curves for a total of 360°.

5. Record the measurement results in a format suitable for calculating the ripple test metric.

6. Record test parameters including: (a) the distance between the measurement and reference antennas, (b) cable losses and other losses associated with the measurement setup, (c) the power of the signal source at the reference antenna connector, and (d) the noise level of the receiver with no signal applied.

7. Repeat steps 1 through 6 above for each reference antenna (polarization and band) for each of the required test positions.  In order to accommodate reference positioning in the lower portion of the quiet zone, support materials with a dielectric constant less than 1.2 may be removed to a maximum distance of 25 cm outside the quiet zone for the tests that require additional clearance.

For the theta-axis ripple test:

1. Place the measurement antenna such that it is boresight with the center of the quiet zone. The measurement antenna should be at the same separation distance to be used for actual pattern measurements. This distance must be at least the minimum measurement distance away from the center of the quiet zone as defined in clause 7.7. Select the polarization of the measurement antenna to correspond to the polarization (theta or phi) to be tested.

2. Mount the reference antenna in the quiet zone using a low permittivity dielectric support and such that rotating the theta positioner will cause the measurement antenna to rotate relative to the reference antenna. Use the sleeve dipole for the phi polarization and the loop for the theta polarization. At each of the specified offset positions, ensure that the axis of the reference antenna is parallel to the theta axis of rotation.

3. Attach a signal source to a coaxial cable feeding the measurement antenna and set the frequency to the appropriate channel. Set the amplitude to a level appropriate for the measurement receiver. Connect a measurement receiver to the reference antenna. The received signal during the ripple test measurement should be at least 40 dB above the noise floor or noise errors greater than 0.1 dB will result. Ensure that all coaxial cables are dressed to minimize effects upon the measurement results.

4. Rotate the reference antenna about the theta axis and record the signal received by the measurement antenna at resolution sufficient to ensure smoothly varying curves for a total of at least ±165° or the equivalent (e.g. 0-360°).  For systems that are unable to rotate a full ±165°, the reference antenna may be mounted to the phi axis and two separate theta cuts from 0 to165° may be taken, after rotating the phi axis 180° between the first and second cut.

5. Record the measurement results in a format suitable for calculating the ripple test metric.

6. Record test parameters including: (a) the distance between the measurement and reference antennas, (b) cable losses and other losses associated with the measurement setup, (c) the power of the signal source at the reference antenna connector, and (d) the noise level of the receiver with no signal applied.

7. Repeat steps 1 through 6 above for each reference antenna (polarization and band) for each of the required test positions. In order to accommodate reference positioning in the lower portion of the quiet zone, support materials with a dielectric constant less than 1.2 may be removed to a maximum distance of 25 cm outside the quiet zone for the tests that require additional clearance.

## 7.7	Minimum Range Length

This sub-section specifies the minimum range lengths for NR FR1 TRP-TRS OTA systems. The range length is defined as the distance from the centre of the quiet zone to the aperture of the measurement probes/antennas, as illustrated in Figure 7.7.1-1.

Figure 7.7.1-1: Illustration of range length definition

The minimum range length shall be the maximum of the following three limits:

-	The phase uncertainty limit: RQZ+2Drad2/

-	The amplitude uncertainty limit: 3D

-	The reactive Near-Field limit: RQZ+2

where RQZ is defined as the radius of the quiet zone, i.e., RQZ=D/2, and Drad is the diameter of the effective radiating aperture. The minimum range length calculations for D=30cm quiet zone size TRP-TRS OTA test systems shall assume that Drad is 30cm below 1GHz and decrease linearly from 30cm to 5cm from 1GHz to 7.125GHz, respectively. The last column of Table 7.7.1-1 shall be considered the minimum range length for NR FR1 TRP-TRS OTA systems with 30cm quiet zone size.

Table 7.7.1-1: Minimum Range Length for NR FR1 TRP-TRS OTA systems with 30cm quiet zone size.

| F [GHz] | Drad [m] | RQZ+2Drad²/λ | 3D = 6RQZ | RQZ+2λ | max(RQZ+2λ,3D,RQZ+2D²/λ) |
| --- | --- | --- | --- | --- | --- |
| 0.41 | 0.30 | 0.40 | 0.9 | 1.61 | 1.61 |
| 0.6 | 0.30 | 0.51 | 0.9 | 1.15 | 1.15 |
| 0.7 | 0.30 | 0.57 | 0.9 | 1.01 | 1.01 |
| 0.8 | 0.30 | 0.63 | 0.9 | 0.90 | 0.90 |
| 1 | 0.30 | 0.75 | 0.9 | 0.75 | 0.90 |
| 1.2 | 0.29 | 0.83 | 0.9 | 0.65 | 0.90 |
| 1.4 | 0.28 | 0.90 | 0.9 | 0.58 | 0.90 |
| 1.6 | 0.28 | 0.96 | 0.9 | 0.52 | 0.96 |
| 1.8 | 0.27 | 1.01 | 0.9 | 0.48 | 1.01 |
| 2 | 0.26 | 1.05 | 0.9 | 0.45 | 1.05 |
| 2.2 | 0.25 | 1.07 | 0.9 | 0.42 | 1.07 |
| 2.4 | 0.24 | 1.09 | 0.9 | 0.40 | 1.09 |
| 2.6 | 0.23 | 1.11 | 0.9 | 0.38 | 1.11 |
| 2.8 | 0.23 | 1.11 | 0.9 | 0.36 | 1.11 |
| 3 | 0.22 | 1.10 | 0.9 | 0.35 | 1.10 |
| 4 | 0.18 | 0.99 | 0.9 | 0.30 | 0.99 |
| 5 | 0.14 | 0.77 | 0.9 | 0.27 | 0.90 |
| 6 | 0.10 | 0.52 | 0.9 | 0.25 | 0.90 |
| 7 | 0.06 | 0.29 | 0.9 | 0.24 | 0.90 |
| 7.125 | 0.05 | 0.27 | 0.9 | 0.23 | 0.90 |

# 8	Reverberation Chamber test methodology

## 8.1	General

This clause defines the test method with a Reverberation Chamber (RC) system as an alternate method for FR1 TRP and TRS testing.

The applicability of alternative RC test method is specified as follows:

-	Comparison between RC system results and an accredited AC system using golden/reference UEs shall be performed on a regular basis.

-	For bands >3GHz, RC test method is harmonized with reference AC method. For the case that directional parameters are not required (e.g. UE antenna pattern, specific directional EIRP/EIS results, partial surface integrated metrics), then the RC test method can be used.

-	For certification testing, it is recommended to verify the results using AC test method as reference in case the UE does not meet the performance requirements using the RC test method.

## 8.2	Test setup

A reverberation chamber (RC) is an electrically large shielded metal enclosure that employs one or several "stirring" methods to randomize the fields, such as moving paddles, turntables, etc. In this way, a large number of uncorrelated samples is obtained. The volume in the room where the field is well-stirred is the working volume. Here the E-field, averaged over an entire stirring cycle, is independent of the location in the room, i.e., the field is spatially uniform.

Figure 8.2-1: Example Test setup diagram for Reverberation Chamber test methodology

Figure 8.2-2: S-parameter test setup diagram

## 8.3	Chamber Characterization

### 8.3.1	S-parameters and power transfer functions

Measurements of S-parameters in the reverberation chamber and derivations of power transfer functions are explained in detail in the calibration procedure.

### 8.3.2	Chamber loading for coherence bandwidth

The reverberation chamber can be loaded to control the power delay profile or coherence bandwidth in the chamber. However, the reverberation chamber should not be loaded to such an extent that the mode statistics in the chamber are impaired. It is important to keep the same loss profile in the chamber during calibration, measurement and test, in order not to change the average power transfer function between these two cases. Examples of lossy objects are blocks of RF absorber, head and hand phantoms. The configuration should be calculated to achieve a minimum coherence bandwidth of [5] times the sub-carrier spacing.

#### 8.3.2.1	Coherence bandwidth calculation

The coherence bandwidth (CBW) is a metric to determine the correlation in frequency within a working volume of a reverberation chamber (RC). To measure CBW, one should calculate the complex autocorrelation of the transmission parameter (S21) for each mode-stirred sample over a minimum [100] MHz bandwidth, as follows [21]:

R(i,n)=j=1P-iS21(fj,n)S21*(fj+i,n)j=1P-iS21(fj,n)S21*(fj,n)         											(8.1)

Where R(i,n) represents the normalized correlation function at the ith frequency step and the nth mode stirring sample. S21(fj,n) represents the measured complex S21 at the frequency step fj with [P] frequency points measured within a given bandwidth (BW) such that f1=fc-(BW2) and fP=fc+(BW2). The symbol * represents the complex conjugation acting on this parameter. n is the mode stirring sample index with a total number equal to [N]. i is the frequency offset index with a frequency resolution of fP-f1P-1 , i.e., Δf=i fP-f1P-1 , and -P-1≤i ≤P-1 [21].

The test points could be the same as in step (a) given in Section 8.3.3 namely, T=12 for chambers with a turntable or T=24 for chambers without a turntable.

To measure CBW, follow these steps:

(a) Select the center frequency fc and BW. A typical measurement frequency range of min [100] MHz channel bandwidth is used as the default.

(b) Determine the number of measurement points P with respect to frequency.

(c) Determine the total number of stirring sequence N.

(d) Measure the S-parameters, specifically S21 and calculate the correlation function using equation 8.1.

(e) Average the autocorrelation functions R(i,n) over all mode stirring sequence (i.e. index n) at each frequency point (i.e., index i).

(f) Identify the coherence bandwidth corresponding to a value of [0.5] in autocorrelation function R(i,n), as shown in Figure 1.

(g) Steps (a) through (f) may be repeated for different chamber loading configurations, e.g., using absorbers or similar materials, to meet specific coherence bandwidth requirements.

Figure 8.3.2.1-1: The CBW plots based on the correlation function for loading with two different amounts of RF absorbers in the chamber. The threshold of 0.5 is chosen as shown by the dotted lines.

### 8.3.3	Chamber spatial uniformity

The reverberation chamber shall have a working volume large enough to support the number of modes needed for the stated accuracy at the lowest operating frequency. The spatial uniformity test defines this working volume as the valid test zone of the chamber which must be large enough to contain the entire DUT plus any test scenario elements such as phantom fixtures. If DUT size is not known at the time of chamber validation, the minimum test zones described in the reference method [24] of 30cm × 30cm or 50cm × 36cm can be used in order to enable the same device size applicability. The test positions T describing this working volume need to maintain the same distance requirements as the calibration antennas throughout the procedure, i.e. reference antenna or DUT shall maintain a distance of more than 0.7 wavelengths from chamber loading and more than 0.5 wavelengths from reflective surfaces.

The value of the uncertainty contribution is determined by repeated calibration measurements for T different positions and orientations of the calibration antenna to determine the statistical variation as a function of frequency, or at least at the frequencies at which the chamber is to be used. It can be assumed that this uncertainty contribution value is normally distributed.

The T calibration configurations refer to T reference antenna positions. The operator should select these positions depending on whether or not a turntable is present in the chamber. The following instructions and explanation will assist the operator to perform the measurements.

If there is a turntable in the RC:

-	Assume a cylindrical working volume by selecting two positions on the turntable (the outermost and innermost radial positions of the turntable are recommended). Then select two elevations (the highest and lowest possible are recommended, but far enough away from absorbent and metallic objects, if any). The cylindrical volume would become a toroid with rectangular cross-section if Rmin is non-zero due to proximity effects, as described in [22]. The role of the turntable is to stir the source to obtain more independent samples.

If there is no turntable in the RC:

-	Choose eight positions as corners of an imaginary cubic volume. This means the recommended heights are the highest and lowest possible, but far enough away from absorbing and metallic objects.

The test steps are as follows.

(a) Choose either a cylindrical or a cubic working volume if there is or is not a turntable in the reverberation chamber. As shown by the red dots in Figure 2, use either the 4 corners of the imaginary rectangle that would form the cylindrical working volume or the 8 corners of the cubic working volume as locations for the measurement. For each location, point the antenna at three different angles, preferably at three orthogonal orientations (e.g., 45 degrees, -45 degrees, and horizontal plane). This results in a number of 12 or 24 measurements i.e., T=12 or T=24.

(b) Measure transmission coefficient S21 for all 12 or 24 in a complete mode stirring sequence.

(c) Calculate the power Pref,t for all 12 or 24 positions. In this way, Pref,t  is the reference power transfer function for position t of the calibration antenna.

(d) Calculate the average of power transfer function Pref over the calibration positions, i.e. T=12 or T=24 using the following relation:

Prefmean= 1T-1 t=1TPref,t

(e) Calculate the standard deviation of the power transfer function over T different calibration antenna positions by

σPref =1T-1t=1TPref,t-Prefmean2

(f) Calculate σPref in dB by

σPrefdB=10log10( Prefmean+ σPrefPrefmean )

(g) Repeat steps (a) through (f) for at least [25] frequency points evenly distributed across the NR FR1 bands.

(h) Steps (a) through (g) are repeated for various chamber loading configurations, e.g. using absorbers or similar materials, to meet specific coherence bandwidth requirements.

Figure 8.3.3-1: Illustration of a working volume with (left) for cylindrical and (right) for cubic volumes.

## 8.4	Calibration procedure

### 8.4.1	S-parameters measurement

This procedure to measure S-parameters through the reverberation chamber must be performed for each measurement setup in which the loading of the chamber has been changed. Examples of objects that change the loading are phantoms, absorbers, antennas, etc. The calibration procedure must be repeated for each frequency range as defined above. Therefore, it is advantageous if the vector network analyzer can be configured to a frequency sweep covering the defined frequencies, so that all frequencies of interest can be measured with a minimal number of measurements runs. Place all objects into the RC which will be used during TRP or TRS measurements, including a head phantom, hand phantom and fixture for the UE. This ensures that the loss in the chamber, which determines the average power transfer level, is the same during both calibration and test measurements, e.g., TRP and TRS measurements.

Figure 8.4.1-1 S-parameter measurement setup in the RC, using a vector network analyzer

Step-1: Place the calibration antenna in the chamber. The calibration antenna should be mounted on a low-loss dielectric fixture (i.e., loss tangent less than 0.05), to avoid effects from the fixture itself which may affect the UE’s radiation efficiency and mismatch factor. The calibration antenna must be placed in the chamber in such a way that it is far enough from any reflective and/or absorbing material such as walls, mode-stirrers, head phantom, or other object. This is to ensure that the environment for the calibration antenna (taken over the complete stirring sequence) resembles a free space environment. “Far enough away” depends on the type of calibration antenna used. For low gain nearly omni-directional antennas like dipoles, it is normally sufficient to ensure that this spacing is larger than 0.5 wavelengths from reflective objects and 0.7 wavelengths from absorbing objects at the lowest operating frequency. More directive calibration antennas should be oriented towards the center of the chamber. The calibration antenna should remain in the chamber during the TRP/TRS measurements.

Step-2: Calibrate the vector network analyzer with a full 2-port calibration in such a way that the vector complex S-parameters between the ports of the fixed measurement antenna and the calibration antenna can be accurately measured. Preferably, the vector network analyzer is set to perform a frequency sweep at each stirrer position. This will enable calibration of several frequency points during the same stirring sequence, thereby reducing calibration time. This will also enable frequency stirring, i.e., averaging the measured power transfer function over a small frequency bandwidth around each measured frequency point (moving frequency window). This will increase accuracy at the expense of (lower) frequency resolution.

Step-3: Place the calibration antenna in the chamber and connect it through the cables. The same procedure should be done for measurement antenna, if necessary. Measure the S-parameters for each fixed measurement antenna. The number of stirrer positions i.e. the number of S-parameter samples at each frequency point, should be chosen in such a way that it is large enough to yield an acceptable statistical contribution to the total measurement uncertainty. To this aim, at least 250 uncorrelated samples shall be used as defined in TR 37.941 clause 7.8.1 [18].

### 8.4.2	Calculation of the chamber reference transfer function

From the S-parameters obtained in the calibration measurement, the chamber reference transfer function for fixed antenna n can be calculated. The chamber reference transfer function for a fixed measurement antenna can be calculated as

Pref =1Mm=1 MS21,m2(1-R)(1-S22,ref2)⋅ 1eref

Where M is total number of stirrer positions, m is S-parameter sample number i.e., sampled at stirrer position m, the operator ⋅ denotes the absolute value, the operator  .   acting on parameter S22,ref denotes the average of complex parameters S22,ref over mode stirring sequence, S22,ref is the reflection coefficient for calibration antenna, S21,m is sample number m of the transfer function for the measurement antenna, eref is the radiation efficiency of the calibration antenna, and R is the reflection coefficient for the fixed antenna and can be calculated as

R= 1Mm=1MS11,m2=S112

Where S11,m  is sample number m of the reflection coefficient for a fixed antenna and S11 is the average of reflection coefficient.

Note that the radiation efficiency of the fixed antenna is not corrected for, because it will be the same both during calibration and measurements. Therefore, the fixed antenna’s radiation efficiency will not affect the final results.

## 8.5	TRP Test procedure

### 8.5.1	TRP for SA and EN-DC

#### 8.5.1.1	Test conditions

The TRP of the DUT is measured by sampling the S-parameters through a mode-stirring sequence and using these parameters to estimate the average power measured by any fixed measurement antenna. This data is used to determine the TRP values measured in the reverberation chamber. The measurement is performed with a step-mode stirring sequence and a frequency stirrer can be used.

#### 8.5.1.2	UE configurations

For devices containing multiple Tx antennas, the Tx Antenna Switching (TAS) function should be OFF, and the TRP should be measured for each Tx antenna individually. The antenna with better TRP is identified as the primary antenna, and the corresponding TRP result will be used to determine the pass/fail compliance. Otherwise, the primary antenna should be selected based on manufacturer declaration.

For Standalone, the NR System Simulator (SS) and DUT shall be configured per TS 38.521-1 [5], section 6.2.1 (UE maximum output power) using the default settings specified in TS 38.521-1 [5] and TS 38.508-1 [7] as applicable. The measurement should be carried out based on the detailed test parameters for each band, as defined in Clause 4.3.

For EN-DC, the SS and DUT shall be configured per TS 38.521-3 [6], Section 6.2B.1 (UE Maximum Output Power for EN-DC) using the default settings specified in TS 38.521-3 [6] and TS 38.508-1 [7] as applicable. The measurement should be carried out based on the detailed test parameters for each band, as defined in Clause 4.3. The UL output power of LTE carrier should be set as a constant power of 10dBm, while measuring NR at maximum output power, i.e., with fixed p-MaxEUTRA-r15=10 dBm, and p-NR-FR1 not configured.

#### 8.5.1.3	Test procedure

For TRP measurement, the evaluations shall be performed at maximum transmit power. The measurement procedure includes the following steps:

1.	Place the DUT inside the test zone.

2.	Connect the DUT through the cables/chamber to the instrument and ensure the DUT is transmitting at its maximum power.

3.	Measure the power at each measurement point, and calculate the TRP by averaging over the sampling sequence. At least [150] uncorrelated measurement samples shall be used. The loss of the entire transmission path must be compensated.

The TRP value is calculated using the TRP equation in the definition outlined in 5.1.2.

## 8.6	TRS Test procedure

### 8.6.1	TRS for SA and EN-DC

#### 8.6.1.1	Test conditions

The TRS of the DUT is based on searching for the lowest power received by UE for a discrete number of field combinations in the chamber. The power received by the UE at each discrete field combination that provides a throughput which is better than the specified target throughput level shall be averaged with other such measurements using different field combinations. The TRS of the DUT is estimated by measuring the transfer function for each fixed measurement antenna, which gives the minimum downlink signal power received at the UE antenna input required to provide a data throughput rate greater than or equal to 95% of the maximum throughput of the reference measurement channel (RMC).

#### 8.6.1.2	UE configurations

For Standalone, the NR System Simulator (SS) and DUT shall be configured per section 7.3.2 (Reference sensitivity power level) of TS 38.521-1 [5] using the defaults specified in TS 38.521-1 [5] and TS 38.508-1 [7] as applicable. The measurement should be carried out based on the detailed test parameters for each band, as defined in Clause 4.

For EN-DC, the EN-DC SS and DUT shall be configured per section 7.3B.2 (Reference Sensitivity for EN-DC) of TS 38.521-3 [6], using the defaults specified in TS 38.521-3 [6] and TS 38.508-1 [7], as applicable. The measurement should be carried out based on the detailed test parameters for each band, as defined in Clause 4. The UL power configuration for LTE and NR is 50%-50% power splitting, i.e.,

- For PC3, p-MaxEUTRA-r15=20 dBm, and p-NR-FR1= 20dBm;

- For PC2, p-MaxEUTRA-r15=23 dBm, and p-NR-FR1= 23dBm.

#### 8.6.1.3	Test procedure

For TRS measurement, the evaluations shall be performed at maximum transmit power.

The measurement procedure includes the following steps:

1)  Place the DUT inside the test zone.

2)  Connect the DUT through the cables/chamber to the instrument and ensure the DUT is transmitting at its maximum power.

3)  Follow steps 1 through 4 in section 7.3.2.4.2 of TS 38.521-1 [5], with the following exception: determine each EIS, i.e., by adjusting the downlink signal level until the minimum power level at which the throughput exceeds or equal to 95% of the maximum throughput of the specified RMC, at each sampling point. The downlink power step size shall be no more than 0.5 dB when the RF power level is near the NR sensitivity level. At least [100] uncorrelated measurement sampling points shall be used.

The TRS value is calculated using the equation outlined in 5.2.2.

## 8.7	Test Volume

The reverberation chamber shall have a volume large enough to support the number of modes needed for the stated accuracy at the lowest operating frequency. If the UE is moved around in the chamber during the measurement, the volume of the reverberation chamber can be reduced. In addition to the physical chamber size considerations, the test volume of the chamber must be large enough to place a reference antenna/test object at a minimum distance of one-half wavelength, i.e., λ/2 from the chamber walls, mode stirrers, and other electromagnetic reflective objects within the chamber, and 0.7 wavelengths from absorbing objects at the lowest test frequency.

# 9	Testing time reduction methodologies

## 9.1	General

<Editor’s note: Alternate test procedure/ test methods to reduce the testing time, or new measurement grid. >

## 9.2	Measurement grids for Anechoic Chamber method

The measurement grid simulations relied on representative antenna patterns for smartphone UEs.

For below 3 GHz, the device simulated is a flip-phone with dual-band GSM antenna and a separate Bluetooth antenna and is shown in Figure 9.2-1 in the BHR configuration. The simulation model for the phone includes many actual components, e.g., LCD, battery, buttons, speaker, vibration motor, shields, hinge, etc. It should be noted that the coordinate system in this simulation is not aligned with the coordinate systems for smartphone UEs in free-space and/or BH configurations, specifically Clause 6. However, given the nature of the simulation analyses, this misalignment is irrelevant.

Figure 9.2-1: Illustration of the antennas integrated inside the DUT mounted (GSM antenna on the top left, Bluetooth antenna on the top right, coordinate system bottom).

The simulations for the free-space configuration use the same coordinate system as shown in Figure 9.2-1 with the phantom completely removed from the simulation setup.

The normalized BHR patterns are shown in Figure 9.2-2 through 9.2-4 for the GSM antenna at 824 MHz and 1800 MHz and for the Bluetooth antenna at 2450 MHz, respectively. Each pattern is normalized to its respective peak, i.e., the pattern is shown with a 15dB dynamic range with a peak value of 0dB.

Figure 9.2-2: Normalized antenna patterns for GSM antenna at 824 MHz in BHR configuration. TRP of normalized pattern of -3.5 dB.

Figure 9.2-3: Normalized antenna patterns for GSM antenna at 1800 MHz in BHR configuration. TRP of normalized pattern of -4.2 dB.

Figure 9.2-4: Normalized antenna patterns for Bluetooth antenna at 2450 MHz in BHR configuration. TRP of normalized pattern of -6.5 dB.

The normalized FS patterns are shown in Figure 9.2-5 through 9.2-7 for the GSM antenna at 824 MHz and 1800 MHz and for the Bluetooth antenna at 2450 MHz, respectively.

Figure 9.2-5: Normalized antenna patterns for GSM antenna at 824 MHz in FS configuration. TRP of normalized pattern of -3.1 dB.

Figure 9.2-6: Normalized antenna patterns for GSM antenna at 1800 MHz in FS configuration. TRP of normalized pattern of -4.7 dB.

Figure 9.2-7: Normalized antenna patterns for Bluetooth antenna at 2450 MHz in FS configuration. TRP of normalized pattern of -4.3 dB.

For beyond 3 GHz, two sets of simulation results were utilized. The first set of results were for a smartphone UE in the FS and the BHR condition, illustrated in Figure 9.2-8.

Figure 9.2-8: Illustration of the simulation model for n78 patterns

The 3D antenna patterns in Figure 9.2-9 using a 30 dB dynamic range.

Figure 9.2-9: Simulated antenna patterns for n78 (3.5 GHz). Left: free space (FS), right: beside head right (BHR)

The second set of patterns are from a smartphone prototype with the antenna placed in the corner of a device with metallic ground plane.  The model is simplified and contains an LCD (glass), ground plane (PEC), camera block (PEC), opposing solid block and battery (PEC) and a plastic housing.  Dimensions of the device are 64 mm x 132 mm x 8 mm.  The antenna element is approximately 3 mm from the ground and has reasonable matching between 4-6 GHz. The peak gain is between 4.6 and 5.6 dBi, but pattern shape varies between the various frequencies.

Figure 9.2-10: Illustration of the simulation model for 4-6 GHz

The corresponding antenna patterns are illustrated in Figure 9.2-11.

Figure 9.2-11 Simulated antenna patterns for 4 GHz (top left), 4.5 GHz (top right), 5 GHz (middle left), 5.5 GHz (middle right), 6 GHz (bottom left).

The analyses in this contribution are closely aligned with those presented in [19]. The steps followed for the analyses are outlined and visualized in Table 9.2-1.

For the TRP analyses, a total of 10,000 rotation angles were applied and analysed. In order to apply uniform rotation vectors to the antenna pattern, the rotation angle around the y axis needed to be scaled. The rotations around the z axis were therefore handled in a completely random fashion, e.g., 360*rand(10000,1), while the distribution of rotations around the z axis needed to be scaled by sin(). The histograms of these two rotation angles are illustrated in Figure 9.2-12, while the uniform rotation vectors are illustrated in 3D in Figure 9.2-13.

Figure 9.2-12: Histogram of random distribution around the z and y axes.

Figure 9.2-13: Illustration of 10k random rotation vectors.

Table 9.2-1: Overview of the various steps for the TRP measurement grid analyses including overview for sample constant-step size measurement grids and pattern.

| Steps | ==1° (Reference) | ==15° | ==45° |
| --- | --- | --- | --- |
| Step 1: Import the reference pattern with very fine discretization of ==1°. |  |  |  |
| Step 2: Determine the reference TRP based on the fine grid pattern |  |  |  |
| Step 3: Apply 10k random rotation angles to the y axis followed by the z axis to the fine pattern | Plot of the discretized pattern (without rotation) | Plot of the discretized pattern (without rotation) | Plot of the discretized pattern (without rotation) |
| Step 4: Discretize the rotated fine pattern to the coarse grid with ,  for the constant-step size grids or N grid points for the constant-density grid |  |  |  |
| Step 5: Calculate the TRP based on the selected quadrature |  |  |  |
| Step 6: Calculate the standard deviation of the 10k different TRPs and the mean error from the reference TRP |  |  |  |

Various constant-step size grids were analysed for two TRP quadratures, i.e., sin(theta), the legacy quadrature, and the Clenshaw-Curtis approach. Both are currently permissible for TRP/TRS testing. The calculated standard deviations and mean errors for both quadratures are tabulated in Table 9.2.-2 for below 3 GHz and in Table 9.2-3 for beyond 3 GHz. The last column includes a background colour with the following formatting:

-	No coloured background means the std. deviation is less than 0.05dB, i.e., that this grid could be used without any additional MU impact

-	Yellow means the std. deviation is between 0.05dB and 0.25dB, i.e., this grid could be used with a small impact in MU.

-	Red means the std. deviation is beyond 0.25dB, i.e., these grids should not be considered any further.

The mean error column includes similar formatting where grids that yield an absolute value mean error between 0.05 and 0.25 dB are highlighted in yellow. These measurement grids might require a new MU element, or the mean error is included in the TIS calculation somehow. Grids with a red highlight have an absolute value of the mean error of greater than 0.25 dB and should be avoided.

Clearly, the Clenshaw-Curtis quadrature has lower uncertainties compared to the sin(theta) quadrature, especially for very coarse grids.

Table 9.2-2: Standard deviations and mean errors of TRPs after applying 10k rotations for various constant-step size measurement grids for below 3 GHz

| Constant Step-Size Grid |  |  | Standard Deviation [dB] | Mean Error [dB] |
| --- | --- | --- | --- | --- |
| = [°] | # of unique points | Quadrature | Aggregate (pooling all Patterns) | Aggregate (pooling all Patterns) |
| 1 | 64442 | Clenshaw-Curtis | 0.00 | 0.00 |
| 5 | 2522 |  | 0.00 | 0.00 |
| 10 | 614 |  | 0.00 | 0.00 |
| 15 | 266 |  | 0.00 | 0.00 |
| 30 | 62 |  | 0.01 | 0.00 |
| 45 | 26 |  | 0.05 | -0.01 |
| 60 | 14 |  | 0.11 | -0.01 |
| 1 | 64442 | sin(theta) | 0.00 | 0.00 |
| 5 | 2522 |  | 0.00 | 0.00 |
| 10 | 614 |  | 0.00 | -0.01 |
| 15 | 266 |  | 0.01 | -0.02 |
| 30 | 62 |  | 0.04 | -0.09 |
| 45 | 26 |  | 0.11 | -0.21 |
| 60 | 14 |  | 0.26 | -0.38 |

Table 9.2-3: Standard deviations and mean errors of TRPs after applying 10k rotations for various constant-step size measurement grids for beyond 3 GHz

| Constant Step-Size Grid |  |  | Standard Deviation [dB] | Mean Error [dB] |
| --- | --- | --- | --- | --- |
| = [°] | # of unique points | Quadrature | Aggregate (pooling all Patterns) | Aggregate (pooling all Patterns) |
| 1 | 64442 | Clenshaw-Curtis | 0.00 | 0.00 |
| 5 | 2522 |  | 0.00 | 0.00 |
| 15 | 266 |  | 0.00 | 0.00 |
| 30 | 62 |  | 0.11 | -0.01 |
| 45 | 26 |  | 0.22 | -0.09 |
| 60 | 14 |  | 0.34 | 0.05 |
| 1 | 64442 | sin(theta) | 0.00 | 0.00 |
| 5 | 2522 |  | 0.00 | 0.00 |
| 15 | 266 |  | 0.01 | -0.03 |
| 30 | 62 |  | 0.11 | -0.11 |
| 45 | 26 |  | 0.27 | -0.33 |
| 60 | 14 |  | 0.43 | -0.37 |

Another important set of analyses was performed for the constant-step size grids beyond 3 GHz where the EIRP at the pole of  = 180° was either set to a very small number/null (due to a potential blockage due to the positioner /pedestal), extrapolated from the second to last cut (mean of all EIRPs), or extrapolated from two neighbouring points 15° off the pole, i.e., EIRP(=165°, =0°) and EIRP(=165°, =180°). Those results for the standard deviations and the mean errors summarized in Table 9.2-4 and 9.2-5, respectively. From the standard deviations in 9.2-4, it can be observed that the impact of how the EIRP at  = 180° is treated has little effect for measurement grids as coarse as 26 grid points while a small increase in standard uncertainty can be seen for the coarsest grid with 14 points regardless of whether the final grid point does not yield a reasonable measurement or whether it is extrapolated from the second to last cut. When the pole point is not measured and considered to be a deep null, unacceptable mean errors are observed for very coarse grids. The mean errors in Table 9.2-5 show that the averaging of the pole from the 2nd to last cut or the extrapolation of the two point 15° off the pole yield very good agreement with the previous simulations that assume the grid point at the pole can be measured without any obstruction/blockage. Those two approaches are illustrated in Figure 9.2-14 for a measurement grid using the ==45° grid where the grid point at=180° (shown in blue) is extrapolated either using 8 existing measurements at=135° (shown in red) on the left or with two new grid points at=165° (shown in red) on the right. The advantage of the more extrapolation approach using two new grid points at =165° becomes more evident when considering the coarsest measurement grid presented here, i.e., ==60°, which could be considered for IoT or RedCap devices in the future. The extrapolation approach of the grid point at=180° that relies on either on the existing 8 measurements at =135° or the two close neighbours at =165° is considered necessary for above 3 GHz TRP measurement grids for the most accurate extrapolation and lowest MUs.

Figure 9.2-14: Illustration of extrapolation approaches for the ==45° grid

Table 9.2-4: Standard deviations for various constant-step size measurement grids and different extrapolation approaches for beyond 3 GHz; results for Clenshaw-Curtis quadrature only

| Constant Step-Size Grid |  | Standard Deviation [dB] |  |  |  |
| --- | --- | --- | --- | --- | --- |
| = [°] | # of unique points | Aggregate (pooling all Patterns) 180° is measured | Aggregate (pooling all Patterns) 180° is not measured | Aggregate (pooling all Patterns) 180° is averaged from previous cut | Aggregate (pooling all Patterns) 180° is extrapolated from two 165° measurements |
| 1 | 64442 | 0.00 | 0.00 | 0.00 | 0.00 |
| 5 | 2522 | 0.00 | 0.00 | 0.00 | 0.00 |
| 15 | 266 | 0.00 | 0.01 | 0.00 | 0.00 |
| 30 | 62 | 0.11 | 0.12 | 0.11 | 0.11 |
| 45 | 26 | 0.22 | 0.27 | 0.25 | 0.23 |
| 60 | 14 | 0.33 | 0.41 | 0.38 | 0.33 |

Table 9.2-5: Mean errors for various constant-step size measurement grids and different extrapolation approaches for beyond 3 GHz; results for Clenshaw-Curtis quadrature only

| Constant Step-Size Grid |  | Mean Error (TRP-TRPref) [dB] |  |  |  |
| --- | --- | --- | --- | --- | --- |
| = [°] | # of unique points | Aggregate (pooling all Patterns) 180° is measured | Aggregate (pooling all Patterns) 180° is not measured | Aggregate (pooling all Patterns) 180° is averaged from previous cut | Aggregate (pooling all Patterns) 180° is extrapolated from two 165° measurements |
| 1 | 64442 | 0.00 | 0.00 | 0.00 | 0.00 |
| 5 | 2522 | 0.00 | 0.00 | 0.00 | 0.00 |
| 15 | 266 | 0.00 | -0.01 | 0.00 | 0.00 |
| 30 | 62 | -0.01 | -0.06 | -0.01 | -0.01 |
| 45 | 26 | -0.09 | -0.21 | -0.08 | -0.08 |
| 60 | 14 | 0.05 | -0.15 | 0.08 | 0.07 |

The treatment of mean errors in the MU budget is FFS and will be finalized in RAN5, e.g., whether to consider these mean errors as a systematic uncertainty as it is currently done for FR2, i.e., the ‘Systematic error due to TRP calculation/quadrature’ MU element, specifically B.2.1.24 of [20].

## 9.3	Other solutions

<Editor’s note: other concluded solutions to reduce testing time, including analysis and outcome >

Annex A:
UE coordinate system

# A.1	Reference coordinate system

This annex defines the measurement coordinate system for NR FR1 TRP and TRS measurements.  The reference coordinate system, as defined in [9] is provided in Figure A.1-1 below while A.1-2 shows the DUT in the default alignment for Free Space.

Figure A.1-1: Reference coordinate system

Figure A.1-2: DUT default alignment to coordinate system

The following aspects are necessary:

-	A basic understanding of the top and bottom of the device is needed in order to define unambiguous DUT positioning requirements for the test, e.g., in the drawings used in this annex, the earpiece is on the top center of the device (front), the three buttons are on the bottom of the device (front) and the camera is on the top of the device (back).

-	An understanding of the origin and alignment of the coordinate system inside the test system, i.e. the directions in which the x, y, z axes point inside the test chamber, is needed in order to define unambiguous DUT orientation and measurement angles.

Annex B:
Estimation of Measurement uncertainty

# B.1	General

Individual uncertainty contributions in the TRP and TRS measurements are discussed and evaluated in this Annex. A technique for calculating the total measurement uncertainty is also presented.

An important part of a standard measurement procedure is the identification of uncertainty sources and the evaluation of the overall measurement uncertainty. There are various individual uncertainty sources in the measurement procedure that introduce a certain uncertainty contribution to the final measurement result. The approach in this standard test procedure is that the test laboratories are not limited to using some specific instruments and antenna positioners, for example.

The TRP/TRS measurement procedure can be considered to include two stages. In Stage 1 the calibration of the absolute level of the DUT measurement results is performed by means of using a calibration antenna whose absolute gain/radiation efficiency is known at the frequencies of interest. In Stage 2 the actual measurement of the 3-D pattern of the Device Under Test (DUT) is performed. The uncertainty contributions are analysed in clause B.2 while the uncertainty budget and example tables related to TRP and TRS are listed in clauses B.4 and B.5 respectively.

The calculation of the uncertainty contribution is based on the Guide to the expression of uncertainty in measurement [10]. Each individual uncertainty is expressed by its Standard Deviation (termed here as ‘standard uncertainty’) and represented by symbol U. The uncertainty contributions can be classified to two categories: Type-A uncertainties, which are statistically determined e.g. by repeated measurements, and Type-B uncertainties, which are derived from existing data e.g. data sheets. Several individual uncertainties are common in Stage 1 and Stage 2 and therefore cancel.

The procedure of forming the uncertainty budget is:

1)	Compile lists of individual uncertainty contributions for TRP or TRS measurement in both Stage 1 and Stage 2.

2)	Determine the standard uncertainty of each contribution by

a)	Determining the distribution of the uncertainty (Actual, U-shaped, rectangular, normal, etc.)

b)	Determining the maximum value of each uncertainty (unless the distribution is Actual)

c)	Calculating the standard uncertainty by dividing the uncertainty by 2 if the distribution is U-shaped, by 3 if the distribution is rectangular, and by 2 if the distribution is normal.

3)	Convert the units (if necessary) of each uncertainty element into the chose unit, i.e. dB.

4)	Combine all the standard uncertainties by the root-sum-squares (RSS) method to derive the ‘combined standard uncertainty’.

5)	Multiply the result by an expansion factor of 1.96 to derive expanded uncertainty at 95% confidence level: 1.96·Uc.

6)	Systematic errors, commonly either DUT or test system dependent, captured in the MU budget are added to the expanded uncertainty to derive the ‘total expanded uncertainty’, i.e.,

uc,total expanded=uc,expanded+uc,systematic=1.96ui2+ui,systematic

NOTE 1: The standard deviation from a data set of N samples is defined as

ui=1N-1k=1Nsk-s2

Where sk are the the respective sample results and s the mean of all N samples. For an uncertainty ui in dB, the dB values (instead of the linear powers) of sk and s are used.

The combination of uncertainties is performed using dB values for simplicity. It has been shown that using dB uncertainty values gives a slightly worse combined uncertainty result than using linear values for the uncertainties. The analysis method therefore errs on the safe side.

# B.2	MU contribution descriptions for Anechoic Chamber method

## B.2.1	Mismatch uncertainty

If the same chain configuration (e.g. including the measurement receiver; the measurement antenna and other elements) is used in both stages, the uncertainty is considered systematic and constant  0.00dB value.

If it is not the case, this uncertainty contribution has to be considered and determined by the following methods.

### B.2.1.1	Mismatch uncertainty between measurement receiver / communication tester and the measurement antenna

In a measurement configuration, when two elements (devices, networks) are connected, if the matching is not ideal, there is an uncertainty in the RF level signal passing through the connection. The magnitude of the uncertainty depends on the VSWR at the junction of the two connectors. In practical measurement system there are probably several connections in a test set-up, they will all interact and contribute to the combined mismatch uncertainty.

The total combined mismatch uncertainty is composed of 2 parts:

1)	The mismatch through the connector between two elements.

2)	The mismatch due to the interaction between two elements.

#### B.2.1.1.1	Mismatch uncertainty through the connector between two elements

Hereunder, a measurement configuration:

Figure B.2.1.1.1-1: Mismatch uncertainty through the connector

ΓMR is the complex reflection coefficient of the Measurement Receiver. This term is also applicable to the communication tester.

Γcable4 is the complex reflection coefficient of the cable4.

S21 is the forward gain in the network between the two reflection coefficients of interest.

S12 is the backward gain in the network between the two reflection coefficients of interest.

Note that S21 and S12 are set to 1 if the two parts are directly connected.

The uncertainty limits of the mismatch are calculated by means of the following formula (equation 6.1 of [11]):

Mismatch limits% voltage=ΓMR·Γcable4·S21·S12·100

These mismatch limits are divided by 2 (equation 6.2 of [11]) because of the U-shaped distribution of the mismatch uncertainty and give the following standard uncertainty:

Umismatch% voltage=ΓMR·Γcable4·S21·S12·1002

To convert this standard uncertainty in dB, we divide it by the standard uncertainty conversion factor (table 1 of [11]):

UmismatchdB=ΓMR·Γcable4·S21·S12·1002·11.5

#### B.2.1.1.2	Mismatch uncertainty due to the interaction of several elements

Previously, we presented how to determine the mismatch uncertainty between two elements through the junction (connector). Now, we introduce the other type of mismatch uncertainty, which is a result of the interaction between several elements.

Hereunder, a measurement configuration:

Figure B.2.1.1.1.2-1: Mismatch uncertainty due to the interaction of several elements

Firstly, we determine the mismatch uncertainty between junctions of the elements:

Between the MR and the cable3:

Umismatch1dB=ΓMR·Γcable3·S21·S12·1002·11.5

Between the cable3 and the cable4:

Umismatch2dB=Γcable3·Γcable4·S21·S12·1002·11.5

S21 and S12 are set to 1 because there is no element between cable3 and cable 4.

Umismatch1dB=ΓMR·Γcable3·1002·11.5

Umismatch2dB=Γcable3·Γcable4·1002·11.5

Each mismatch uncertainty due to the interaction between the measurement receiver and the cable4 is determined by means of the following formula:

Umismatch_interaction1dB=ΓMR·Γcable4·S21Cable3·S12Cable3·1002·11.5

S21 and S12 are equal and correspond to the cable3 attenuation.

Umismatch_interaction1dB=ΓMR·Γcable4·S21Cable32·1002·11.5

We consider in the general case, the following measurement configuration:

Figure B.2.1.1.1.2-2: Mismatch uncertainty measurement configuration

In the general case, this uncertainty contribution can be calculated by:

Umismatch_interactionNdB=ΓMR·Γantenna·S21cable1·S12cable1·…·S21cableN·S12cableN·1002·11.5

S21=S12 for passive elements (cables…)

Umismatch_interactionNdB=ΓMR·Γantenna·S21cable12·…·S21cableN2·1002·11.5

#### B.2.1.1.3	Total combined mismatch uncertainty

The two kinds of mismatch uncertainty contributions are combined by the root-sum-squares (RSS) method to derive the total combined mismatch uncertainty.

The total combined mismatch uncertainty is equal to:

Umismatch12+…+UmismatchN2+Umismatch_interaction12+…+Umismatch_interactionN2

This formula shows that the uncertainty is frequency dependent by the way of the forward and the backward gains in the network between the two components. The uncertainty upon S21 and S12 increases with frequency.

Note that for an anechoic chamber, horn antennas are frequently used as measurement antennas. There are two kinds of horn antennas: single-polarized and dual-polarized. With the second one, it is possible to measure the co-polarized and cross-polarized signals without any movement of the measurement antenna, which reduces the cable antenna uncertainty contribution and improves the measurement stability.

To conduct the signals to the measurement receiver, the measurement system configuration using a dual-polarized horn antenna has to be completed with an RF Relay. This device will include new mismatch uncertainty contributions, which have to be determined with the previously presented calculation methods, completed by the RF relay parameters contributions, and described in the following.

### B.2.1.2	Mismatch uncertainty of the RF relay

If the same receiver chain configuration (including the measurement receiver; the measurement antenna and other elements) is used in both stages, the uncertainty is considered systematic and constant  0.00dB value.

If it is not the case, this uncertainty contribution has to be considered and determined by the following method.

The following figure describes the RF Relay with its S-parameters and the complex reflection coefficient of the inputs and output:

Figure B.2.1.2-1: Mismatch uncertainty of the RF relay

The RF relay is used to switchover the cross and direct polarization signals from the measurement antenna. To determine RF Relay mismatch uncertainty contributions, reflection coefficients for each port and the cross-talk attenuation have to be known.

The total combined mismatch uncertainty is composed of two parts:

1)	The mismatch uncertainty contributions when the RF Relay switches on the direct polarization signal

2)	The mismatch uncertainty contributions when the RF Relay switches on the cross-polarization signal

Each part is composed of two types of uncertainties introduced in the previous paragraph: the mismatch through the connector between two elements and the mismatch due to the interaction between several elements.

#### B.2.1.2.1	First part: RF Relay switched on the co-polarized signal

##### B.2.1.2.1.1	The mismatch through the connector between two elements

Between the Input1 and the port1:

Umismatch1dB=Γ1·S11·1002·11.5

Between the port3 and the Output:

Umismatch2dB=Γ3·S33·1002·11.5

Between the Input2 and the port2:

The RF Relay switchovers on the direct polarization signal. As a result, there is no mismatch uncertainty contribution.

##### B.2.1.2.1.2	Mismatch due to the interaction between two elements or more

Between the Input1 and the Output:

Umismatch_interaction1dB=Γ1·Γ3·S31·S13·1002·11.5

Between the Input1 and the Input2:

Umismatch_interaction2dB=Γ1·Γ2·S21·S12·1002·11.5

The RF Relay switchovers on the cross-polarization signal. As a result, this uncertainty contribution is usually disregarded because of the high crosstalk attenuation which is characterized by S21 and S12 S-parameters. If the crosstalk attenuation is low, this uncertainty contribution has to be considered.

Between the Input2 and the Output:

Umismatch_interaction3dB=Γ2·Γ3·S23·S32·1002·11.5

The RF Relay switchovers on the cross polarization signal. As a result; this uncertainty contribution is usually disregarded because of the high cross-talk attenuation, which is characterized by S23 and S32 S-parameters. If the crosstalk attenuation is low, this uncertainty contribution has to be considered.

#### B.2.1.2.2	Second part: RF relay switched on the cross-polarized signal

##### B.2.1.2.2.1	The mismatch through the connector between two elements

Between the Input1 and the port1: the RF Relay switchovers on the direct polarization signal. As a result, there is no mismatch uncertainty contribution.

Between the port3 and the Output:

Umismatch3dB=Γ3·S33·1002·11.5

Between the Input2 and the port2:

Umismatch4dB=Γ2·S22·1002·11.5

##### B.2.1.2.2.2	Mismatch due to the interaction between two elements or more

Between the Input1 and the Output:

Umismatch_interaction4dB=Γ1·Γ3·S31·S13·1002·11.5

The RF Relay switchovers on the cross-polarization signal. As a result, this uncertainty contribution is usually disregarded because of the high crosstalk attenuation which is characterized by S31 and S13 S-parameters. If the crosstalk attenuation is low, this uncertainty contribution has to be considered.

Between the Input1 and the Input2:

Umismatch_interaction5dB=Γ1·Γ2·S21·S12·1002·11.5

The RF Relay switchovers on the cross-polarization signal. As a result, this uncertainty contribution is usually disregarded because of the high crosstalk attenuation which is characterized by S21 and S12 S-parameters. If the crosstalk attenuation is low, this uncertainty contribution has to be considered.

Between the Input2 and the Output:

Umismatch_interaction6dB=Γ2·Γ3·S23·S32·1002·11.5

#### B.2.1.2.3	Total combined mismatch uncertainty

Each non-zero mismatch uncertainty contribution from both parts (RF Relay switched on the cross and direct polarization signal) are combined by the root-sum-squares (RSS) method to derive the total combined mismatch uncertainty.

The total combined mismatch uncertainty is equal to:

Umismatch12+…+UmismatchN2+Umismatch_interaction12+…+Umismatch_interactionN2

If a RF Relay is used to drive the cross and direct polarization signals from the dual-polarized antenna, this total combined mismatch uncertainty has to be added with all the uncertainty measurement contributions for the total combined measurement uncertainty.

## B.2.2	Insertion loss

### B.2.2.1	Insertion loss of the measurement antenna cable

If the measurement antenna cable does not move between the calibration and the DUT measurement stage, the uncertainty due to the insertion loss of the cable is assumed to be systematic. Moreover, this uncertainty is common and constant in both stages and that is why this leads to 0.00dB value.

If a different cable is used in the calibration measurement and in the DUT measurement, and the difference of the insertion loss is used in the calculations, then the overall combined standard uncertainty of the insertion loss measurement should be used in the uncertainty budget. The distribution of this uncertainty is assumed to be rectangular, in which case the standard uncertainty can be calculated as the maximum value divided by 3.

### B.2.2.2	Insertion loss of the measurement antenna attenuator (if used)

See Insertion loss of the measurement antenna cable.

If the measurement antenna attenuator is used in both stages, the uncertainty is considered systematic and constant  0.00dB value.

### B.2.2.3	Insertion loss of the RF relays (if used)

See Insertion loss of the measurement antenna cable.

If the RF relay is used in both stages, the uncertainty is considered systematic and constant  0.00dB value.

### B.2.2.4	Insertion loss: calibration antenna feed cable

The feed cable of the calibration antenna only appears in Stage 1. As a result, this uncertainty has to be considered.

This uncertainty will be measured or calculated from the manufacturer’s data in logs with a rectangular distribution (see clause 5.1.2 in [11]).

### B.2.2.5	Insertion loss: calibration antenna attenuator (if used)

If a calibration antenna attenuator is used, it only appears in Stage 1. As a result, this uncertainty has to be considered.

This uncertainty will be calculated from the manufacturer’s data in logs with a rectangular distribution (see clause 5.1.2 in [11]).

## B.2.3	Influence of the antenna cable

### B.2.3.1	Measurement antenna cable

If the measurement antenna is directional (i.e. peak gain >+5dBi e.g. horn, LPDA, etc.) and the same measurement antenna cable configuration is used for both stages, the uncertainty is considered systematic and constant  0.00dB value.

### B.2.3.2	Calibration antenna cable

If an efficiency calibration is performed, influence of the calibration antenna feed cable can be assumed to be negligible, due to data averaging.

In the case of gain calibration, the influence of the calibration antenna feed cable must be assessed by measurements. A gain calibration measurement is repeated with a reasonably differing routing of the feed cable. Largest difference between the results is entered to the uncertainty budget with a rectangular distribution.

## B.2.4	Measurement receiver: uncertainty of the absolute level

The receiving device is used to measure the received signal level in TRP tests either as an absolute level or as a relative level. Receiving device used is typically a communication tester, spectrum analyser (SA), or power meter (PM). Generally, there occurs an uncertainty contribution from limited absolute level accuracy and non-linearity.

This uncertainty will be determined from the manufacturer’s datasheet and the distribution used (see clause 5.1.2 in [11]) shall match that provided in the datasheet which are commonly quoting MUs/accuracies with a “95% confidence level” and/or a “coverage factor of 2”. In the absence of a declared distribution in the datasheet, the rectangular distribution should be used. Given the wide bandwidths of the modulated signal transmitted during the testing, the level flatness and not just the CW level accuracy shall be considered.

## B.2.5	Communication tester: uncertainty of the absolute level

The transmitter device (typically a communication tester or BS simulator) is used to drive a signal to the measurement antenna in sensitivity tests either as an absolute level or as a relative level. Receiving device used is the UE. Generally, there occurs uncertainty contribution from limited absolute level accuracy and non-linearity of the communication tester.

For practical reasons, the calibration measurement (Stage 1) should be only performed with the measurement antenna as a receiver. Hence, the uncertainty on the absolute level of the transmitter device cannot be assumed as systematic. This uncertainty will be determined from the manufacturer’s datasheet and the distribution used (see clause 5.1.2 in [11]) shall match that provided in the datasheet which are commonly quoting MUs/accuracies with a “95% confidence level” and/or a “coverage factor of 2”. In the absence of a declared distribution in the datasheet, the rectangular distribution should be used. Furthermore, the uncertainty of the non-linearity of the device is included in the absolute level uncertainty. Given the wide bandwidths of the modulated signal transmitted during the testing, the level flatness and not just the CW level accuracy shall be considered.

## B.2.6	Sensitivity measurement: output level step resolution

When output power of the communication tester is swept to reach the throughput target that defines the sensitivity threshold, the final power step resolution represents an asymmetric uncertainty contribution that can be corrected since this uncertainty is device and test system independent. The lab shall correct the TRS results by -0.5·output level step resolution of the final power step search and note the correction in the test report.

## B.2.7	Measurement distance

The uncertainty contribution from a finite measurement distance is estimated differently for Stage 1 and Stage 2.

### B.2.7.1	Offset of phase centre from axis(es) of rotation

#### B.2.7.1.1	Offset of DUT phase centre from axis(es) of rotation

In all the DUT measurements (Stage 2) defined in this test procedure the DUT and head phantom combination is rotated about the ear reference point of SAM phantom, which is also assumed to be the location of the phase centre in both angular directions of the measurements.

For some positioning systems this may be practically impossible in which case a measurement uncertainty contribution can arise because the phase centre will rotate on a non-zero radius about the centre of rotation, thereby giving a variable measurement distance. Data averaging process may lead to a partial self-cancel of this uncertainty.

The uncertainty limits of this effect are calculated by means of the following formula (uj22 of [12]):

Phase centre limits%=± estimated offset from the axis of rotationrange length·100

To convert this standard uncertainty in dB, we divide it by the standard uncertainty conversion factor (table 1 of [11]):

Uphase centre offsetdB=± estimated offset from the axis of rotationrange length·11.5·100

Because of the phase center can be anywhere between the offset limits, the distribution is assumed to have a rectangular distribution.

For hand phantom measurements, the device is aligned with the centre of the quiet zone; thus, this MU element is not applicable to hand phantoms and the MU can be considered 0 dB.

#### B.2.7.1.2	Offset of calibration antenna phase centre from axis(es) of rotation

If a gain calibration is performed in Stage 1 with a directive antenna (e.g. horn antenna), the uncertainty contribution of calibration antenna’s phase centre displacement is estimated by means of the following formula (uj21 of [12]):

Phase centre limits%=± maximum dimension of the antenna2·range length·100

To convert this standard uncertainty in dB, we divide it by the standard uncertainty conversion factor (table 1 of [11]):

Uphase centre offsetdB=± maximum dimension of the antenna2·range length·11.5·100

Because the phase centre can be anywhere between the offset limits, the distribution is assumed to have a rectangular distribution.

If a gain calibration is performed in Stage 1 with omnidirectional calibration antenna (e.g. sleeve dipoles), uncertainty should be 0.00 dB provided that care is taken in their positioning since the phase centre are easily identifiable.

For an efficiency calibration with an omnidirectional calibration antenna, the Uphase centre offset is calculated similarly as for gain calibration but the uncertainty may be divided by factor 2. This is due to correcting impact of data averaging in this type of calibration.

### B.2.7.2	Mutual coupling

In measurement of radio performances of UMTS mobile phones in speech mode, the mutual coupling uncertainty for this frequency band is a 0.00dB value (see annex A.2 in [13]).

The 0.00dB value can be extended for NR FR1 band frequencies.

### B.2.7.3	Phase curvature

This uncertainty originates from the finite far-field measurement distance, which causes phase curvature across the DUT. If the minimum measurement distance is respected, this error is assumed to be negligible.

## B.2.8	Quality of quiet zone

The uncertainty contribution due to unwanted reflections and obstructions within the anechoic chamber, including imperfect absorber treatments and the impact of positioning equipment support structure, is determined from data acquired using the ripple test methodology in clause 7.4. This data consists of single-axis pattern cuts that represent the sum of direct and reflected rays from a highly symmetrical omnidirectional radiation pattern measured at various points throughout the test volume. The data must be measured in sufficient spatial or angular resolution to accurately capture the peaks and nulls of the pattern to within a small fraction of the overall ripple contribution. In general, the worst-case peak-to-null ripple will reflect the potential error in a peak EIRP or EIS measurement for an omnidirectional DUT pattern located anywhere within the test volume. Note however that nulls in the pattern can exhibit considerably larger errors due to reflected signals being stronger than the line-of-sight signal from the null.

When measuring the range path loss in Stage 1 using a dipole pattern, the associated measurement uncertainty may be determined from the peak-to-null ripple, after relative path loss compensation, of a single radial offset ripple test, where the range is configured as for range calibration (e.g. with any extraneous support structure removed) using a rectangular distribution.

For spherically integrated quantities such as TRP and TRS, the peak-to-null ripple would overestimate the measurement uncertainty due to the inherent averaging of the various peaks and nulls as the spherical pattern is integrated. In this case, the surface standard deviation (SSD) [15] is used to obtain a statistical representation of the expected impact of ripple on the integrated power from an isotropic radiator placed anywhere within the test volume. Due to the impracticality of maintaining a constant path loss reference between individual ripple test cuts, each resultant pattern is treated individually and then the worst case SSD result is chosen as the standard uncertainty of the quiet zone.

For the phi-axis ripple tests, the pattern can be considered an equatorial (theta = 90°) cut of the isotropic pattern where every point has equal weighting on an evenly spaced spherical surface.  Thus, the standard deviation of the single cut should be equivalent to the standard deviation of the entire spherical surface.  Defining pk as the kth ripple measurement point in linear power units pk = 10PkdB10, and p as the average of all  pk values in the associated ripple test, then the standard deviation of the corresponding cut is given by:

spk=1N-1k-0N-1pkp-12

On the theta-axis ripple test each ripple test cut can be considered as a great circle cut through an isotropic pattern with the symmetrical distortions that would be produced by revolving the pattern about the phi (0-180°) axis. Thus, it becomes apparent that ripple near the poles impacts a smaller total surface area on the sphere than that near theta = 90 and 270°. In this case, sin(theta) weighting is used to generate the spherical surface weighted standard deviation as:

spk = 1N-1k=0N-1pkp-1sinθk2

Note that this equation simplifies to the previous equation when theta = 90°, so the two formulations are in fact the same, regardless of which orientation of the ripple test is used.

The standard uncertainty for the quiet zone ripple contribution to the TRP/TRS measurement is then given by the maximum SSD from all of the ripple test measurements.

## B.2.9	DUT Tx-power drift

A single point power reference measurement in the beginning and at the end of the measurement procedure is recommended to monitor the power drift of the DUT. Based on TX-power drift measurements for typical 3G UE, an uncertainty of 0.2 dB shall be entered to uncertainty budget with a rectangular distribution. If the drift measurement indicates larger drift, the actual drift shall be included to uncertainty.

In order to minimize Tx-power drift error it is recommended to interleave sensitivity and power measurement of multiple channels. This spreads the measurements over a longer period, which helps to average the drift of the TX-power.

Typical TX-power drifts of 3G UE, measured in a single angular point, DUT placed against phantom head are shown in Figure B.2.9-1.

Figure B.2.9-1: Output power variation of typical 3G UE during battery life

## B.2.10	DUT sensitivity drift

Due to statistical uncertainty of sensitivity measurement, drift in the TRS cannot be monitored similarly to TRP. An uncertainty value of 0.2dB can be used with a rectangular distribution, or the TRS drift should be measured, with a setup corresponding to the actual TRS measurement.

## B.2.11	Uncertainty related to the use of phantoms

The following elements of uncertainty contributions (when applicable) are combined by the root-sum-squares (RSS) method to derive the total combined uncertainty related to the use of phantoms.

### B.2.11.1	Uncertainty from using different types of SAM phantom

This uncertainty contribution originates from the fact that different laboratories may use the two different versions of SAM head: the SAM head phantom or the SAM phantom including the head and the shoulders. The standard SAM head is the specified phantom. However, the use of the other type of SAM is also allowed with the requirement that the resulting uncertainty contribution is considered in the uncertainty budget.

### B.2.11.2	Simulated tissue liquid uncertainty

This uncertainty will occur, if the laboratory uses a liquid which has dielectric parameters deviating more than ±15% of the target parameters.

### B.2.11.3	Uncertainty of dielectric properties and shape of the hand phantom

The hand phantom contributes to OTA measurement uncertainty due to the manufacturing tolerances of its dielectric properties and shape. The dielectric properties on the surface of the hand may differ from those of its interior, so both are included in the evaluation. The moulded exterior surface of the hand shall be measured directly with an open-ended coaxial probe. The interior hand material is evaluated indirectly, by substituting a cube-shaped sample moulded from the same material and having some exterior surfaces removed. Following procedure will be used to evaluate the dielectric properties of the hand phantom:

1.	Each hand shall be manufactured together with a reference cube of the same material. The sides of the reference cube shall be not less than 40 mm in length.

2.	The moulded surface on three orthogonal sides of the cube shall be sliced away to a depth of at least 3 mm, in order to expose interior material for evaluation. The remaining three sides of the cube shall be left untreated.

3.	Relative permittivity and conductivity shall be measured at ten different points on each of the three cut, exposed surfaces of the reference cube, and the combined interior averages (εintavg, σintavg, 30 points) and standard deviations (εintstd,σintstd,30 points) shall be calculated. Individual interior averages for each of these three sides (εinti,σinti,10 points) shall also be calculated.

4.	Relative permittivity and conductivity shall be measured at ten points on the hand phantom exterior. A measurement point shall be located to each fingertip or as close to the tip as applicable. One measurement point shall be located to the back of the hand and one to the inner surface of wrist area. The exterior averages (εext_avg, σext_avg, 10 points) and standard deviations (εext_std, σext_std, 10 points) calculated accordingly.

5.	The total averages (εavg , σavg) shall be calculated as the average of exterior and interior values by either evaluating all data points or using equations :

εavg=εext_avg+3·εint_avg4

σavg=σext_avg+3·σint_avg4

6.	The total standard deviations (εstd , σstd) shall be calculated as the statistical combination of exterior and interior values by either evaluating all data points or using equations:

εstd=14εext_std2+εext_avg2+3*εint_std2+εint_avg2-εavg2

σstd=14σext_std2+σext_avg2+3*σint_std2+σint_avg2-σavg2

7.	The hands are acceptable for radiated performance testing, i.e., meet the minimal requirements, if

a.	εavg deviates by less than 15% from the target values

b.	σavg deviates by less than 25% from the target values

c.	the difference between the averaged permittivity of each 10-point interior surface εint_avg deviates by less than 10% and εext_avg by less than 20% from the total average εavg

d.	the difference between the averaged conductivity of each 10-point interior surfaceσint_avg deviates by less than 20% and σext_avg by less than 30% from the total average σavg

e.	the standard deviation of the combined measurements (30 interior points and 10 exterior points) is less than 20% for permittivity εstd and less than 40% for conductivity σstd

8.	For the hands meeting the minimal requirements of step 7, the following approximations shall be used to determine the hand uncertainty due to dielectric properties.

UεdB=c1·10log101+∆εavg2+εunc2+a1εstd2ε

UσdB=c2·10log101+∆σavg2+σunc2+a1σstd2ε

∆εavg, ∆σavg, εstd, σstd are the values determined as defined above and εunc and σunc are expanded measurement uncertainties (k = 2) of the dielectric parameter measurement method. The cube will be provided together with the hand such that the user can evaluate if the interior (cube) properties of the hand has degenerated over time by performing the test above. Coefficient c1=0.78, c2=0.39 and a1=0.50 were determined by numeric simulations.

In case the hand phantoms are manufactured within CAD models, the tolerance is 2% and therefore the effects shape errors are negligible. If the tolerance is larger, a numerical study must be conducted.

The assessment of uncertainties for the phantoms defined in CTIA Certification OTA Test Plan 01.72 Section 2 [25], is further detailed in CTIA Certification OTA Test Plan 01.70 Section 2 [26]. Values are from the CTIA Certification Test Plan for Wireless Device Over-the-Air Performance © CTIA Certification. Reproduced with permission.

### B.2.11.4	Uncertainty from using different types of Laptop Ground Plane phantom

This uncertainty contribution originates from the fact that different laboratories may use different variations of Laptop Ground Plane phantom. The standard Laptop Ground Plane is the specified phantom.

### B.2.11.5	Positioning Uncertainty from using Phantoms

Some uncertainty also occurs from the positioning of the DUT against the SAM phantom, as the DUT cannot be attached exactly in the same way every time. This uncertainty depends on how much the DUT’s positioning against the SAM phantom and hand phantoms varies from the specified testing positions. It is noted that the uncertainty of the phone positioning depends on the phone holder and the measurement operator and is in fact difficult to distinguish from random uncertainty. Some uncertainty also occurs from the positioning of the DUT plugged into the Laptop Ground Plane phantom, as the DUT may not be plugged into the USB connector and positioned exactly in the same way every time. This uncertainty depends on how much the DUT’s position plugged into the Laptop Ground Plane phantom varies from the specified plug-in position. Therefore, the positioning uncertainty is included in random uncertainty.

To estimate this uncertainty for the SAM phantom, it is suggested to perform at least five evaluations of TRP/TRS whereby the device shall be dismounted and newly positioned with a fully charged battery before each test. This measurement set has to be carried out in mid channel of lowest and highest frequency bands utilized by the testing lab, for at least three phones with different type of mechanical design. The values have to be normalized by the mean for each measurement set. As a result, the uncertainty contribution entered to uncertainty budget is the difference between the maximum and minimum normalized values.

With head and hand phantoms, random uncertainty evaluation may be done separately for each measurement configuration i.e. head only, browsing mode or speech mode. A speech mode random uncertainty evaluation, were both head and hand phantoms are used, can reasonably be considered to be the worst-case scenario and thus random uncertainties in other configurations to be less.

To estimate this uncertainty for the Laptop Ground Plane phantom, it is suggested to perform at least five evaluations of TRP/TRS for the plug-in position whereby the device shall be dismounted and newly positioned before each test. This measurement set has to be carried out in mid channel of lowest and highest frequency bands utilized by the testing lab, for at least three USBs with different type of mechanical design. The values have to be normalized by the mean for each measurement set. As a result, the uncertainty contribution entered to uncertainty budget is the difference between the maximum and minimum normalized value.

For the forearm phantom, the assessment of uncertainties for the phantoms defined in CTIA Certification OTA Test Plan 01.72 Section 2 [25], is further detailed in CTIA Certification OTA Test Plan 01.70 Section 2 [26]. Values are from the CTIA Certification Test Plan for Wireless Device Over-the-Air Performance © CTIA Certification. Reproduced with permission.

### B.2.11.6	Uncertainty of dielectric properties and shape of the forearm phantom

The assessment of uncertainties for the phantoms defined in CTIA Certification OTA Test Plan 01.72 Section 2 [25], is further detailed in CTIA Certification OTA Test Plan 01.70 Section 2 [26]. Values are from the CTIA Certification Test Plan for Wireless Device Over-the-Air Performance © CTIA Certification. Reproduced with permission.

## B.2.12	Coarse sampling grid

This contributor describes the uncertainty of the measured TRP/TRS value due to the finite number of measurement grid points. Decreasing of sampling density to finite number of samples affects the measurement uncertainty by two different errors. First is due to inadequate number of samples and second is a systematic discrimination approximation error in TRP and TRS equations. Different TRP quadratures also have an effect on the MU.

The grid options for TRP/TRS with associated MUs for constant-step size grids are summarized in Table B.2.12-1.

Table B.2.12-1: Grid Options for TRP/TRS with constant-step size grids

| Test Metric | Frequency Range | Quadrature | [°] | [°] | Min. Number of Grid Points (Note 1) | Std. Uncertainty [dB] | ¦Mean Error¦ (Note 3) [dB] |
| --- | --- | --- | --- | --- | --- | --- | --- |
| TRP | < 3GHz | sin() | 15 | 15 | 266 | 0 | 0 |
| TRS |  |  | 30 | 30 | 62 | 0.13 | 0 |
| TRP |  | Clenshaw-Curtis | 15 | 15 | 266 | 0 | 0 |
| TRS |  |  | 30 | 30 | 62 | 0.10 | 0 |
| TRP (Note 2) |  |  | 30 | 30 | 62 | 0.10 | 0 |
| TRS (Note 2) |  |  | 30 | 60 | 32 | 0.15 | 0 |
| TRS (Note 2) |  |  | 45 | 45 | 26 | 0.19 | 0 |
| TRP | > 3GHz | sin() | 15 | 15 | 266 | 0 | 0 |
| TRS |  |  | 30 | 30 | 62 | 0.21 | 0 |
| TRP |  | Clenshaw-Curtis | 15 | 15 | 266 | 0 | 0 |
| TRS |  |  | 30 | 30 | 62 | 0.18 | 0 |
| TRP (Note 2) |  |  | 30 | 30 | 62 | 0.18 | 0 |
| TRS (Note 2) |  |  | 30 | 60 | 32 | 0.25 | 0 |
| TRS (Note 2, 4) |  |  | 45 | 45 | 27 | 0.23 | 0.08 |
| TRS (Note 2, 5) |  |  | 45 | 45 | 25 | 0.25 | 0.08 |
| Note 1: The exact number of grid points depends on how the back pole EIRP(=180°)/EIS(=180°) is approximated due to obstruction and/or blocking. Note 2: The overall MU shall not be larger than the maximum MU limits if the coarsest measurement grid is adopted. Note 3: The inclusion of the mean error into the MU template/budget is FFS. Note 4: The EIS value at 180˚ is determined from two 165º measurements. Note 5: The EIS value at 180˚ is averaged from previous cut. |  |  |  |  |  |  |  |

The mean error in Table B.2.12-1 shall be considered a systematic uncertainty that cannot be corrected and thus shall be included in the uncertainty budget table as a systematic uncertainty added to the combined expanded uncertainty.

The legacy grid with ==15° for TRP and ==30° for TRS should be considered the default measurement grid and is recommended for certification testing.

## B.2.13	Random uncertainty

This contribution is used to account for all the unknown, unquantifiable, etc. uncertainties associated with the measurements.

Random uncertainty MU contributions are normally distributed.

The random uncertainty term, by definition, cannot be measured, or even isolated completely. However, past system definitions provide an empirical basis for a value. A fixed value of 0.25 dB (half that of FR2 SISO OTA measurements) is suggested for TRP measurements while a fixed value of 0.4 dB is suggested for TRS measurement to include a digital error rate uncertainty.

## B.2.14	Frequency response

Test systems might not be able to ensure flat frequency response across the entire channel bandwidth required for testing (e.g. up to 100MHz for NR FR1). When a frequency response correction based on the results from the system calibration measurements in Stage 1 is not possible or practical, this uncertainty has to be considered.

This uncertainty term can be estimated as described in [14] using the following formula:

εj=10logk=j-N2j+N2PLkN+1PLj

where εj is the expected relative error in the average power result for a given channel in dB, PLj, is the linear path loss at the center frequency of the given channel, PLk, is the linear path loss at each frequency point across the corresponding channel, and N is the number of frequency steps across a given channel bandwidth. The maximum deviation εj across all of the possible channels in a band shall be used to estimate the required frequency response uncertainty contribution with a rectangular distribution.

This error may be removed directly at each frequency, fj, by using the average path loss across the channel as the range loss correction rather than the path loss at the center frequency as described in [14].

For sensitivity measurements, this effect is included in the output level step resolution.

## B.2.15	Uncertainty of network analyser

This uncertainty includes all uncertainties involved in the S21 measurement with a network analyser, and will be determined from the manufacturer’s datasheet which is now commonly quoting MUs/accuracies with a “95% confidence level” and/or a “coverage factor of 2” and the distribution used (see clause 5.1.2 in [11]) shall match that provided in the datasheet. In the absence of a declared distribution in the datasheet, the rectangular distribution should be used.

## B.2.16	Uncertainty of the gain/efficiency of the calibration antenna

The calibration antenna only appears in Stage 1. Therefore, the gain/efficiency uncertainty has to be considered.

This uncertainty shall come from a calibration report (which is now commonly quoting MUs/accuracies with a “95% confidence level” and/or a “coverage factor of 2”) with traceability to a National Metrology Institute with measurement uncertainty budgets generated following the guidelines outlined in internationally accepted standards. In the absence of a declared distribution in the report, the rectangular distribution should be used.

# B.3	MU contribution descriptions for Alternative method

Some of the terms for measurement uncertainty in the Reverberation Chamber method are identical to the reference method (i.e. Anechoic Chamber). Therefore, this section only lists the additional MU contribution description that are unique to the Reverberation Chamber method.

## B.3.1	Additional Power Loss in EUT Chassis

When the EUT is small and does not add noticeable loss to the chamber, the calibration procedure outlined in section 8.4, is performed without the EUT present in the chamber. The possible difference in average chamber transmission level between the EUT measurement and the reference measurement must in this case be considered in the uncertainty evaluation. The uncertainty value for this contribution can be tested empirically by choosing a unit within a set of samples which is considered to incur the highest amount of loss (normally the largest unit), and measure the average transmission loss in the chamber with and without the test unit present in the chamber including phantoms if applicable. The difference between the two cases shall be used in the uncertainty calculation and the distribution should be assumed to be rectangular.

Alternatively, a fixed value of 0.2 dB with a rectangular distribution can be used in the uncertainty calculations.

## B.3.3	Quality of Spatial Uniformity

The standard deviation over calibration antenna positions and rotations of the transfer function, chamber reference or chamber loss S212. This measures the lack of spatial uniformity over the test zone in the reverberation chamber which is also referred as lack of isotropy, statistical ripple, or uniformity of the transfer function. This procedure is detailed in section 8.3.3.

According to the significance tests performed in [23], the uncertainty due to lack of spatial uniformity is not negligible. In this case, equation 12 in [23] holds and leads to the following uncertainty:

U=σPrefdBT

The expanded uncertainty covers the 95th percentile and is given by

U95 = K95 *U

Where K95 is the coverage factor for the two-sided 95th percentile of Student’s t-distribution given T-1 degrees of freedom and corresponds to 2.201 at T=12 (11 degrees of freedom) or 2.069 at T=24 (23 degrees of freedom).

## B.3.4	Sensitivity measurement: output level step resolution

When output power of the communication tester is swept to reach the throughput target that defines the sensitivity threshold, used power step resolution creates this uncertainty.  Output power step used in the sensitivity measurement is divided by factor 2 and then a rectangular distribution applied to obtain the uncertainty.

# B.4	MU Assessment for TRP

## B.4.1	MU Assessment for TRP in Anechoic Chamber

The uncertainty contributions related to TRP are listed in Table B.4.1-1. A preliminary example uncertainty budgets are presented in Table B.4.1-2 and Table B.4.1-3.

Table B.4.1-1 Uncertainty contributions in TRP measurement for anechoic chamber method

| UID | Description of uncertainty contribution | Details in clause |
| --- | --- | --- |
| Stage 2: DUT measurement (Figure 7.2-1, Figure 7.2-2) |  |  |
| 1 | Mismatch of receiver chain | B.2.1 |
| 2 | Insertion loss of receiver chain | B.2.2 |
| 3 | Influence of the measurement antenna cable | B.2.3 |
| 4 | Measurement Receiver: uncertainty of the absolute level | B.2.4 |
| 5 | Measurement distance | B.2.7 |
| 6 | Quality of quiet zone | B.2.8 |
| 7 | DUT Tx-power drift | B.2.9 |
| 8 | Uncertainty related to the use of phantoms | B.2.11 |
| 9 | Coarse sampling grid | B.2.12 |
| 10 | Random uncertainty | B.2.13 |
| 11 | Frequency Response | B.2.14 |
| Stage 1: Calibration measurement, network analyzer method (Figure 7.3-1) |  |  |
| 12 | Uncertainty of network analyzer | B.2.15 |
| 13 | Mismatch of receiver chain | B.2.1 |
| 14 | Insertion loss of receiver chain | B.2.2 |
| 15 | Mismatch in the connection of calibration antenna | B.2.1 |
| 16 | Influence of the calibration antenna feed cable | B.2.3 |
| 17 | Influence of the measurement antenna cable | B.2.3 |
| 18 | Uncertainty of the absolute gain/ radiation efficiency of the calibration antenna | B.2.16 |
| 19 | Measurement distance | B.2.7 |
| 20 | Quality of the Quiet Zone | B.2.8 |
| Systematic Errors |  |  |
| 21 | Systematic Error related to TRP grids | B.2.12 |

Table B.4.1-2 Preliminary example of uncertainty budget for TRP hand only (browsing mode) measurement for anechoic chamber method for NR FR1 bands

| UID | Uncertainty Source | Comment | Uncertainty Value [dB] |  | Prob Distr | Div | ci | Standard Uncertainty [dB] |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  | Below 3GHz | Above 3GHz |  |  |  | Below 3GHz | Above 3GHz |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |  |
| 1 | Mismatch of receiver chain | Гreceiver < 0.33 Гmeasurement antenna < 0.5 Cable attenuation > 3dB | 0.26 | 0.26 | U-shaped | 1.4 | 1 | 0.18 | 0.18 |
| 2 | Insertion loss of receiver chain | Systematic with Stage 1 (=> cancels) | 0 | 0 | Rectangular | 1.7 | 1 | 0.00 | 0.00 |
| 3 | Influence of the measurement antenna cable | Systematic with Stage 1 (=> cancels) | 0 | 0 | Rectangular | 1.7 | 1 | 0.00 | 0.00 |
| 4 | Measurement Receiver: uncertainty of the absolute level | Spectrum Analyzer | 0.42 | 0.54 | Normal | 2 | 1 | 0.21 | 0.27 |
| 5 | Measurement distance | DUT is not offset for hand-only phantom testing | 0 | 0 | Rectangular | 1.7 | 1 | 0.00 | 0.00 |
| 6 | Quality of quiet zone | Surface standard deviation of power measurements in ripple test | 0.5 | 0.5 | Actual | 1 | 1 | 0.50 | 0.50 |
| 7 | DUT Tx-power drift | Drift | 0.2 | 0.2 | Rectangular | 1.7 | 1 | 0.12 | 0.12 |
| 8 | Uncertainty related to the use of phantoms | Material Dielectric Constant, Material Conductivity, Geometry/Shape (incl. spacer), Data Mode Fixture | 0.64 | 0.64 | Rectangular | 1.7 | 1 | 0.37 | 0.37 |
| 9 | Coarse sampling grid | Sampling grids per Table B.2.12-1 | 0 | 0 | Actual | 1 | 1 | 0.00 | 0.00 |
| 10 | Random Uncertainty | Fixed MU to account for all the unknown, unquantifiable, etc. uncertainties | 0.25 | 0.25 | Normal | 2 | 1 | 0.13 | 0.13 |
| 11 | Frequency Response | Average path loss corrected | 0 | 0 | Rectangular | 1.7 | 1 | 0.00 | 0.00 |
| Stage 1: Calibration measurement, network analyzer method |  |  |  |  |  |  |  |  |  |
| 12 | Uncertainty of network analyzer | From datasheet of VNA with assessed transmission coefficients | 0.2 | 0.5 | Normal | 2 | 1 | 0.10 | 0.25 |
| 13 | Mismatch of receiver chain | Taken into account in VNA uncertainty term | 0 | 0 | U-shaped | 1.4 | 1 | 0.00 | 0.00 |
| 14 | Insertion loss of receiver chain | Systematic with Stage 2 (=> cancels) | 0 | 0 | Rectangular | 1.7 | 1 | 0.00 | 0.00 |
| 15 | Mismatch in the connection of calibration antenna | Taken into account in VNA setup uncertainty | 0 | 0 | U-shaped | 1.4 | 1 | 0.00 | 0.00 |
| 16 | Influence of the calibration antenna feed cable | Gain calibration with a dipole | 0.3 | 0.3 | Rectangular | 1.7 | 1 | 0.17 | 0.17 |
| 17 | Influence of the measurement antenna cable | Systematic with Stage 2 (=> cancels) | 0 | 0 | Rectangular | 1.7 | 1 | 0.00 | 0.00 |
| 18 | Uncertainty of the absolute gain/ radiation efficiency of the calibration antenna | Calibration report with traceability to a National Metrology Institute | 0.58 | 0.58 | Normal | 2 | 1 | 0.29 | 0.29 |
| 19 | Measurement distance | Dipole: aligned with phase center | 0 | 0 | Rectangular | 1.7 | 1 | 0.00 | 0.00 |
| 20 | Quality of the Quiet Zone | Peak-to-null ripple | 0.5 | 0.5 | Rectangular | 1.7 | 1 | 0.29 | 0.29 |
| Combined standard uncertainty [dB] |  |  |  |  |  |  |  | 0.84 | 0.88 |
| Expanded uncertainty [dB] (Confidence interval of 95 %) |  |  |  |  |  |  |  | 1.64 | 1.73 |
| 21 | Systematic Error related to TRP grids | mean error | 0 | 0 | Actual | 1 | 1 | 0.00 | 0.00 |
| Total Expanded uncertainty [dB] (Confidence interval of 95 %) |  |  |  |  |  |  |  | 1.64 | 1.73 |

Table B.4.1-3 Preliminary example of uncertainty budget for TRP Beside Head and Hand (Talk mode) measurement for anechoic chamber method for NR FR1 bands

| UID | Uncertainty Source | Comment | Uncertainty Value [dB] |  | Prob Distr | Div | ci | Standard Uncertainty [dB] |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  | Below 3GHz | Above 3GHz |  |  |  | Below 3GHz | Above 3GHz |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |  |
| 1 | Mismatch of receiver chain | Гreceiver < 0.33 Гmeasurement antenna < 0.5 Cable attenuation > 3dB | 0.26 | 0.26 | U-shaped | 1.4 | 1 | 0.18 | 0.18 |
| 2 | Insertion loss of receiver chain | Systematic with Stage 1 (=> cancels) | 0 | 0 | Rectangular | 1.7 | 1 | 0.00 | 0.00 |
| 3 | Influence of the measurement antenna cable | Systematic with Stage 1 (=> cancels) | 0 | 0 | Rectangular | 1.7 | 1 | 0.00 | 0.00 |
| 4 | Measurement Receiver: uncertainty of the absolute level | Spectrum Analyzer | 0.42 | 0.54 | Normal | 2 | 1 | 0.21 | 0.27 |
| 5 | Measurement distance | d=1.6m, Δd=0.05m | 0.27 | 0.27 | Rectangular | 1.7 | 1 | 0.16 | 0.16 |
| 6 | Quality of quiet zone | Surface standard deviation of power measurements in ripple test | 0.5 | 0.5 | Actual | 1 | 1 | 0.50 | 0.50 |
| 7 | DUT Tx-power drift | Drift | 0.2 | 0.2 | Rectangular | 1.7 | 1 | 0.12 | 0.12 |
| 8 | Uncertainty related to the use of phantoms | Material Dielectric Constant, Material Conductivity, Geometry/Shape (incl. spacer), Beside Head and Hand | 0.99 | 0.99 | Rectangular | 1.7 | 1 | 0.57 | 0.57 |
| 9 | Coarse sampling grid | Sampling grids per Table B.2.12-1 | 0 | 0 | Actual | 1 | 1 | 0.00 | 0.00 |
| 10 | Random Uncertainty | Fixed MU to account for all the unknown, unquantifiable, etc. uncertainties | 0.25 | 0.25 | Normal | 2 | 1 | 0.13 | 0.13 |
| 11 | Frequency Response | Average path loss corrected | 0 | 0 | Rectangular | 1.7 | 1 | 0.00 | 0.00 |
| Stage 1: Calibration measurement, network analyzer method |  |  |  |  |  |  |  |  |  |
| 12 | Uncertainty of network analyzer | From datasheet of VNA with assessed transmission coefficients | 0.2 | 0.5 | Normal | 2 | 1 | 0.10 | 0.25 |
| 13 | Mismatch of receiver chain | Taken into account in VNA uncertainty term | 0 | 0 | U-shaped | 1.4 | 1 | 0.00 | 0.00 |
| 14 | Insertion loss of receiver chain | Systematic with Stage 2 (=> cancels) | 0 | 0 | Rectangular | 1.7 | 1 | 0.00 | 0.00 |
| 15 | Mismatch in the connection of calibration antenna | Taken into account in VNA setup uncertainty | 0 | 0 | U-shaped | 1.4 | 1 | 0.00 | 0.00 |
| 16 | Influence of the calibration antenna feed cable | Gain calibration with a dipole | 0.3 | 0.3 | Rectangular | 1.7 | 1 | 0.17 | 0.17 |
| 17 | Influence of the measurement antenna cable | Systematic with Stage 2 (=> cancels) | 0 | 0 | Rectangular | 1.7 | 1 | 0.00 | 0.00 |
| 18 | Uncertainty of the absolute gain/ radiation efficiency of the calibration antenna | Calibration report with traceability to a National Metrology Institute | 0.58 | 0.58 | Normal | 2 | 1 | 0.29 | 0.29 |
| 19 | Measurement distance | Dipole: aligned with phase center | 0 | 0 | Rectangular | 1.7 | 1 | 0.00 | 0.00 |
| 20 | Quality of the Quiet Zone | Peak-to-null ripple | 0.5 | 0.5 | Rectangular | 1.7 | 1 | 0.29 | 0.29 |
| Combined standard uncertainty [dB] |  |  |  |  |  |  |  | 0.96 | 1.00 |
| Expanded uncertainty [dB] (Confidence interval of 95 %) |  |  |  |  |  |  |  | 1.88 | 1.96 |
| 21 | Systematic Error related to TRP grids | mean error | 0 | 0 | Actual | 1 | 1 | 0.00 | 0.00 |
| Total Expanded uncertainty [dB] (Confidence interval of 95 %) |  |  |  |  |  |  |  | 1.88 | 1.96 |

Table B.4.1-4 Preliminary example of uncertainty budget for TRP Wrist-Worn device measurement for anechoic chamber method for NR FR1 bands

| UID | Uncertainty Source | Comment | Uncertainty Value [dB] |  | Prob Distr | Div | ci | Standard Uncertainty [dB] |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  | Below 3GHz | Above 3GHz |  |  |  | Below 3GHz | Above 3GHz |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |  |
| 1 | Mismatch of receiver chain | Гreceiver < 0.33 Гmeasurement antenna < 0.5 Cable attenuation > 3dB | 0.26 | 0.26 | U-shaped | 1.4 | 1 | 0.18 | 0.18 |
| 2 | Insertion loss of receiver chain | Systematic with Stage 1 (=> cancels) | 0 | 0 | Rectangular | 1.7 | 1 | 0.00 | 0.00 |
| 3 | Influence of the measurement antenna cable | Systematic with Stage 1 (=> cancels) | 0 | 0 | Rectangular | 1.7 | 1 | 0.00 | 0.00 |
| 4 | Measurement Receiver: uncertainty of the absolute level | Spectrum Analyzer | 0.42 | 0.54 | Normal | 2 | 1 | 0.21 | 0.27 |
| 5 | Measurement distance | DUT is not offset for forearm phantom testing | 0 | 0 | Rectangular | 1.7 | 1 | 0 | 0 |
| 6 | Quality of quiet zone | Surface standard deviation of power measurements in ripple test | 0.5 | 0.5 | Actual | 1 | 1 | 0.50 | 0.50 |
| 7 | DUT Tx-power drift | Drift | 0.2 | 0.2 | Rectangular | 1.7 | 1 | 0.12 | 0.12 |
| 8 | Uncertainty related to the use of phantoms | Material Dielectric Constant, Material Conductivity, Geometry/Shape (incl. spacer), Forearm | 0.74 | 0.74 | Rectangular | 1.7 | 1 | 0.43 | 0.43 |
| 9 | Coarse sampling grid | Sampling grids per Table B.2.12-1 | 0 | 0 | Actual | 1 | 1 | 0.00 | 0.00 |
| 10 | Random Uncertainty | Fixed MU to account for all the unknown, unquantifiable, etc. uncertainties | 0.25 | 0.25 | Normal | 2 | 1 | 0.13 | 0.13 |
| 11 | Frequency Response | Average path loss corrected | 0 | 0 | Rectangular | 1.7 | 1 | 0.00 | 0.00 |
| Stage 1: Calibration measurement, network analyzer method |  |  |  |  |  |  |  |  |  |
| 12 | Uncertainty of network analyzer | From datasheet of VNA with assessed transmission coefficients | 0.2 | 0.5 | Normal | 2 | 1 | 0.10 | 0.25 |
| 13 | Mismatch of receiver chain | Taken into account in VNA uncertainty term | 0 | 0 | U-shaped | 1.4 | 1 | 0.00 | 0.00 |
| 14 | Insertion loss of receiver chain | Systematic with Stage 2 (=> cancels) | 0 | 0 | Rectangular | 1.7 | 1 | 0.00 | 0.00 |
| 15 | Mismatch in the connection of calibration antenna | Taken into account in VNA setup uncertainty | 0 | 0 | U-shaped | 1.4 | 1 | 0.00 | 0.00 |
| 16 | Influence of the calibration antenna feed cable | Gain calibration with a dipole | 0.3 | 0.3 | Rectangular | 1.7 | 1 | 0.17 | 0.17 |
| 17 | Influence of the measurement antenna cable | Systematic with Stage 2 (=> cancels) | 0 | 0 | Rectangular | 1.7 | 1 | 0.00 | 0.00 |
| 18 | Uncertainty of the absolute gain/ radiation efficiency of the calibration antenna | Calibration report with traceability to a National Metrology Institute | 0.58 | 0.58 | Normal | 2 | 1 | 0.29 | 0.29 |
| 19 | Measurement distance | Dipole: aligned with phase center | 0 | 0 | Rectangular | 1.7 | 1 | 0.00 | 0.00 |
| 20 | Quality of the Quiet Zone | Peak-to-null ripple | 0.5 | 0.5 | Rectangular | 1.7 | 1 | 0.29 | 0.29 |
| Combined standard uncertainty [dB] |  |  |  |  |  |  |  | 0.86 | 0.91 |
| Expanded uncertainty [dB] (Confidence interval of 95 %) |  |  |  |  |  |  |  | 1.69 | 1.78 |
| 21 | Systematic Error related to TRP grids | mean error | 0 | 0 | Actual | 1 | 1 | 0.00 | 0.00 |
| Total Expanded uncertainty [dB] (Confidence interval of 95 %) |  |  |  |  |  |  |  | 1.69 | 1.78 |

## B.4.2	MU Assessment for TRP in Reverberation Chamber

Table B.4.2-1 Uncertainty contributions in TRP measurement for reverberation chamber method

| UID | Description of uncertainty contribution | Details in clause |
| --- | --- | --- |
| Stage 2: DUT measurement |  |  |
| 1 | Mismatch of receiver chain | B.2.1 |
| 2 | Insertion loss of receiver chain | B.2.2 |
| 3 | Influence of the measurement antenna cable | B.2.3 |
| 4 | Measurement Receiver: uncertainty of the absolute level | B.2.4 |
| 5 | Quality of Spatial Uniformity | B.3.3 |
| 6 | Additional power loss in EUT chassis | B.3.1 |
| 7 | DUT Tx-power drift | B.2.9 |
| 8 | Uncertainty related to the use of phantoms | B.2.11 |
| 9 | Random uncertainty | B.2.13 |
| Stage 1: Calibration measurement |  |  |
| 10 | Uncertainty of network analyzer | B.2.15 |
| 11 | Mismatch of receiver chain | B.2.1 |
| 12 | Insertion loss of receiver chain | B.2.2 |
| 13 | Mismatch in the connection of calibration antenna | B.2.1 |
| 14 | Influence of the calibration antenna feed cable | B.2.3 |
| 15 | Influence of the measurement antenna cable | B.2.3 |
| 16 | Uncertainty of the absolute gain/ radiation efficiency of the calibration antenna | B.2.16 |
| 17 | Uniformity of transfer function | B.3.3 |

Table B.4.2-2 Preliminary example of uncertainty budget for TRP measurement hand only (browsing mode) for reverberation chamber method for NR FR1 bands

| UID | Uncertainty Source | Comment | Uncertainty Value [dB] | Prob Distr | Div | ci | Standard Uncertainty [dB] |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |
| 1 | Mismatch of receiver chain | Гreceiver <0.33 Гmeasurement antenna connection <0.5 Cable loss > 3dB | 0.26 | U-shaped | 1.41 | 1 | 0.18 |
| 2 | Insertion loss of receiver chain | Systematic with Stage 1 (=> cancels) | 0.1 | Rectangular | 1.73 | 1 | 0.06 |
| 3 | Influence of the fixed measurement antenna cable | Systematic with Stage 1 (=> cancels) | 0 | Rectangular | 1.73 | 1 | 0.00 |
| 4 | Measurement Receiver: uncertainty of the absolute level | From datasheet of communication tester or Spectrum Analyzer | 0.54 | Normal | 2 | 1 | 0.27 |
| 5 | Quality of Spatial Uniformity | Statistics of chamber | 0.5 | Actual | 1 | 1 | 0.5 |
| 6 | Additional power loss in EUT chassis | The EUT not present in the chamber during calibration measurement | 0.2 | Rectangular | 1.73 | 1 | 0.12 |
| 7 | DUT Tx-power drift | Drift | 0.2 | Rectangular | 1.73 | 1 | 0.12 |
| 8 | Uncertainty related to the use of phantoms | Material and Geometry | 0.64 | Rectangular | 1.73 | 1 | 0.37 |
| 9 | Random uncertainty | Using the same setup and stirring sequence | 0.25 | Rectangular | 1.73 | 1 | [0.14] |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |
| 10 | Uncertainty of network analyzer | Manufacturer’s uncertainty calculator, covers whole NA setup | 0.5 | Normal | 2 | 1 | 0.25 |
| 11 | Mismatch of receiver chain | Taken in to account in NA setup uncertainty | 0 | U-shaped | 1.41 | 1 | 0.00 |
| 12 | Insertion loss of receiver chain | Systematic with Stage 1 (=> cancels) | 0 | Rectangular | 1.73 | 1 | 0.00 |
| 13 | Mismatch in the connection of calibration antenna | Taken into account in NA setup uncertainty | 0 | U-shaped | 1.41 | 1 | 0.00 |
| 14 | Influence of the feed cable of the calibration antenna | Efficiency calibration used | 0 | Rectangular | 1.73 | 1 | 0.00 |
| 15 | Influence of the measurement antenna cable | Systematic with Stage 1 (=> cancels) | 0 | Rectangular | 1.73 | 1 | 0.00 |
| 16 | Uncertainty of the absolute gain/ radiation efficiency of the calibration antenna | Calibration certificate | 0.58 | Normal | 2 | 1 | [0.29] |
| 17 | Uniformity of transfer function | Statistics of chamber | 0.5 | Actual | 1 | 1 | [0.5] |
| Combined standard uncertainty |  |  |  |  |  |  | [0.97] |
| Expanded uncertainty (Confidence interval of 95 %) |  |  |  |  |  |  | [1.90] |

Table B.4.2-2 Preliminary example of uncertainty budget for TRP Beside Head and Hand (Talk mode) measurement for reverberation chamber method for NR FR1 bands

| UID | Uncertainty Source | Comment | Uncertainty Value [dB] | Prob Distr | Div | ci | Standard Uncertainty [dB] |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |
| 1 | Mismatch of receiver chain | Гreceiver <0.33 Гmeasurement antenna <0.5 | 0.26 | U-shaped | 1.41 | 1 | 0.18 |
| 2 | Insertion loss of receiver chain | Typically measured | 0.1 | Rectangular | 1.73 | 1 | 0.06 |
| 3 | Influence of the fixed measurement antenna cable | Systematic with Stage 1 (=> cancels) | 0 | Rectangular | 1.73 | 1 | 0.00 |
| 4 | Measurement Receiver: uncertainty of the absolute level | Spectrum Analyzer | 0.54 | Normal | 2 | 1 | 0.27 |
| 5 | Quality of Spatial Uniformity | Statistics of chamber | 0.5 | Actual | 1 | 1 | [0.5] |
| 6 | Additional power loss in EUT chassis | The EUT not present in the chamber during calibration measurement | 0.2 | Rectangular | 1.73 | 1 | [0.12] |
| 7 | DUT Tx-power drift | Drift | 0.2 | Rectangular | 1.73 | 1 | 0.12 |
| 8 | Uncertainty related to the use of phantoms | Material and Geometry, Beside Head/Hand | 0.99 | Rectangular | 1.73 | 1 | 0.57 |
| 9 | Random Uncertainty | Using the same setup and stirring sequence | 0.25 | Rectangular | 1.73 | 1 | [0.14] |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |
| 10 | Uncertainty of network analyzer | Manufacturer’s uncertainty calculator, covers whole NA setup | 0.5 | Normal | 2 | 1 | 0.25 |
| 11 | Mismatch of receiver chain | Taken in to account in NA setup uncertainty | 0 | U-shaped | 1.41 | 1 | 0.00 |
| 12 | Insertion loss of receiver chain | Systematic with Stage 1 (=> cancels) | 0 | Rectangular | 1.73 | 1 | 0.00 |
| 13 | Mismatch in the connection of calibration antenna | Taken into account in NA setup uncertainty | 0 | U-shaped | 1.41 | 1 | 0.00 |
| 14 | Influence of the feed cable of the calibration antenna | Efficiency calibration used | 0 | Rectangular | 1.73 | 1 | 0.00 |
| 15 | Influence of the measurement antenna cable | Systematic with Stage 1 (=> cancels) | 0 | Rectangular | 1.73 | 1 | 0.00 |
| 16 | Uncertainty of the absolute gain/ radiation efficiency of the calibration antenna | Calibration certificate | 0.58 | Normal | 2 | 1 | [0.29] |
| 17 | Uniformity of transfer function | Statistics of chamber | 0.5 | Actual | 1 | 1 | [0.5] |
| Combined standard uncertainty |  |  |  |  |  |  | [1.06] |
| Expanded uncertainty (Confidence interval of 95 %) |  |  |  |  |  |  | [2.08] |

# B.5	MU Assessment for TRS

## B.5.1	MU Assessment for TRS in Anechoic Chamber

The uncertainty contributions related to TRS are listed in Table B.5.1-1. A preliminary example uncertainty budgets are presented in Table B.5.1-2 and Table B.5.1-3.

Table B.5.1-1 Uncertainty contributions in TRS measurement for anechoic chamber method

| UID | Description of uncertainty contribution | Details in clause |
| --- | --- | --- |
| Stage 2: DUT measurement (Figure 7.2-1, Figure 7.2-2) |  |  |
| 1 | Mismatch of transmitter chain | B.2.1 |
| 2 | Insertion loss of transmitter chain | B.2.2 |
| 3 | Influence of the measurement antenna cable | B.2.3 |
| 4 | Communication Tester: uncertainty of the absolute output level | B.2.5 |
| 5 | Sensitivity measurement: output level step resolution | B.2.6 |
| 6 | Measurement distance | B.2.7 |
| 7 | Quality of quiet zone | B.2.8 |
| 8 | DUT sensitivity drift | B.2.10 |
| 9 | Uncertainty related to the use of phantoms | B.2.11 |
| 10 | Coarse sampling grid | B.2.12 |
| 11 | Random uncertainty | B.2.13 |
| 12 | Frequency Response | B.2.14 |
| Stage 1: Calibration measurement, network analyzer method (Figure 7.3-1) |  |  |
| 13 | Uncertainty of network analyzer | B.2.15 |
| 14 | Mismatch of transmitter chain | B.2.1 |
| 15 | Insertion loss of transmitter chain | B.2.2 |
| 16 | Mismatch in the connection of calibration antenna | B.2.1 |
| 17 | Influence of the calibration antenna feed cable | B.2.3 |
| 18 | Influence of the measurement antenna cable | B.2.3 |
| 19 | Uncertainty of the absolute gain/radiation efficiency of the calibration antenna | B.2.16 |
| 20 | Measurement distance | B.2.7 |
| 21 | Quality of quiet zone | B.2.8 |
| Systematic Errors |  |  |
| 22 | Systematic Error related to TRS grids | B.2.12 |

Table B.5.1-2: Preliminary example of uncertainty budget for TRS hand only (browsing mode) measurement for anechoic chamber method for NR FR1 bands

| UID | Uncertainty Source | Comment | Uncertainty Value [dB] |  | Prob Distr | Div | ci | Standard Uncertainty [dB] |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  | Below 3GHz | Above 3GHz |  |  |  | Below 3GHz | Above 3GHz |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |  |
| 1 | Mismatch of receiver chain | Гcomm tester < 0.29 Гmeasurement antenna < 0.5 Cable attenuation > 3dB | 0.22 | 0.22 | U-shaped | 1.4 | 1 | 0.16 | 0.16 |
| 2 | Insertion loss of receiver chain | Systematic with Stage 1 (=> cancels) | 0 | 0 | Rectangular | 1.7 | 1 | 0.00 | 0.00 |
| 3 | Influence of the measurement antenna cable | Systematic with Stage 1 (=> cancels) | 0 | 0 | Rectangular | 1.7 | 1 | 0.00 | 0.00 |
| 4 | Communication Tester: uncertainty of the absolute output level | Manufacturer’s data sheet | 1.3 | 1.3 | Normal | 2 | 1 | 0.65 | 0.65 |
| 5 | Sensitivity measurement: output level step resolution | Systematic error that can be corrected | 0 | 0 | Rectangular | 1.73 | 1 | 0.00 | 0.00 |
| 6 | Measurement distance | DUT is not offset for hand-only phantom testing | 0 | 0 | Rectangular | 1.7 | 1 | 0.00 | 0.00 |
| 7 | Quality of quiet zone | Surface standard deviation of power measurements in ripple test | 0.5 | 0.5 | Actual | 1 | 1 | 0.50 | 0.50 |
| 8 | DUT sensitivity drift | Drift | 0.2 | 0.2 | Rectangular | 1.7 | 1 | 0.12 | 0.12 |
| 9 | Uncertainty related to the use of phantoms | Material Dielectric Constant, Material Conductivity, Geometry/Shape (incl. spacer), Data Mode Fixture | 0.64 | 0.64 | Rectangular | 1.7 | 1 | 0.37 | 0.37 |
| 10 | Coarse sampling grid | Sampling grids per Table B.2.12-1 | 0.10 | 0.18 | Actual | 1 | 1 | 0.10 | 0.18 |
| 11 | Random Uncertainty | Fixed MU to account for all the unknown, unquantifiable, etc. uncertainties including digital error rate | 0.4 | 0.4 | Normal | 2 | 1 | 0.20 | 0.20 |
| 12 | Frequency Response | Included in the output level step resolution | 0 | 0 | Rectangular | 1.7 | 1 | 0.00 | 0.00 |
| Stage 1: Calibration measurement, network analyzer method |  |  |  |  |  |  |  |  |  |
| 13 | Uncertainty of network analyzer | From datasheet of VNA with assessed transmission coefficients | 0.2 | 0.5 | Normal | 2 | 1 | 0.10 | 0.25 |
| 14 | Mismatch of receiver chain | Taken into account in VNA uncertainty term | 0 | 0 | U-shaped | 1.4 | 1 | 0.00 | 0.00 |
| 15 | Insertion loss of receiver chain | Systematic with Stage 2 (=> cancels) | 0 | 0 | Rectangular | 1.7 | 1 | 0.00 | 0.00 |
| 16 | Mismatch in the connection of calibration antenna | Taken in to account in VNA setup uncertainty | 0 | 0 | U-shaped | 1.4 | 1 | 0.00 | 0.00 |
| 17 | Influence of the calibration antenna feed cable | Gain calibration with a dipole | 0.3 | 0.3 | Rectangular | 1.7 | 1 | 0.17 | 0.17 |
| 18 | Influence of the measurement antenna cable | Systematic with Stage 2 (=> cancels) | 0 | 0 | Rectangular | 1.7 | 1 | 0.00 | 0.00 |
| 19 | Uncertainty of the absolute gain/ radiation efficiency of the calibration antenna | Calibration report with traceability to a National Metrology Institute | 0.58 | 0.58 | Normal | 2 | 1 | 0.29 | 0.29 |
| 20 | Measurement distance | Dipole: aligned with phase center | 0 | 0 | Rectangular | 1.7 | 1 | 0.00 | 0.00 |
| 21 | Quality of the Quiet Zone | Peak-to-null ripple | 0.5 | 0.5 | Rectangular | 1.7 | 1 | 0.29 | 0.29 |
| Combined standard uncertainty [dB] |  |  |  |  |  |  |  | 1.05 | 1.09 |
| Expanded uncertainty [dB] (Confidence interval of 95 %) |  |  |  |  |  |  |  | 2.06 | 2.13 |
| 22 | Systematic Error related to TRS grids | mean error | 0 | 0 | Actual | 1 | 1 | 0.00 | 0.00 |
| Total Expanded uncertainty [dB] (Confidence interval of 95 %) |  |  |  |  |  |  |  | 2.06 | 2.13 |

Table B.5.1-3: Preliminary example of uncertainty budget for TRS Beside Head and Hand (Talk mode) measurement for anechoic chamber method for NR FR1 bands

| UID | Uncertainty Source | Comment | Uncertainty Value [dB] |  | Prob Distr | Div | ci | Standard Uncertainty [dB] |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  | Below 3GHz | Above 3GHz |  |  |  | Below 3GHz | Above 3GHz |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |  |
| 1 | Mismatch of receiver chain | Гcomm tester < 0.29 Гmeasurement antenna < 0.5 Cable attenuation > 3dB | 0.22 | 0.22 | U-shaped | 1.4 | 1 | 0.16 | 0.16 |
| 2 | Insertion loss of receiver chain | Systematic with Stage 1 (=> cancels) | 0 | 0 | Rectangular | 1.7 | 1 | 0.00 | 0.00 |
| 3 | Influence of the measurement antenna cable | Systematic with Stage 1 (=> cancels) | 0 | 0 | Rectangular | 1.7 | 1 | 0.00 | 0.00 |
| 4 | Communication Tester: uncertainty of the absolute output level | Manufacturer’s data sheet | 1.3 | 1.3 | Normal | 2 | 1 | 0.65 | 0.65 |
| 5 | Sensitivity measurement: output level step resolution | Systematic error that can be corrected | 0 | 0 | Rectangular | 1.73 | 1 | 0.00 | 0.00 |
| 6 | Measurement distance | d=1.6m, Δd=0.05m | 0.27 | 0.27 | Rectangular | 1.7 | 1 | 0.16 | 0.16 |
| 7 | Quality of quiet zone | Surface standard deviation of power measurements in ripple test | 0.5 | 0.5 | Actual | 1 | 1 | 0.50 | 0.50 |
| 8 | DUT sensitivity drift | Drift | 0.2 | 0.2 | Rectangular | 1.7 | 1 | 0.12 | 0.12 |
| 9 | Uncertainty related to the use of phantoms | Material Dielectric Constant, Material Conductivity, Geometry/Shape (incl. spacer), Beside Head and Hand | 0.99 | 0.99 | Rectangular | 1.7 | 1 | 0.57 | 0.57 |
| 10 | Coarse sampling grid | Sampling grids per Table B.2.12-1 | 0.1 | 0.18 | Actual | 1 | 1 | 0.1 | 0.18 |
| 11 | Random Uncertainty | Fixed MU to account for all the unknown, unquantifiable, etc. uncertainties including digital error rate | 0.4 | 0.4 | Normal | 2 | 1 | 0.20 | 0.20 |
| 12 | Frequency Response | Included in the output level step resolution | 0 | 0 | Rectangular | 1.7 | 1 | 0.00 | 0.00 |
| Stage 1: Calibration measurement, network analyzer method |  |  |  |  |  |  |  |  |  |
| 13 | Uncertainty of network analyzer | From datasheet of VNA with assessed transmission coefficients | 0.2 | 0.5 | Normal | 2 | 1 | 0.10 | 0.25 |
| 14 | Mismatch of receiver chain | Taken into account in VNA uncertainty term | 0 | 0 | U-shaped | 1.4 | 1 | 0.00 | 0.00 |
| 15 | Insertion loss of receiver chain | Systematic with Stage 2 (=> cancels) | 0 | 0 | Rectangular | 1.7 | 1 | 0.00 | 0.00 |
| 16 | Mismatch in the connection of calibration antenna | Taken in to account in VNA setup uncertainty | 0 | 0 | U-shaped | 1.4 | 1 | 0.00 | 0.00 |
| 17 | Influence of the calibration antenna feed cable | Gain calibration with a dipole | 0.3 | 0.3 | Rectangular | 1.7 | 1 | 0.17 | 0.17 |
| 18 | Influence of the measurement antenna cable | Systematic with Stage 2 (=> cancels) | 0 | 0 | Rectangular | 1.7 | 1 | 0.00 | 0.00 |
| 19 | Uncertainty of the absolute gain/ radiation efficiency of the calibration antenna | Calibration report with traceability to a National Metrology Institute | 0.58 | 0.58 | Normal | 2 | 1 | 0.29 | 0.29 |
| 20 | Measurement distance | Dipole: aligned with phase center | 0 | 0 | Rectangular | 1.7 | 1 | 0.00 | 0.00 |
| 21 | Quality of the Quiet Zone | Peak-to-null ripple | 0.5 | 0.5 | Rectangular | 1.7 | 1 | 0.29 | 0.29 |
| Combined standard uncertainty [dB] |  |  |  |  |  |  |  | 1.15 | 1.18 |
| Expanded uncertainty [dB] (Confidence interval of 95 %) |  |  |  |  |  |  |  | 2.25 | 2.31 |
| 22 | Systematic Error related to TRS grids | mean error | 0 | 0 | Actual | 1 | 1 | 0.00 | 0.00 |
| Total Expanded uncertainty [dB] (Confidence interval of 95 %) |  |  |  |  |  |  |  | 2.25 | 2.31 |

Table B.5.1-4: Preliminary example of uncertainty budget for TRS Wrist-Worn device measurement for anechoic chamber method for NR FR1 bands

| UID | Uncertainty Source | Comment | Uncertainty Value [dB] |  | Prob Distr | Div | ci | Standard Uncertainty [dB] |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  | Below 3GHz | Above 3GHz |  |  |  | Below 3GHz | Above 3GHz |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |  |
| 1 | Mismatch of receiver chain | Гcomm tester < 0.29 Гmeasurement antenna < 0.5 Cable attenuation > 3dB | 0.22 | 0.22 | U-shaped | 1.4 | 1 | 0.16 | 0.16 |
| 2 | Insertion loss of receiver chain | Systematic with Stage 1 (=> cancels) | 0 | 0 | Rectangular | 1.7 | 1 | 0.00 | 0.00 |
| 3 | Influence of the measurement antenna cable | Systematic with Stage 1 (=> cancels) | 0 | 0 | Rectangular | 1.7 | 1 | 0.00 | 0.00 |
| 4 | Communication Tester: uncertainty of the absolute output level | Manufacturer’s data sheet | 1.3 | 1.3 | Normal | 2 | 1 | 0.65 | 0.65 |
| 5 | Sensitivity measurement: output level step resolution | Systematic error that can be corrected | 0 | 0 | Rectangular | 1.73 | 1 | 0.00 | 0.00 |
| 6 | Measurement distance | DUT is not offset for forearm phantom testing | 0 | 0 | Rectangular | 1.7 | 1 | 0 | 0 |
| 7 | Quality of quiet zone | Surface standard deviation of power measurements in ripple test | 0.5 | 0.5 | Actual | 1 | 1 | 0.50 | 0.50 |
| 8 | DUT sensitivity drift | Drift | 0.2 | 0.2 | Rectangular | 1.7 | 1 | 0.12 | 0.12 |
| 9 | Uncertainty related to the use of phantoms | Material Dielectric Constant, Material Conductivity, Geometry/Shape (incl. spacer), Forearm | 0.74 | 0.74 | Rectangular | 1.7 | 1 | 0.43 | 0.43 |
| 10 | Coarse sampling grid | Sampling grids per Table B.2.12-1 | 0.1 | 0.18 | Actual | 1 | 1 | 0.1 | 0.18 |
| 11 | Random Uncertainty | Fixed MU to account for all the unknown, unquantifiable, etc. uncertainties including digital error rate | 0.4 | 0.4 | Normal | 2 | 1 | 0.20 | 0.20 |
| 12 | Frequency Response | Included in the output level step resolution | 0 | 0 | Rectangular | 1.7 | 1 | 0.00 | 0.00 |
| Stage 1: Calibration measurement, network analyzer method |  |  |  |  |  |  |  |  |  |
| 13 | Uncertainty of network analyzer | From datasheet of VNA with assessed transmission coefficients | 0.2 | 0.5 | Normal | 2 | 1 | 0.10 | 0.25 |
| 14 | Mismatch of receiver chain | Taken into account in VNA uncertainty term | 0 | 0 | U-shaped | 1.4 | 1 | 0.00 | 0.00 |
| 15 | Insertion loss of receiver chain | Systematic with Stage 2 (=> cancels) | 0 | 0 | Rectangular | 1.7 | 1 | 0.00 | 0.00 |
| 16 | Mismatch in the connection of calibration antenna | Taken in to account in VNA setup uncertainty | 0 | 0 | U-shaped | 1.4 | 1 | 0.00 | 0.00 |
| 17 | Influence of the calibration antenna feed cable | Gain calibration with a dipole | 0.3 | 0.3 | Rectangular | 1.7 | 1 | 0.17 | 0.17 |
| 18 | Influence of the measurement antenna cable | Systematic with Stage 2 (=> cancels) | 0 | 0 | Rectangular | 1.7 | 1 | 0.00 | 0.00 |
| 19 | Uncertainty of the absolute gain/ radiation efficiency of the calibration antenna | Calibration report with traceability to a National Metrology Institute | 0.58 | 0.58 | Normal | 2 | 1 | 0.29 | 0.29 |
| 20 | Measurement distance | Dipole: aligned with phase center | 0 | 0 | Rectangular | 1.7 | 1 | 0.00 | 0.00 |
| 21 | Quality of the Quiet Zone | Peak-to-null ripple | 0.5 | 0.5 | Rectangular | 1.7 | 1 | 0.29 | 0.29 |
| Combined standard uncertainty [dB] |  |  |  |  |  |  |  | 1.07 | 1.11 |
| Expanded uncertainty [dB] (Confidence interval of 95 %) |  |  |  |  |  |  |  | 2.10 | 2.17 |
| 22 | Systematic Error related to TRS grids | mean error | 0 | 0 | Actual | 1 | 1 | 0.00 | 0.00 |
| Total Expanded uncertainty [dB] (Confidence interval of 95 %) |  |  |  |  |  |  |  | 2.10 | 2.17 |

## B.5.2	MU Assessment for TRS in Reverberation Chamber

Table B.5.2-1 Uncertainty contributions in TRS measurement for reverberation chamber method

| UID | Description of uncertainty contribution | Details in clause |
| --- | --- | --- |
| Stage 2: DUT measurement |  |  |
| 1 | Mismatch of transmitter chain | B.2.1 |
| 2 | Insertion loss of transmitter chain | B.2.2 |
| 3 | Influence of the measurement antenna cable | B.2.3 |
| 4 | Communication Tester: uncertainty of the absolute output level | B.2.5 |
| 5 | Sensitivity measurement: output level step resolution | B.3.4 |
| 6 | Quality of Spatial Uniformity | B.3.3 |
| 7 | Additional power loss in EUT chassis | B.3.1 |
| 8 | DUT sensitivity drift | B.2.10 |
| 9 | Uncertainty related to the use of phantoms | B.2.11 |
| 10 | Random uncertainty | B.2.13 |
| Stage 1: Calibration measurement |  |  |
| 11 | Uncertainty of network analyzer | B.2.15 |
| 12 | Mismatch of transmitter chain | B.2.1 |
| 13 | Insertion loss of transmitter chain | B.2.2 |
| 14 | Mismatch in the connection of calibration antenna | B.2.1 |
| 15 | Influence of the calibration antenna feed cable | B.2.3 |
| 16 | Influence of the measurement antenna cable | B.2.3 |
| 17 | Uncertainty of the absolute gain/radiation efficiency of the calibration antenna | B.2.16 |
| 18 | Uniformity of transfer function | B.3.3 |

Table B.5.2-2: Preliminary example of uncertainty budget for TRS hand only (browsing mode) measurement for reverberation chamber method for NR FR1 bands

| UID | Uncertainty Source | Comment | Uncertainty Value [dB] | Prob Distr | Div | ci | Standard Uncertainty [dB] |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |
| 1 | Mismatch of transmitter chain | Гcomm tester <0.29 Г meas antenna <0.5 Cable loss > 3dB | 0.22 | U-shaped | 1.41 | 1 | 0.16 |
| 2 | Insertion loss of transmitter chain | Systematic with Stage 1 (=> cancels) | 0 | Rectangular | 1.73 | 1 | 0 |
| 3 | Influence of the fixed measurement antenna cable | Systematic with Stage 1 (=> cancels) | 0 | Rectangular | 1.73 | 1 | 0 |
| 4 | Communication tester: uncertainty of the absolute level | From datasheet of communication tester | 1.3 | Normal | 2 | 1 | 0.65 |
| 5 | Sensitivity measurement: output level step resolution | Step 0.5dB | 0.25 | Rectangular | 1.73 | 1 | 0.14 |
| 6 | Quality of Spatial Uniformity | Statistics of chamber | 0.5 | Actual | 1 | 1 | 0.5 |
| 7 | Additional power loss in EUT chassis | The EUT not present in the chamber during calibration measurement | 0.2 | Rectangular | 1.73 | 1 | [0.12] |
| 8 | DUT sensitivity drift | Drift measurement | 0.2 | Rectangular | 1.73 | 1 | 0.12 |
| 9 | Uncertainty related to the use of phantoms | Material and Geometry for Data Mode Fixture | 0.64 | Rectangular | 1.73 | 1 | 0.37 |
| 10 | Random Uncertainty | Using the same setup and stirring sequence | 0.4 | Normal | 2 | 1 | 0.2 |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |
| 11 | Uncertainty of network analyzer | Manufacturer’s uncertainty calculator, covers NA setup | 0.5 | Normal | 2 | 1 | 0.25 |
| 12 | Mismatch of transmitter chain | Taken in to account in NA setup uncertainty | 0 | U-shaped | 1.41 | 1 | 0 |
| 13 | Insertion loss of transmitter chain | Systematic with Stage 2 (=> cancels) | 0 | Rectangular | 1.73 | 1 | 0 |
| 14 | Mismatch in the connection of calibration antenna | Taken into account in NA setup uncertainty | 0 | U-shaped | 1.73 | 1 | 0 |
| 15 | Influence of the feed cable of the calibration antenna | Gain calibration with dipole | 0.3 | Rectangular | 1.73 | 1 | 0.17 |
| 16 | Influence of the fixed measurement antenna cable | Systematic with Stage 2 (=> cancels) | 0 | Rectangular | 1.73 | 1 | 0 |
| 17 | Uncertainty of the absolute gain of the calibration antenna | Calibration certificate | 0.58 | Normal | 2 | 1 | [0.29] |
| 18 | Uniformity of transfer function | Statistics of chamber | 0.5 | Actual | 1 | 1 | 0.5 |
| Combined standard uncertainty |  |  |  |  |  |  | [1.16] |
| Expanded uncertainty (Confidence interval of 95 %) |  |  |  |  |  |  | [2.28] |

Table B.5.2-2: Preliminary example of uncertainty budget for TRS Beside Head and Hand (Talk mode) measurement for reverberation chamber method for NR FR1 bands

| UID | Uncertainty Source | Comment | Uncertainty Value [dB] | Prob Distr | Div | ci | Standard Uncertainty [dB] |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |
| 1 | Mismatch of transmitter chain | Гcomm tester <0.29 Г meas antenna <0.5 Cable loss >3dB | 0.22 | U-shaped | 1.41 | 1 | 0.16 |
| 2 | Insertion loss of transmitter chain | Systematic with Stage 1 (=> cancels) | 0 | Rectangular | 1.73 | 1 | 0 |
| 3 | Influence of the fixed measurement antenna cable | Systematic with Stage 1 (=> cancels) | 0 | Rectangular | 1.73 | 1 | 0 |
| 5 | Communication tester: uncertainty of the absolute level | Manufacturer’s data sheet | 1.3 | Normal | 2 | 1 | 0.65 |
| 6 | Sens. measurement: output level step resolution | Step 0.5dB | 0.25 | Rectangular | 1.73 | 1 | 0.14 |
| 6 | Quality of Spatial Uniformity | Statistics of chamber | 0.5 | Actual | 1 | 1 | 0.5 |
| 7 | Additional power loss in EUT chassis | The EUT not present in the chamber during calibration measurement | 0.2 | Rectangular | 1.73 | 1 | [0.12] |
| 8 | DUT sensitivity drift | Drift measurement | 0.2 | Rectangular | 1.73 | 1 | 0.12 |
| 9 | Uncertainty related to the use of phantoms | Material and Geometry for Data Mode Fixture | 0.99 | Rectangular | 1.73 | 1 | 0.57 |
| 10 | Random Uncertainty | Using the same setup and stirring sequence | 0.4 | Normal | 2 | 1 | 0.20 |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |
| 11 | Uncertainty of network analyzer | Manufacturer’s uncertainty calculator, covers NA setup | 0.5 | Normal | 2 | 1 | 0.25 |
| 12 | Mismatch of transmitter chain | Taken in to account in NA setup uncertainty | 0 | U-shaped | 1.41 | 1 | 0 |
| 13 | Insertion loss of transmitter chain | Systematic with Stage 1 (=> cancels) | 0 | Rectangular | 1.73 | 1 | 0 |
| 14 | Mismatch in the connection of calibration antenna | Taken into account in NA setup uncertainty | 0 | Rectangular | 1.73 | 1 | 0 |
| 15 | Influence of the feed cable of the calibration antenna | Gain calibration with dipole | 0.3 | Rectangular | 1.73 | 1 | 0.17 |
| 16 | Influence of the fixed measurement antenna cable | Systematic with Stage 1 (=> cancels) | 0 | Rectangular | 1.73 | 1 | 0 |
| 17 | Uncertainty of the absolute gain of the calibration antenna | Calibration certificate | [0.58] | Normal | 2 | 1 | [0.29] |
| 18 | Quality of Spatial Uniformity | Statistics of chamber | 0.5 | Gaussian | 1 | 1 | 0.5 |
| Combined standard uncertainty |  |  |  |  |  |  | [1.24] |
| Expanded uncertainty (Confidence interval of 95 %) |  |  |  |  |  |  | [2.43] |

Annex C:
Environmental requirements

# C.1	Scope

The requirements in this clause apply to all types of UE(s) in FR1 for SA and EN-DC mode.

# C.2	Ambient temperature

All the test cases defined in this technical report should be measured in room temperature e.g. 18C - 28C.

# C.3	Operating voltage

For FR1 TRP TRS, test cases shall be performed with the DUT operated in stand-alone battery powered mode. It is preferable if the UE is fully charged in the beginning of the test.

Annex D:
Phantom Definition

# D.1	Head Phantom

The basic head phantom is based on the “SAM” head phantom in IEEE Std 1528-2003, which is also described in TS 37.544 Annex A.2 [17]. For TRP TRS test, the IEEE SAM head model has been extended with the neck region, which is specified in CTIA Certification OTA Test Plan [14].

The Head phantom defined in CTIA Certification OTA Test Plan section C.2 [14], is used for FR1 TRP TRS testing in this technical report.

Figure D.1-1 Head Phantom (© 2001 – 2022 CTIA Certification. Reproduced with permission.), defined in the CTIA Certification OTA Test Plan

# D.2	Hand Phantom

## D.2.1	PDA Grip Hand

The PDA Grip Hand described in TS 37.544 Annex A.2.2 [17], which is identical to that defined in CTIA Certification OTA Test Plan section C.4.8 [14], is adopted for TRP TRS testing for the UE with width ≥56mm and ≤72mm.

## D.2.2	Wide Grip Hand

The Wide Grip hand defined in CTIA Certification OTA Test Plan section C.4.9 [14], is used for FR1 TRP TRS testing for UE with Width >72mm and ≤92mm in this technical report.

Figure D.2.2-1 Wide Grip hand (© 2001 – 2022 CTIA Certification. Reproduced with permission.), defined in the CTIA Certification OTA Test Plan

# D.3	Forearm Phantom

The Forearm phantom defined in CTIA Certification OTA Test Plan 01.72 section 2.3 [25], is used for FR1 TRP TRS testing for wrist-worn RedCap UE in this technical report.

Figure D.3-1: Forearm Phantom with Target Test Position, (© 2001 – 2022 CTIA Certification. Reproduced with permission.), defined in the CTIA Certification OTA Test Plan

Annex E:
Harmonization outcome of Alternative method and Reference method

# E.1	General

This clause records the lab alignment activity of RC test method and harmonization activity between RC and reference AC test method.

# E.2	Rel-18 Harmonization activity

## E.2.1	Framework and workplan

This clause defines the working procedure on how to proceed the Rel-18 RC lab alignment and harmonization activity. The purpose of RC lab alignment activity is to ensure there is no unexpected lab deviation among different RC systems. The purpose of RC harmonization is for comparison between RC results with reference AC results to demonstrate the gap between different methodologies.

Rel-18 RC lab alignment and RC harmonization activity are performed and analysed based on RC measurement results  in Annex E.2.2 and the AC lab alignment results in Annex F.2 and F.3.

RC Test labs are invited to participate to the lab alignment and harmonization test campaign, the following conditions should be fulfilled:

-	Participating lab should have Reverberation chamber(s) ready to support testing based on latest version of 3GPP TR 38.870.

-	Participating lab should have sufficient test resource to provide the on-time measurement results without delay.

Test methodology:

-	Test plan: latest version of 3GPP TR 38.870;

Test cases for RC harmonization Campaign:

-	Test bands: n28, n78; one low and one high band is sufficient

-	Number of test cases per band for each scenario: left and right at low/mid/high channel, total 6 test cases;

-	Use scenario: Both Head and Hand phantom (Talk mode, BHHL and BHHR) and Hand phantom only (Browsing mode, HL and HR);

-	Hand Phantom: Wide Grip hand

-	Operation mode: NR Standalone (SA)

-	Number of Tx chain: UE with 1Tx.

Harmonization Devices:

-	Same devices as Rel-18 AC lab alignment activity, listed in Annex F.3.1

Test results submitting:

-	Using the same worksheet template in [R4-2313893] to submit the measurement results

-	The measurement results of LADs should be submitted to RAN4 by anonymous approach (the UE model should not be disclosed)

-	Results shall not be shared between labs before submitting to RAN4 meetings or sharing in the RAN4 reflector. Comparison and lab alignment analysis should only be done in RAN4 meetings/discussions

RC lab alignment criteria:

-	The pass/fail criteria are defined as the maximum deviation between the measurement result and the RC reference value

-	Confirm the RC reference value derived based on the per-band per-PC averaging approach (linear average with dBm) of RC measurement data pool.

-	Apparent outliers will not be considered in averaging process. The value deviates over 1.5*MU from all the other lab’s results should be identified as apparent outlier.

-	Pass/fail limit for RC lab alignment should be defined as 0.75*MU as baseline. MU value is the preliminary expanded MU for talk mode and/or browsing mode, defined in Annex B.4.2 and B.5.2.

RC vs AC test methods harmonization criteria:

-	The pass/fail criteria are defined as the comparison between each RC results with reference AC value

-	Average of AC value per-band per-device in R4-2404644 should be reference value

-	Pass fail limits as 0.75*Total harmonization MU

-	The harmonization MU value is equal to AC MU+0.2dB

Test lab procedures:

-	LADs delivery scheme

-	LAD delivery scheme is decided after the confirmation of all the RC volunteer labs.

-	LAD delivery will also consider the parallel Rel-18 AC lab alignment activity to ensure efficiency

-	LAD measurement time in each test lab: finalize LAD measurement within [5] workdays, and deliver to next lab ASAP with LAD delivery In/Out information shared in reflector.

-	Encourage test labs to share resulting combined MU based on their own systems

## E.2.2	Measurement results

RAN4 use the same LADs in Annex F.3.1 for RC harmonization to ensure the measurement results comparison between AC and RC are based on the same set of devices.

The measurement results from RC test labs are summarized as follows:

Figure E.2.2-1: Rel-18 RC Talk mode measurement results (no n28 results from Lab9)

Figure E.2.2-2: Rel-18 RC Browsing mode measurement results (no n28 results from Lab9)

Based on the measured results, the RC lab alignment is analysed and the outcome is as following:

Figure E.2.2-3: Rel-18 Talk mode RC lab alignment analysis, deviation between each test lab and RC reference value (no n28 results from Lab9)

Figure E.2.2-4: Rel-18 Browsing mode RC lab alignment analysis, deviation between each test lab and RC reference value (no n28 results from Lab9)

Based on above analysis in Figure E.2.2-3 and E.2.2-4, the following outcome for RC lab alignment can be observed:

-	For n78 and n28 TRP, the RC reference value is derived by the agreed linear average (with dBm) of all the available results, the maximum deviation between test labs and RC reference value is 1.47dB for BHH TRP and 1.18dB for BH TRP.

-	For n78 and n28 TRS, the RC reference value is derived by the agreed linear average (with dBm) of all the available results, the maximum deviation between test labs and reference value 1.70dB for BHH TRS and 1.58dB for BH TRS.

Furthermore, the comparison of RC vs AC is also analysed as following:

Figure E.3.2-5: Rel-18 Talk mode RC harmonization analysis, deviation between each test lab and AC reference value (no n28 talk mode results from Lab9)

Figure E.3.2-6: Rel-18 Browsing mode RC harmonization analysis, deviation between each test lab and AC reference value (no n28 AC reference value)

Based on above analysis in Figure E.3.2-5 and E.3.2-6, the following outcome for RC harmonization can be observed:

-	For Talk mode TRP, the maximum deviation between each RC test lab and AC reference value is 1.03dB for n78 and 1.45dB for n28

-	For Browsing mode TRP, the maximum deviation between each RC test lab and AC reference value is 0.89dB for n78

-	For Talk mode TRS, the maximum deviation between each RC test lab and AC reference value is 1.62dB for n78 and 3.72dB for n28

-	For Browsing mode TRS, the maximum deviation between each RC test lab and AC reference value is 1.46dB for n78

Additional observation is that there is a consistency trend for n28 talk mode TRS, all results (both LAD3 and LAD4) from RC labs are worse than AC reference value.

## E.2.3	Pass/fail limit

For RC lab alignment:

For RC lab alignment, the reference value is defined as average value of RC value per-band per-device from all labs in R4-2404644. Pass fail limits is defined as 0.75*RC MU as following.

-	TRP: 2.08 dB for talk mode, and 1.90 dB for browsing mode (all FR1 bands)

-	TRS: 2.43 dB for talk mode, and 2.28 dB for browsing mode (all FR1 bands)

For RC harmonization:

For RC harmonization, the reference value is defined as average value of AC value per-band per-device from all AC labs in R4-2404644. Pass fail limits is defined as 0.75*Total harmonization MU.

Based on the MU assessment of AC and RC test method in Annex B, the Total harmonization MU is defined as AC MU+0.2dB. The final pass/fail limits for Rel-18 RC harmonization are specified as following:

-	Talk mode TRP: 1.62 dB for bands >3GHz, and 1.56 dB for bands <3GHz

-	Browsing mode TRP: 1.45 dB for bands >3GHz, and 1.38 dB for bands <3GHz

-	Talk mode TRS: 1.88 dB for bands >3GHz, and 1.84 dB for bands <3GHz

-	Browsing mode TRS: 1.75 dB for bands >3GHz, and 1.70 dB for bands <3GHz

## E.2.4	Conclusions

For RC lab alignment:

Based on the analysis in Annex E.2.2 and pass fail limits defined in Annex E.2.3, 3GPP Rel-18 RC lab alignment activity can be successfully concluded as following:

-	All the 6 labs with Reverb chamber system are well aligned at FR1 high band

-	All 5 labs with available n28 measurement results from Reverb chamber system are well aligned at FR1 low band

For RC harmonization:

Based on the analysis in Annex E.2.2 and harmonization pass fail limits defined in Annex E.2.3, 3GPP Rel-18 RC harmonization activity can be successfully concluded as following:

-	RC test method can be harmonized with reference AC test method at FR1 band >3GHz for TRP and TRS measurements.

-	RC test method harmonization with reference AC test method at FR1 FDD bands below 3GHz is not confirmed.

Annex F:
Lab alignment outcome of Reference method

< Editor’s note: Detailed structure of the subclause is TBD. Framework, pass/fail limits, and final conclusions>

# F.1	General

This clause records the lab alignment activity of AC test method. The purpose of lab alignment campaign is to ensure there is no unexpected lab deviation and establish full trust and confidence on the measurement results, which is an essential action to define TRP TRS requirements based on measurement data.

The test campaign is performed among volunteered test labs those are accredited under ISO 17025 (ISO 17025 accredited labs) and have any of 3GPP TS 37.544, CCSA YD/T 1484, and CTIA OTA Test Plan listed on its accreditation scope.

# F.2	Rel-17 lab alignment campaign (Browsing mode)

## F.2.1	Framework and workplan

This clause defines the working procedure on how to proceed the lab alignment campaign in Rel-17. The purpose of Lab Alignment Campaign is to ensure there is no unexpected lab deviation and establish full trust and confidence on the results.

Test labs are invited to participate to the lab alignment and test campaign, the following conditions should be fulfilled:

-	Participating lab should be accredited under ISO 17025 (ISO 17025 accredited labs) and have any of 3GPP TS 37.544, CCSA YD/T 1484.6, and CTIA OTA Test Plan listed on its accreditation scope.

-	Participating lab should have anechoic chamber(s) ready to support testing based on 3GPP TR 38.834. Participating lab should have sufficient test resource to provide the on-time measurement results without delay.

-	Other test methodology besides the methodology captured in existing TR 38.834 is not precluded and can be considered as long as it can meet the endorsed timeline based on the input from companies.

Test methodology:

-	Test plan: 3GPP TR 38.834;

Test cases for Lab Alignment Campaign:

-	Test bands: n41 and n78;

-	Number of test cases: maximum 3 devices per-band

-	Use scenario: Hand phantom only (Browsing mode), i.e., Hand Left and Hand Right

-	Hand Phantom: Corresponding phantom depends on UE size

-	Operation mode: NR Standalone (SA)

Lab Alignment Device (LAD) selection criteria:

-	Smartphone DUT size: two sizes, both width >72mm and ≤92mm, and width ≥56mm and ≤72mm;

-	DUT capability: support for all the Bands n41, n28, n78, and n79 those listed in the WID is preferred, but devices supporting only a subset of the above bands can equally be used in the lab alignment campaign.

-	Intended for which market: no limitation

-	Tx Antenna switching: if the DUT support TAS, the LAD provider should also provide the software/method to lock the UE primary antenna, or the primary antenna has already been locked before submitting to test lab and kept unchanged during whole alignment campaign

-	Power Class: PC2 (n41 and n78)

-	TxD is not allowed

-	For LAD selection: all application will be first come first served;

-	For each device, all the supported bands information should be shared

Test results submitting:

-	Using the same worksheet template in R4-2207327 to submit the measurement results

-	The measurement results should be submitted to RAN4 by anonymous approach (the UE model should not be disclosed)

-	Results shall not be shared between labs before submitting to RAN4 meetings or sharing in the RAN4 reflector. Comparison and lab alignment analysis should only be done in RAN4 meetings/discussions

Lab alignment criteria:

-	The pass/fail criteria are defined as the maximum deviation between the measurement result and the reference value

-	Confirm the reference value derived based on the per-band per-PC averaging approach (linear average with dBm) of lab alignment data pool from ≥3 labs submitted before 16th May 2022 as baseline.

-	If the unfinished volunteer labs submit LAD measurement results before RAN4#104 meeting and the results are not identified as apparent outliers, consider to update the reference values as the average of the LAD measurement results from all the labs.

-	RAN4 allows the unfinished volunteer labs to submit LAD measurement results after RAN4#103-e meeting, before RAN4#104-e meeting.

-	Labs who submitted data to RAN4#103-e are confirmed as the aligned labs according to the currently agreed reference values and pass/fail limits.

-	Apparent outliers will not be considered in averaging process. The value deviates over 1.5*MU from all the other lab’s results should be identified as apparent outlier.

-	Pass/fail limit for lab alignment should be defined as ±0.75*MU (i.e., 1.34dB for TRP, and 1.65dB for TRS) as baseline. MU value is the expanded MU in TR38.834, i.e., 1.78dB for TRP and 2.20dB for TRS.

-	The pass/fail limit and reference value shall be considered together if further update identified based on more data input

-	The summation form for TRP and TRS lab alignment should keep consistent during the calculation process of TRP TRS lab alignment from each company, i.e., sin weights approach or Clenshaw-Curtis quadrature integral approximation. Only traditional approach should be used during lab alignment activity to reduce unnecessary uncertainty.

-	How to treat late submission results and confirm the alignment:

-	The reference value and pass/fail limit should be defined in RAN4#103-e meeting. The late submission deadline for phase 2 is Tdoc submission deadline of RAN4#104-e meeting.

-	An offline RAN4 conference call may be needed to review lab alignment data to tentatively confirm the pass or fail of late-submitted labs. Final endorsement will happen during RAN4#104e.

-	Only measurement results from aligned labs in RAN4#104-e meeting are considered as TRP TRS Performance data pool to define final requirements.

Test lab procedures:

-	LAD delivery scheme

-	Decide LAD delivery scheme after all the test lab and LAD information being confirmed.

-	The available LADs can be split among labs to multiplex the testing effort

-	LAD measurement time in each test lab: finalize LAD measurement within 5 workdays, and deliver to next lab ASAP with LAD delivery In/Out information shared in reflector.

-	Encourage test labs to share resulting combined MU based on their own systems

## F.2.2	Measurement results

Considering the test burden and urgent timeline, RAN4 decided to select two smartphones (named as LADs, Lab Alignment Devices) as reference DUTs to perform Rel-17 AC lab alignment activity. The devices’ size belongs to width >72mm and ≤92mm.

In Rel-17, the lab alignment campaign is focused on Browsing mode (Hand phantoms only) at band n41 and n78. There are 8 test labs participate the lab alignment campaign, the reference value is derived based on linear average (with dBm) of all the 8 labs results. The summary of the lab alignment analysis outcome is shown in Figure F.2.2-1 for TRP, and Figure F.2.2-2 for TRS.

Figure F.2.2-1: NR FR1 TRP lab alignment analysis, deviation between each test lab and reference value

Figure F.2.2-2: NR FR1 TRS lab alignment analysis, deviation between each test lab and reference value

The maximum deviation between test labs and reference value is 1.08dB for TRP and 0.93dB for TRS.

## F.2.3	Pass/fail limit

Based on the preliminary MU assessment of AC test method with browsing mode in Annex B, and lab alignment measurement results in Annex F.2.2, RAN4 decided the final pass/fail limits for Rel-17 FR1 TRP TRS lab alignment activity (Browsing mode) as following:

-	TRP: 1.2 dB

-	TRS: 1.5 dB

## F.2.4	Conclusions

RAN4 conclude the successful Rel-17 FR1 TRP TRS lab alignment activity, all the 8 test labs with anechoic chamber system are well aligned. Rel-17 TRP TRS requirements are specified based on the measurement results submitted by above aligned test labs.

# F.3	Rel-18 lab alignment (Talk mode)

## F.3.1	Framework and workplan

This clause defines the working procedure on how to proceed the lab alignment campaign in Rel-18. The purpose of Lab Alignment Campaign is to ensure there is no unexpected lab deviation and establish full trust and confidence on the results.

Rel-18 lab alignment activity is performed with Anechoic Chamber based test method.

Test labs are invited to participate to the lab alignment and test campaign, the following conditions should be fulfilled:

-	Participating lab should be accredited under ISO 17025 (ISO 17025 accredited labs) and have any of 3GPP TS 37.544, CCSA YD/T 1484.6, and CTIA OTA Test Plan listed on its accreditation scope.

-	Participating lab should have anechoic chamber(s) ready to support testing based on 3GPP TS 38.161.

-	Participating lab should have sufficient test resource to provide the on-time measurement results without delay.

Test methodology:

-	Test plan: 3GPP TS 38.161;

Test cases for Lab Alignment Campaign:

-	Test bands: n28, n78; one low and one high band is sufficient

-	Number of test cases per band: BHHL and BHHR at low/mid/high channel;

-	Use scenario: Head and Hand phantom (Talk mode), i.e., BHHL and BHHR

-	Hand Phantom: Wide Grip hand

-	Operation mode: NR Standalone (SA)

-	Number of Tx chain: UE configured with 1Tx

Lab Alignment Device (LAD): in total 4 devices

-	Rel-17 LADs (LAD1 and LAD2) for band n78;

-	2 more devices (LAD 3 and LAD 4) for band n28;

-	UE information listed in Table F.3.1-1

Test results submitting:

-	Using the same worksheet template in [R4-2313892] to submit the measurement results

-	The measurement results of LADs should be submitted to RAN4 by anonymous approach (the UE model should not be disclosed)

-	Results shall not be shared between labs before submitting to RAN4 meetings or sharing in the RAN4 reflector. Comparison and lab alignment analysis should only be done in RAN4 meetings/discussions

-	Sharing of the measurements results of a specific device from test labs is permitted only towards the 3GPP member that provided such device (i.e., device’s owner) before submitting to RAN4

Lab alignment criteria:

-	The pass/fail criteria are defined as the maximum deviation between the measurement result and the reference value

-	Confirm the reference value derived based on the per-band per-PC averaging approach (linear average with dBm) of lab alignment data pool.

-	Apparent outliers will not be considered in averaging process. The value deviates over 1.5*MU from all the other lab’s results should be identified as apparent outlier.

-	Pass/fail limit for lab alignment should be defined as 0.75*MU as baseline. MU value is the expanded MU for BHH (defined in Annex of TR38.870).

-	The summation form for TRP and TRS lab alignment should keep consistent during the calculation process of TRP TRS lab alignment from each company i.e. sin weights approach. Only traditional approach (15-degrees TRP and 30-degrees TRS) should be used during lab alignment activity to reduce unnecessary uncertainty.

-	The AC alignment activity is finally concluded in RAN4#110bis meeting.

Test lab procedures:

-	LAD delivery scheme

-	Decide LAD delivery scheme after all the test lab and LAD information being confirmed (after the confirmation of volunteered labs and LADs).

-	LAD measurement time in each test lab: finalize LAD measurement within [5] workdays and deliver to next lab ASAP with LAD delivery In/Out information shared in reflector.

-	Encourage test labs to share resulting combined MU based on their own systems.

Table F.3.1-1: Rel-18 LAD information

| Rel-18 Lab Alignment Devices (LADs) |  |  |  |
| --- | --- | --- | --- |
| DUT | Provider/Owner | Contact | Note |
| LAD1 (same device used in Rel-17) | vivo | Ruixin Wang ruixin. wang@vivo.com | For n78 testing. |
| LAD2 (same device used in Rel-17) | vivo | Ruixin Wang ruixin. wang@vivo.com | For n78 testing. |
| LAD3 (new device for Rel-18) | vivo | Ruixin Wang ruixin. wang@vivo.com | For n28 testing. |
| LAD4 (new device for Rel-18) | vivo | Ruixin Wang ruixin. wang@vivo.com | For n28 testing. |
| Note: To keep the test burden, most 4 test cases, i.e., two devices per band. |  |  |  |

## F.3.2	Measurement results

Considering the test burden and urgent timeline, RAN4 decided to select two smartphones per band (in total 4, named as LADs, Lab Alignment Devices) as reference DUTs to perform Rel-18 AC lab alignment activity. The devices’ size belongs to width >72mm and ≤92mm.

In Rel-18, the lab alignment campaign is focused on Talk mode (Head and Hand phantoms) at one low band and one high band (i.e., band n28 and n78). There are 8 test labs participate the lab alignment campaign, the reference value is derived based on linear average (with dBm) of all the 8 labs results. The summary of the lab alignment analysis outcome is shown in Figure F.3.2-1 for TRP, and Figure F.3.2-2 for TRS.

Figure F.3.2-1: Rel-18 NR FR1 TRP lab alignment analysis, deviation between each test lab and reference value

Figure F.3.2-2: Rel-18 NR FR1 TRS lab alignment analysis, deviation between each test lab and reference value

The maximum deviation between test labs and reference value is 1.22 dB for TRP and 1.55 dB for TRS.

## F.3.3	Pass/fail limit

Based on the MU assessment of AC test method with talk mode in Annex B, and lab alignment measurement results in Annex F.3.2, RAN4 decided the final pass/fail limits for Rel-18 FR1 TRP TRS lab alignment activity (Talk mode) as following:

-	TRP: 1.47dB for n78, and 1.41dB for n28

-	TRS: 1.73dB for n78 and 1.69 dB for n28

## F.3.4	Conclusions

RAN4 conclude the successful Rel-18 FR1 TRP TRS lab alignment activity, all the 8 test labs with anechoic chamber system are well aligned at NR FR1. Rel-18 TRP TRS requirements are specified based on the measurement results submitted by above aligned test labs.

Annex G (informative):
Change history

| Change history |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Date | Meeting | TDoc | CR | Rev | Cat | Subject/Comment | New version |
| 2022-10 | RAN4#104bis-e | R4-2216104 |  |  |  | Initial Skeleton | 0.0.1 |
| 2022-11 | RAN4#105 | R4-2218845 |  |  |  | R4-2218850 TP to TR 38.870 on supplement of basic AC test method | 0.1.0 |
| 2023-03 | RAN4#106 | R4-2301555 |  |  |  | R4-2301556 TP to TR 38.870 on RedCap OTA test parameters R4-2302983 TP to TR 38.870 Reverberation Chamber Method (Alternate Method) R4-2302498 TP to TR 38.870 on contents for Annex B | 0.2.0 |
| 2023-04 | RAN4#106bis-e | R4-2305102 |  |  |  | R4-2304034 Text proposal for Annex C on environment in TR 38.870 R4-2305908 TP to TR 38.870 on General parts R4-2305785 TP on Test reductions with Measurement Grids | 0.3.0 |
| 2023-05 | RAN4#107 | R4-2308250 |  |  |  | R4-2308251 TP to TR 38.870 on supplement of coordinate system and phantom definition R4-2309814 TP for CBW and spatial uniformity test procedures in TR38.870 R4-2308824 TP on Measurement Uncertainty Associated with Grid Size R4-2309289 TP to TR 38.870 on contents for Annex B | 0.4.0 |
| 2023-08 | RAN4#108 | R4-2312565 |  |  |  | R4-2311673 TP to TR 38.870 on contents for Annex B R4-2313988 TP to TR 38.870 on AC lab alignment campaign | 0.5.0 |
| 2023-10 | RAN4#108bis | R4-2315838 |  |  |  | R4-2315839 TP to TR 38.870 on phantoms R4-2317003 TP to TR 38.870 on 2Tx test configuration | 0.6.0 |
| 2023-11 | RAN4#109 | R4-2318965 |  |  |  | R4-2321097 TP to TR 38.870 on TRP TRS test procedure for CA R4-2321098 TP to TR 38.870 on TRP TRS test method R4-2320707 TP to TR 38.870 on contents for Annex B | 0.7.0 |
| 2023-12 | RAN#102 | RP-233136 |  |  |  | For one-step approval | 1.0.0 |
| 2023-12 | RAN#102 | RP-234063 |  |  |  | Approved with additional editor note | 1.1.0 |

| Change history |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Date | Meeting | TDoc | CR | Rev | Cat | Subject/Comment | New version |
| 2023-12 | RAN#102 |  |  |  |  | Approved by plenary – Rel-18 spec under change control | 18.0.0 |
| 2024-03 | RAN#103 | RP-240608 | 0002 |  | F | CR to TR 38.870 on TRP TRS test method | 18.1.0 |
| 2024-06 | RAN#104 | RP-241487 | 0005 | 2 | F | CR to TR38.870 on UL MIMO radiated output power metric | 18.2.0 |
| 2024-06 | RAN#104 | RP-241441 | 0006 | 1 | F | CR to TR38.870 on CBW scaling of the TRS requirement | 18.2.0 |
| 2024-06 | RAN#104 | RP-241487 | 0008 | 1 | B | CR to TR 38.870 to add CA combinations to section 4.3.5 and editorial corrections | 18.2.0 |
| 2024-06 | RAN#104 | RP-241487 | 0009 |  | F | CR to TR 38.870 on Rel-18 AC lab alignment activity outcome | 18.2.0 |
| 2024-06 | RAN#104 | RP-241487 | 0010 | 2 | F | CR to TR 38.870 on Rel-18 RC lab alignment and harmonization outcome | 18.2.0 |
| 2024-06 | RAN#104 | RP-241487 | 0011 |  | F | CR to TR 38.870 to update RC TRS definition in 5.22 and 8.6.1.1 | 18.2.0 |
| 2024-06 | RAN#104 | RP-241487 | 0012 |  | F | Editorial CR to TR 38.870 | 18.2.0 |