3GPP TSG-RAN5 Meeting #103	R5-243312

# Fukuoka City, Fukuoka , Japan, 20th – 24th May, 2024

| CR-Form-v12.3 |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CHANGE REQUEST |  |  |  |  |  |  |  |  |
|  | 38.551 | CR | 0015 | rev | - | Current version: | 17.1.0 |  |
| For HELP on using this form: comprehensive instructions can be found at http://www.3gpp.org/Change-Requests. |  |  |  |  |  |  |  |  |

| Proposed change affects: | UICC apps |  | ME |  | Radio Access Network |  | Core Network |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |

| Title: | Clarification of Calibration and Ripple Test Procedures |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Source to WG: | Keysight Technologies UK Ltd, ETS-Lindgren |  |  |  |  |  |  |  |  |  |
| Source to TSG: | R5 |  |  |  |  |  |  |  |  |  |
| Work item code: | TEI17_Test, NR_MIMO_OTA-UEConTest |  |  |  |  |  | Date: |  |  | 2024-04-28 |
| Category: | F |  |  |  |  |  | Release: |  |  | Rel-17 |
|  | Use one of the following categories: F (correction) A (mirror corresponding to a change in an earlier release) B (addition of feature), C (functional modification of feature) D (editorial modification) Detailed explanations of the above categories can be found in 3GPP TR 21.900. |  |  |  |  |  |  |  | Use one of the following releases: Rel-8 (Release 8) Rel-9 (Release 9) Rel-10 (Release 10) Rel-11 (Release 11) … Rel-17 (Release 17) Rel-18 (Release 18) Rel-19 (Release 19) Rel-20 (Release 20) |  |
| Reason for change: |  | The need to repeat the calibration and ripple test procedures were missing. |  |  |  |  |  |  |  |  |
| Summary of change: |  | Added clarifications on the need to repeat calibration and ripple test procedures. |  |  |  |  |  |  |  |  |
| Consequences if not approved: |  | Measurement results could be invalid. |  |  |  |  |  |  |  |  |
| Clauses affected: |  | A.2.2, B.2.3 |  |  |  |  |  |  |  |  |
|  |  | Y | N |  |  |  |  |  |  |  |
| Other specs |  |  | X | Other core specifications |  |  |  | TS/TR ... CR ... |  |  |
| affected: |  |  | X | Test specifications |  |  |  | TS/TR ... CR ... |  |  |
| (show related CRs) |  |  | X | O&M Specifications |  |  |  | TS/TR ... CR ... |  |  |
| Other comments: |  |  |  |  |  |  |  |  |  |  |
| This CR's revision history: |  |  |  |  |  |  |  |  |  |  |

<<< Skip Unchanged Sections >>>

<<< START OF CHANGES >>>

## A.2.2	Calibration procedure

The system needs to be calibrated by using a reference calibration antenna with known gain values in order to ensure that the downlink signal power is correct. In Non-Standalone (NSA) mode, the LTE link antenna provides a stable LTE signal without precise path loss or polarization control.

Unlike traditional TRP/TRS testing where the path loss corrections can all be applied as a post processing step to the measured data, the path loss for each probe in the MPAC system has to be balanced at test time in order to generate the desired channel model environment within the test zone of the chamber. The imbalance of each path during testing would result in an alteration of the angular dependence of the channel model (i.e. varied characteristics of generated channel model) within the test zone of the chamber:

1)	Place a vertical reference dipole in the centre of the test zone, connected to a VNA port, with the other VNA port connected to the input of the channel emulator unit.

2)	Configure the channel emulator for bypass mode.

3)	Measure the response of each path from each vertical polarization probe to the reference antenna in the centre of test zone.

4)	Adjust the power on all vertical polarization branches of the channel emulator so that the powers received at the centre are equal.

5)	Repeat the steps 1 to 4 with the magnetic loop or horizontally polarized reference dipole instead and adjust the horizontal polarization branches of the channel emulator.

6)	The worst-case path loss becomes the reference path loss of the entire system; this loss is used to compute the power in the centre of the test zone relative to the output power of the Base Station simulator. Besides, based on the reference path loss, the relative offset of each path loss shall be corrected.

NOTE:	Calibration based on other antennas, e.g. horn antennas is not precluded.

This procedure shall be repeated on a regular basis, i.e., at least one per year, and any time the conducted and/or radiated signal paths have been affected, e.g., change in propagation conditions, replacement of cables, change in measurement antennas/probes, etc. It is recommended to perform this procedure rather frequently, e.g., monthly, to assess system stability.

<<< Skip Unchanged Sections >>>

## B.2.3	Quality of quiet zone

The quality of the quiet zone procedure characterizes the quiet zone performance of the anechoic chamber, specifically the effect of reflections within the anechoic chamber including any positioners and support structures. For FR1 quality of quiet zone measurements, reference antennas of sleeve dipole or magnetic loop are used. For more information, see clauses A.3.7 and A.4.2.8 of 3GPP TS 38.561 [8] with the exception that only the phi-axis ripple test is performed for MIMO OTA.

The ripple test shall be repeated when the RF/propagation conditions inside the chamber have changed, e.g., the chamber has been disassembled and reassembled, portions of the absorber been replaced, measurement antennas/probes been replaced, positioning system been replaced, etc.

<<< END OF CHANGES >>>