3GPP TSG-RAN WG4 Meeting #104-e	R4-2212817
Electronic Meeting, August 15 – August 26, 2022

Agenda item:	9.2.3.1

Source:	vivo

Title:	Outcome of 3GPP TRP TRS lab alignment activity

Document for:	Approval

1	Introduction

The phase 1 lab alignment activity for FR1 TRP TRS was concluded in RAN4#103-e meeting [1], it was also agreed that phase 2 labs should be finalized and concluded in RAN4#104-e meeting [2]. All the 3 labs for phase 2 alignment have submitted the measurement results [3-5].

With the submitted results from phase 1 [6-10], this contribution presents the analysis of the full sets of LAD measurement data (phase 1 and phase 2) and proposes the conclusions for the TRP TRS lab alignment activity.

2	Discussion

In the latest approved working procedure for TRP TRS requirement development [2], it was agreed that the lab alignment activity should be confirmed in RAN4#104-e meeting:

- Lab alignment criteria:

  - The pass/fail criteria are defined as the maximum deviation between the measurement result and the reference value

  - Confirm the reference value derived based on the per-band per-PC averaging approach (linear average with dBm) of lab alignment data pool from ≥3 labs submitted before 16th May 2022 as baseline.

    - If the unfinished volunteer labs submit LAD measurement results before RAN4#104 meeting and the results are not identified as apparent outliers, consider to update the reference values as the average of the LAD measurement results from all the labs.

    - RAN4 allows the unfinished volunteer labs to submit LAD measurement results after RAN4#103-e meeting, before RAN4#104-e meeting.

    - Labs who submitted data to RAN4#103-e are confirmed as the aligned labs according to the currently agreed reference values and pass/fail limits.

  - Apparent outliers will not be considered in averaging process. The value deviates over 1.5*MU from all the other lab’s results should be identified as apparent outlier.

  - Pass/fail limit for lab alignment should be defined as ±0.75*MU (i.e. 1.34dB for TRP, and 1.65dB for TRS) as baseline. MU value is the expanded MU in TR38.834, i.e. 1.78dB for TRP and 2.20dB for TRS.

    - The pass/fail limit and reference value shall be considered together if further update identified based on more data input

  - The summation form for TRP and TRS lab alignment should keep consistent during the calculation process of TRP TRS lab alignment from each company, i.e. sin weights approach or Clenshaw-Curtis quadrature integral approximation. Only traditional approach should be used during lab alignment activity to reduce unnecessary uncertainty.

  - How to treat late submission results and confirm the alignment:

    - The reference value and pass/fail limit should be defined in RAN4#103-e meeting. The late submission deadline for phase 2 is Tdoc submission deadline of RAN4#104-e meeting.

    - An offline RAN4 conference call may be needed to review lab alignment data to tentatively confirm the pass or fail of late-submitted labs. Final endorsement will happen during RAN4#104e.

    - Only measurement results from aligned labs in RAN4#104-e meeting are considered as TRP TRS Performance data pool to define final requirements.

Based on the above working procedure, the full sets of lab alignment results [3-10] have been analysed and summarized. The reference value is derived based on linear average of all the 8 labs results. The detailed analysis file is in the Excel worksheet of this contribution.

The summary of the key alignment analysis outcome is shown in Figure 1 for TRP, and Figure 2 for TRS.

Figure 1: NR FR1 TRP lab alignment analysis, deviation between each test lab and reference value

Figure 2: NR FR1 TRS lab alignment analysis, deviation between each test lab and reference value

Observation 1: The reference value is derived by the agreed linear average (with dBm) of all the 8 labs results, the maximum deviation between test labs and reference value is 1.08dB for TRP and 0.93dB for TRS.

Observation 2: The maximum deviation between test labs and reference value are within the agreed pass/fail limits (i.e., 0.75*MU, TRP 1.34dB, TRS 1.65dB).

With above analysis and observations, we can see that the performance deviation between lab and reference value is quite small, and the lab alignment activity presents a good harmonized performance of FR1 TRP TRS anechoic chamber systems.

With above analysis, the following proposal can be made as the 3GPP FR1 TRP TRS lab alignment outcome:

Proposal 1: 3GPP Rel-17 FR1 TRP TRS lab alignment activity can be successfully concluded and all the 8 labs (vivo, CAICT, CMCC, SRTC, OPPO, Sporton, Huawei, Element) with anechoic chamber system are well aligned.

3 Conclusion

Based on the submitted results in [3-10] and the analysis in this paper, the following outcome of 3GPP FR1 TRP TRS lab alignment activity is proposed:

Observation 1: The reference value is derived by the agreed linear average (with dBm) of all the 8 labs results, the maximum deviation between test labs and reference value is 1.08dB for TRP and 0.93dB for TRS.

Observation 2: The maximum deviation between test labs and reference value are within the agreed pass/fail limits (i.e., 0.75*MU, TRP 1.34dB, TRS 1.65dB).

Proposal 1: 3GPP Rel-17 FR1 TRP TRS lab alignment activity can be successfully concluded and all the 8 labs (vivo, CAICT, CMCC, SRTC, OPPO, Sporton, Huawei, Element) with anechoic chamber system are well aligned.

4	References

- R4-2207326, WF on FR1 TRP TRS, vivo, RAN4#103-e

- R4-2210941, Further updated Working procedure for TRP TRS requirement development, vivo, RAN4#103-e

- R4-2211570, TRP TRS lab alignment measurement results from Huawei, Huawei, RAN4#104-e, August 15-26, 2022

- R4-2211536, 3GPP FR1 TRP/TRS Lab Alignment Measurement Results from Sporton International, Sporton, RAN4#104-e, August 15-26, 2022

- R4-2211613, Element FR1 TRP/TRS Lab Alignment Campaign Measurement Results, Element, RAN4#104-e, August 15-26, 2022

- R4-2208323, TRP TRS lab alignment measurement results from CAICT, RAN4#103-e

- R4-2208411, TRP TRS lab alignment measurement results from CMCC, RAN4#103-e

- R4-2208628, TRP TRS lab alignment measurement results from vivo, RAN4#103-e

- R4-2209382, TRP TRS lab alignment measurement results from SRTC, RAN4#103-e

- R4-2210352, OPPO LAD test results for FR1 TRP/TRS Lab Alignment Campaign, OPPO, RAN4#103-e