3GPP TSG-RAN WG4 Meeting # 106-bis                                                                    R4-2304027

Electronic Meeting, 17 - 26 April, 2023

Title: 	Proposals for spatial uniformity and CBW test procedures in RC

Source: 	EMITE, Bluetest, Huawei, HiSilicon

Agenda item:	5.16.2.2

Document for:	Agreement

# Introduction

Issue 2-1-2: How to calculate Coherence bandwidth of RC

Proposals:

  - Proposal 1: The current method proposed to validate RC isotropy or Spatial Uniformity is not sufficient for a proper characterization of a RC loaded with absorbers to achieve proper Coherent Bandwidth, which is a pre-requisite to test modulated signals. Another RC Isotropy or Spatial Uniformity validation method compliant with loaded RC shall be determined prior to move forward with RC alignment test efforts, e.g.: Standard Uncertainty on RC Transfer Function (uref) [4] in R4-2300350.

  - Proposal 2: Review and correct the isotropy and CBW test procedure proposal in Annex in R4-2300139.

Agreements:

  - Further check next meeting the above two approaches based on detailed descriptions in TP.

Issue 2-1-3: Verification procedure for Coherence bandwidth of RC

Agreements:

  - Encourage companies to provide corresponding verification procedure of coherence bandwidth based on the proposed method to calculate Coherence bandwidth in Issue 2-1-2.

  - Text Proposals are encouraged for detailed discussion and endorsement.

This contribution further discusses test issues associated with RC.

# Discussion

Test procedure for spatial uniformity in appendix 1 is based on [2] and is for discussion and/or correction before agreeing to formal text proposal.

Measurement in [4] shows that CBW of RC without absorbers is sufficient for TRP and TRS evaluation. The reason could be that the intrinsic CBW without absorbers is typically above 500kHz and is still much larger than the modulation bandwidth or SCS (Sub-Carrier Spacing) of 15kHz, 30kHz or 60kHz used in FR1. The proposed procedure in appendix 1 is therefore without absorber loading, but includes turntable inside the working volume.

Proposal 1: review and correct the spatial uniformity test procedure proposal in appendix 1 of R4-2304027.

Test procedure for CBW (Coherence Bandwidth) in appendix 2 is based on [3] and is for discussion and/or correction before agreeing to formal text proposal. Values for discussion are in square brackets.

As the spectrum shape is a sinc function in frequency domain, it would be prudent to set a minimum CBW much larger than the values of sub-carrier spacing. A factor of [5] is proposed below.

Proposal 2: the minimum CBW of RC should be [5] times of Sub-Carrier Spacing.

Proposal 3: review and correct the CBW test procedure proposal in appendix 2 of R4-2304027.

# Conclusions

This contribution makes the following proposals.

Proposal 1: review and correct the spatial uniformity test procedure proposal in appendix 1 of R4-2304027.

Proposal 2: the minimum CBW of RC should be [5] times of Sub-Carrier Spacing.

Proposal 3: review and correct the CBW test procedure proposal in appendix 2 of R4-2304027.

# References

- R4-2302917 WF for Rel-18 TRP TRS WI

- “Reverberation Chambers through the Magnifying Glass: An Overview and Classification of Performance Indicators”, Ramiro Serra, IEEE Electromagnetic Compatibility Magazine, Volume 6, Quarter 2, 2017, pp.76-88

- “Reverberation Chamber Metrology for Wireless Internet of Things Devices”, Anouk Hubrechsen, Kate A. Remley and Sara Catteau, IEEE Microwave Magazine, February 2022, pp.75-85

- R4-2301557 Reverberation chamber test methodology

# Appendix 1 Test procedure for spatial uniformity

The test steps are as follows.

- Select either a cylindrical or cubic working volume, with when there is a turn table inside or not, respectively. As shown in Figure 1 by red dots, either the 4 corner points of the imaginary rectangle that would generate the cylindrical working volume or the 8 corner points will be used as locations for measurement. This amounts to either 4 or 8 total number of measurement points.

- Measure transmission coefficient S21 at three orthogonal antenna positions in a complete mode stirring sequence.

- Calculate the power Pref,t for all 12 or 24 positions. In this way, Pref,t  is the reference power transfer function for position t of the calibration antenna.

- Calculate the average of power transfer function Pref over the calibration positions i.e., T=12 or T=24 using the following relation:

Prefmean= 1T-1 t=1TPref,t

- Calculate the standard deviation of the power transfer function over T different calibration antenna positions by

σPref =1T-1t=1TPref,t-Prefmean2

- Calculate σPref in dB by

σPrefdB=10log10( Prefmean+ σPrefPrefmean )

- Repeat steps (a) to (e) for at least [25] frequency points evenly across the NR FR1 bands.

(0,0)

(R,0)

(0,Z)

(R,Z)

- (b)

Figure 1: Illustration of a working volume with (a) for cylindrical and (b) for cubic volumes.

# Appendix 2 Test procedure for coherence bandwidth

CBW or Coherence Bandwidth is a metric to determine the correlation in frequency within a given working volume of a reverberation chamber (RC) and defined in [3] as

R(i,n)=j=1P-iS21(fj,n)S21*(fj+i,n)j=1P-iS21(fj,n)S21*(fj,n)         (1)

Where

S21 (fj, n) corresponds to the measured complex S21 at the frequency step fj with [P] frequency points measured within a given bandwidth (BW) such that f1 = fc – (BW/2) and fp = fc + (BW/2).

n is the mode stirring sample index with a total equal to [N].

i is the frequency offset index with a frequency resolution of (f1 - fp)/(P - 1). (discussion point: because typical CBW is a few MHz, frequency resolution can be selected as [50-100 kHz])

Test steps are as follows.

- Select the center frequency fc and BW. A typical measurement frequency range of [100MHz] would be used as a default.

- Determine the number of measurement points P.

- Determine the total number of stirring sequence N.

- Perform measurement and process the data using equation (1)

- Average the auto-correlation R(i, n) over a complete stirring sequence (i.e. index n) for each frequency point index i.

- Identify the coherence bandwidth corresponding to a value of [0.5] in auto-correlation R(i, n)