| 3GPP TR 38.761 V0.1.0 (2023-11) |  |
| --- | --- |
| Technical Report |  |
| 3rd Generation Partnership Project; Technical Specification Group Radio Access Network; Measurements of Multiple Input Multiple Output (MIMO) Over-the-Air (OTA) performance of User Equipment (UE); (Release 18) |  |
| The present document has been developed within the 3rd Generation Partnership Project (3GPP TM) and may be further elaborated for the purposes of 3GPP. The present document has not been subject to any approval process by the 3GPP Organizational Partners and shall not be implemented. This Specification is provided for future development work within 3GPP only. The Organizational Partners accept no liability for any use of this Specification. Specifications and Reports for implementation of the 3GPP TM system should be obtained via the 3GPP Organizational Partners' Publications Offices. |  |

3GPP Postal address 3GPP support office address 650 Route des Lucioles - Sophia Antipolis Valbonne - FRANCE Tel.: +33 4 92 94 42 00 Fax: +33 4 93 65 47 16 Internet https://www.3gpp.org

Copyright Notification No part may be reproduced except as authorized by written permission. The copyright and the foregoing restriction extend to reproduction in all media. © 2023, 3GPP Organizational Partners (ARIB, ATIS, CCSA, ETSI, TSDSI, TTA, TTC). All rights reserved. UMTS™ is a Trade Mark of ETSI registered for the benefit of its members 3GPP™ is a Trade Mark of ETSI registered for the benefit of its Members and of the 3GPP Organizational Partners LTE™ is a Trade Mark of ETSI registered for the benefit of its Members and of the 3GPP Organizational Partners GSM® and the GSM logo are registered and owned by the GSM Association

Contents

Foreword	4

1	Scope	6

2	References	6

3	Definitions of terms, symbols and abbreviations	6

3.1	Terms	6

3.2	Symbols	6

3.3	Abbreviations	7

4	General	7

5	Measurement setup	7

5.1	General	7

5.2	Measurement setup for FR1 MIMO OTA	7

5.3	Measurement setup for FR2 MIMO OTA	7

6	Channel Model Validation	7

6.1	General	7

6.2	Verification of Channel Model implementation of FR1	7

6.3	Verification of Channel Model implementation of FR2	7

7	Lab alignment of FR1 MIMO OTA	8

7.1	General	8

7.2	Noise impact in MPAC on MIMO performance	8

7.3	Lab alignment campaign for frequency bands > 1GHz	8

7.3.1	Framework	8

7.3.2	Measurement results	8

7.3.3	Pass/fail limits	8

7.3.4	Conclusion	8

7.4	Lab alignment campaign for frequency bands < 1GHz	8

7.4.1	Framework	8

7.4.2	Measurement results	8

7.4.3	Pass/fail limits	8

7.4.4	Conclusion	8

8	Lab alignment of FR2 MIMO OTA	9

8.1	Framework	9

8.2	Measurement results	9

8.3	Pass/fail limits	9

8.4	Conclusion	9

Annex <A>: Change history	9

Foreword

This Technical Report has been produced by the 3rd Generation Partnership Project (3GPP).

The contents of the present document are subject to continuing work within the TSG and may change following formal TSG approval. Should the TSG modify the contents of the present document, it will be re-released by the TSG with an identifying change of release date and an increase in version number as follows:

Version x.y.z

where:

x	the first digit:

1	presented to TSG for information;

2	presented to TSG for approval;

3	or greater indicates TSG approved document under change control.

y	the second digit is incremented for all changes of substance, i.e. technical enhancements, corrections, updates, etc.

z	the third digit is incremented when editorial only changes have been incorporated in the document.

In the present document, modal verbs have the following meanings:

shall	indicates a mandatory requirement to do something

shall not	indicates an interdiction (prohibition) to do something

The constructions "shall" and "shall not" are confined to the context of normative provisions, and do not appear in Technical Reports.

The constructions "must" and "must not" are not used as substitutes for "shall" and "shall not". Their use is avoided insofar as possible, and they are not used in a normative context except in a direct citation from an external, referenced, non-3GPP document, or so as to maintain continuity of style when extending or modifying the provisions of such a referenced document.

should	indicates a recommendation to do something

should not	indicates a recommendation not to do something

may	indicates permission to do something

need not	indicates permission not to do something

The construction "may not" is ambiguous and is not used in normative elements. The unambiguous constructions "might not" or "shall not" are used instead, depending upon the meaning intended.

can	indicates that something is possible

cannot	indicates that something is impossible

The constructions "can" and "cannot" are not substitutes for "may" and "need not".

will	indicates that something is certain or expected to happen as a result of action taken by an agency the behaviour of which is outside the scope of the present document

will not	indicates that something is certain or expected not to happen as a result of action taken by an agency the behaviour of which is outside the scope of the present document

might	indicates a likelihood that something will happen as a result of action taken by some agency the behaviour of which is outside the scope of the present document

might not	indicates a likelihood that something will not happen as a result of action taken by some agency the behaviour of which is outside the scope of the present document

In addition:

is	(or any other verb in the indicative mood) indicates a statement of fact

is not	(or any other negative verb in the indicative mood) indicates a statement of fact

The constructions "is" and "is not" do not indicate requirements.

1	Scope

The present document is a Technique Report for Measurements of Multiple Input Multiple Output (MIMO) Over-the-Air (OTA) performance of User Equipment (UE). It includes analysis and measurement results of MIMO OTA in Frequency Range 1 and Frequency Range 2, such as verification of channel models, lab alignment activities, etc..

2	References

The following documents contain provisions which, through reference in this text, constitute provisions of the present document.

-	References are either specific (identified by date of publication, edition number, version number, etc.) or non-specific.

-	For a specific reference, subsequent revisions do not apply.

-	For a non-specific reference, the latest version applies. In the case of a reference to a 3GPP document (including a GSM document), a non-specific reference implicitly refers to the latest version of that document in the same Release as the present document.

[1]	3GPP TR 21.905: "Vocabulary for 3GPP Specifications".

…

[x]	<doctype> <#>[ ([up to and including]{yyyy[-mm]|V<a[.b[.c]]>}[onwards])]: "<Title>".

3	Definitions of terms, symbols and abbreviations

This clause and its three (sub) clauses are mandatory. The contents shall be shown as "void" if the TS/TR does not define any terms, symbols, or abbreviations.

3.1	Terms

For the purposes of the present document, the terms given in TR 21.905 [1] and the following apply. A term defined in the present document takes precedence over the definition of the same term, if any, in TR 21.905 [1].

Definition format (Normal)

<defined term>: <definition>.

example: text used to clarify abstract rules by applying them literally.

3.2	Symbols

For the purposes of the present document, the following symbols apply:

Symbol format (EW)

<symbol>	<Explanation>

3.3	Abbreviations

For the purposes of the present document, the abbreviations given in TR 21.905 [1] and the following apply. An abbreviation defined in the present document takes precedence over the definition of the same abbreviation, if any, in TR 21.905 [1].

Abbreviation format (EW)

<ABBREVIATION>	<Expansion>

4	General

<Editor’s note: Detailed structure of the subclause is TBD. Testing configuration, testing bands should be included for reference.>

5	Measurement setup

5.1	General

<Editor’s note: Clause content is FFS.>

5.2	Measurement setup for FR1 MIMO OTA

<Editor’s note: Detailed structure of the subclause is TBD. Refer to TS38.151, including baseline setup of FR1 MIMO OTA testing.>

5.3	Measurement setup for FR2 MIMO OTA

<Editor’s note: Detailed structure of the subclause is TBD. Refer to TS38.151, including baseline setup of FR2 MIMO OTA testing.>

6	Channel Model Validation

## 6.1	General

This clause describes the FR1 and FR2 MIMO OTA channel model validation measurements. The purpose of channel model validation is to ensure that the channel models are correctly implemented and hence capable of generating the propagation environment, as described by the model, within the test zone.

### 6.2	Verification of Channel Model implementation of FR1

### 6.2.1	Framework

This clause describes how to proceed Channel Model Validation for FR1 MIMO OTA with MPAC system.

- The channel model validation measurements shall be performed according to the procedures in Annex C.3 of TS 38.151, including:

  - Power Delay Profile (PDP)

  - Doppler/Temporal correlation

  - Spatial correlation

  - Cross-polarization

  - Power validation

- Channel model: FR1 UMa CDL-C and UMi CDL-C, as specified in Annex C.1 of TS 38.151

- Test frequency: as specified in Tables C.3.1-1 and C.3.1-2 of TS 38.151

- Pass/fail limits: as defined in Annex C.4 of TS 38.151

### 6.2.2 Channel Model Validation Results

This clause presents channel model validation results from different labs with different types of equipment and setup vendors. All the validation results are included for comparison. Table 6.2.2-1 lists equipment and setup vendors of different labs.

Table 6.2.2-1: Equipment and setup vendors of different labs

| Labs | Channel Emulator | BS Emulator | System |
| --- | --- | --- | --- |
| Lab 1 | Vendor A | Vendor I | Vendor 1 |
| Lab 2 | Vendor B (for bands n41, n78) Vendor A (for band n28) | Vendor II | Vendor 4 |
| Lab 3 | Vendor B | Vendor II | Vendor 1 |
| Lab 4 | Vendor A | Vendor II | Vendor 3 |
| Lab 5 | Vendor A | Vendor III | Vendor 3 |
| Lab 6 | Vendor B | Vendor I | Vendor 2 |
| Lab 7 | Vendor B | Vendor I | Vendor 1 |

#### 6.2.2.1 Power Delay Profile (PDP)

The PDP should be validated in a Beam-Specific manner. It is assumed that the beams are mapped to the inputs of the channel emulator as follows:

-	Beam 1: Input 1 and Input 2

-	Beam 2: Input 3 and Input 4 (CDL-C UMa only)

The PDP measurement results of UMa CDL-C for bands n41 and n78 are presented in Figures 6.2.2.1-1~5.

(a) Band n41

(b) Band n78

Figure 6.2.2.1-1: Lab 1: PDP measurement results for CDL-C UMa, (a) Band n41, beam 1 and beam 2; (b) Band n78, beam 1 and beam 2. CE bandwidth: 100MHz

(a) Band n41

(b) Band n78

Figure 6.2.2.1-2: Lab 2: PDP measurement results for CDL-C UMa, (a) Band n41, beam 1 and beam 2; (b) Band n78, beam 1 and beam 2

(a) Band n41

(b) Band n78

Figure 6.2.2.1-3: Lab 3: PDP measurement results for CDL-C UMa, (a) Band n41, beam 1 and beam 2; (b) Band n78, beam 1 and beam 2. CE bandwidth: 40MHz

(a) Band n41

(b) Band n78

Figure 6.2.2.1-4: Lab 4: PDP measurement results for CDL-C UMa, (a) Band n41, beam 1 and beam 2; (b) Band n78, beam 1 and beam 2.

(a) Band n41

(b) Band n78

Figure 6.2.2.1-5: Lab 6: PDP measurement results for CDL-C UMa, (a) Band n41, beam 1 and beam 2; (b) Band n78, beam 1 and beam 2.

#### 6.2.2.2 Doppler/Temporal correlation

The Doppler should be validated in a Beam-Specific manner. It is assumed that the beams are mapped to the inputs of the channel emulator as follows:

-	Beam 1: Input 1 and Input 2

-	Beam 2: Input 3 and Input 4 (CDL-C UMa only)

The Doppler measurement results of UMa CDL-C for bands n41 and n78 are presented in Figures 6.2.2.2-1~6.

(a-1) Beam 1, H11                                                 (a-2) Beam 2, H11

(a-3) Beam 1, H12                                                    (a-4) Beam 2, H12

(a-5) Beam 1, H12                                                          (a-6) Beam 2, H12

(a-7) Beam 1, X2V                                                       (a-8) Beam 2, X2V

(a) Band n41

(b-1) Beam 1, H11                                                  (b-2) Beam 2, H11

(b-3) Beam 1, X2V                                                    (b-4) Beam 2, X2V

(b) Band n78

Figure 6.2.2.2-1: Lab 1: Doppler measurement results for CDL-C UMa, (a) Band n41 with different beams and different polarizations; (b) Band n78 with different beams and different polarizations

(a) Band n41

(b) Band n78

Figure 6.2.2.2-2: Lab 2: Doppler measurement results for CDL-C UMa, (a) Band n41, beam 1 and beam 2; (b) Band n78, beam 1 and beam 2

(a) Band n41

(b) Band n78

(c) Band n79

Figure 6.2.2.2-3: Lab 3: Doppler measurement results for CDL-C UMa, (a) Band n41, beam 1, beam 2, and combined beams; (b) Band n78, beam 1 and beam 2; (c) Band n79, beam 1 and beam 2

(a) Band n41

(b) Band n78

Figure 6.2.2.2-4: Lab 4: Doppler measurement results for CDL-C UMa, (a) Band n41, beam 1 and beam 2; (b) Band n78, beam 1 and beam 2

(a) Band n41

(b) Band n78

Figure 6.2.2.2-5: Lab 5: Doppler measurement results for CDL-C UMa, (a) Band n41, beam 1 and beam 2; (b) Band n78, beam 1 and beam 2

(a) Band n41

(b) Band n78

Figure 6.2.2.2-6: Lab 6: Doppler measurement results for CDL-C UMa, (a) Band n41, beam 1 and beam 2; (b) Band n78, beam 1 and beam 2

#### 6.2.2.3 Spatial correlation

The Spatial correlation should be validated in a Beam-Combined manner. It is assumed that the beams are mapped to the inputs of the channel emulator as follows:

-	Beam 1: Input 1 and Input 2

-	Beam 2: Input 3 and Input 4 (CDL-C UMa only)

-	Combined beam for CDL-C UMa: Input 1 + Input 2 + Input 3 + Input 4

-	Combined beam for CDL-C UMi: Input 1 + Input 2

The Spatial correlation measurement results of UMa CDL-C for bands n41 and n78 are presented in Figures 6.2.2.3-1~6.

(a) Band n41

(b) Band n78

Figure 6.2.2.3-1: Lab 1: Spatial correlation measurement results for CDL-C UMa, (a) Band n41, beam 1, beam 2, and combined beams; (b) Band n78, combined beams

(a) Band n41                                                                (b) Band n78

Figure 6.2.2.3-2: Lab 2: Spatial correlation measurement results for CDL-C UMa, (a) Band n41 with combined beams; (b) Band n78 with combined beams

(a) Band n41

(b) Band n78

Figure 6.2.2.3-3: Lab 3: Spatial correlation measurement results for CDL-C UMa, (a) Band n41, beam 1, beam 2, and combined beams; (b) Band n78, combined beams

(a) Band n41

(b) Band n78

Figure 6.2.2.3-4: Lab 4: Spatial correlation measurement results for CDL-C UMa, (a) Band n41, combined beams; (b) Band n78, combined beams

(a) Band n41

(b) Band n78

Figure 6.2.2.3-5: Lab 5: Spatial correlation measurement results for CDL-C UMa, (a) Band n41, combined beams; (b) Band n78, combined beams

(a) Band n41

(b) Band n78

Figure 6.2.2.3-6: Lab 6: Spatial correlation measurement results for CDL-C UMa, (a) Band n41, combined beams; (b) Band n78, combined beams

#### 6.2.2.4 Cross-polarization

The Cross-polarization should be validated in a Beam-Specific manner. It is assumed that the beams are mapped to the inputs of the channel emulator as follows:

-	Beam 1: Input 1 and Input 2

-	Beam 2: Input 3 and Input 4 (CDL-C UMa only)

The Cross-polarization measurement results of UMa CDL-C for bands n41 and n78 are presented in Table 6.2.2.4-1~6.

Table 6.2.2.4-1: Lab 1: Cross-polarization verification results for CDL-C UMa, bands n41 and n78

| Frequency | Beam | Reference | Measurement result |
| --- | --- | --- | --- |
| fc ≤ 2.5 GHz (n41) | Beam 1 | Input 1+2: V/H = 0 dB | 0.44 dB |
|  | Beam 2 | Input 3+4: V/H = 0 dB | 0.55 dB |
| fc > 2.5 GHz (n78) | Beam 1 | Input 1+2: V/H = 0 dB | 0.49 dB |
|  | Beam 2 | Input 3+4: V/H = 0 dB | 0.58 dB |

Table 6.2.2.4-2: Lab 2: Cross-polarization verification results for CDL-C UMa, bands n41 and n78

| Frequency | Beam | Reference | Measurement result |
| --- | --- | --- | --- |
| fc ≤ 2.5 GHz (n41) | Beam 1 | Input 1: V/H = -0.5 dB Input 2: V/H = 0.6 dB | Input 1: V/H = -0.96 dB Input 2: V/H = 0.94 dB |
|  | Beam 2 | Input 3: V/H = -0.6 dB Input 4: V/H = 0.7 dB | Input 3: V/H = -0.11 dB Input 4: V/H = 0.71 dB |
| fc > 2.5 GHz (n78) | Beam 1 | Input 1: V/H = -0.6 dB Input 2: V/H = 0.7 dB | Input 1: V/H = -0.51 dB Input 2: V/H = 1.19 dB |
|  | Beam 2 | Input 3: V/H = -0.7 dB Input 4: V/H = 0.8 dB | Input 3: V/H = -0.41 dB Input 4: V/H = 0.47 dB |

Table 6.2.2.4-3: Lab 3: Cross-polarization verification results for CDL-C UMa, bands n41 and n78

| Frequency | Beam | Reference | Measurement result |
| --- | --- | --- | --- |
| fc ≤ 2.5 GHz (n41) | Beam 1 | Input 1: V/H = -0.5 dB Input 2: V/H = 0.6 dB Input 1+ 2: V/H = 0 dB | Input 1: V/H = -0.98 dB Input 2: V/H = 0.63 dB Input 1+ 2: V/H = -0.21 dB |
|  | Beam 2 | Input 3: V/H = -0.6 dB Input 4: V/H = 0.7 dB Input 3+4: V/H = 0 dB | Input 3: V/H = -1.1dB Input 4: V/H = -0.07 dB Input 3+4: V/H = -0.60 dB |
| fc > 2.5 GHz (n78) | Beam 1 | Input 1: V/H = -0.6 dB Input 2: V/H = 0.7 dB Input 1+ 2: V/H = 0 dB | Input 1: V/H = -0.45 Input 2: V/H = 1.32 dB Input 1+ 2: V/H = 0.39 dB |
|  | Beam 2 | Input 3: V/H = -0.7 dB Input 4: V/H = 0.8 dB Input 3+4: V/H = 0 dB | Input 3: V/H = 0.25 Input 4: V/H = 1.61 Input 3+4: V/H = 0.91 dB |

Table 6.2.2.4-4: Lab 4: Cross-polarization verification results for CDL-C UMa, bands n41 and n78

| Frequency | Beam | Reference | Measurement result |
| --- | --- | --- | --- |
| fc ≤ 2.5 GHz (n41) | Beam 1 | Input 1: V/H = -0.5 dB | -0.48 dB |
|  |  | Input 2: V/H = 0.6 dB | 0.57 dB |
|  | Beam 2 | Input 3: V/H = -0.6 dB | -0.69 dB |
|  |  | Input 4: V/H = 0.7 dB | 0.71 dB |
| fc > 2.5 GHz (n78) | Beam 1 | Input 1: V/H = -0.6 dB | -0.48 |
|  |  | Input 2: V/H = 0.7 dB | 0.99 |
|  | Beam 2 | Input 3: V/H = -0.7 dB | -0.34 |
|  |  | Input 4: V/H = 0.8 dB | 1.13 |

Table 6.2.2.4-5: Lab 5: Cross-polarization verification results for CDL-C UMa, band n41

| Frequency | Beam | Reference | Measurement result |
| --- | --- | --- | --- |
| fc ≤ 2.5 GHz (n41) | Beam 1 | Input 1: V/H = -0.5 dB | -0.48 dB |
|  |  | Input 2: V/H = 0.6 dB | 0.61 dB |
|  | Beam 2 | Input 3: V/H = -0.6 dB | -0.64 dB |
|  |  | Input 4: V/H = 0.7 dB | 0.68 dB |

Table 6.2.2.4-6: Lab 6: Cross-polarization verification results for CDL-C UMa, bands n41 and n78

| Frequency | Beam | Reference | Measurement result |
| --- | --- | --- | --- |
| fc ≤ 2.5 GHz (n41) | Beam 1 | Input 1+2: V/H = 0 dB | -0.6555 dB |
|  | Beam 2 | Input 3+4: V/H = 0 dB | -0.0676 dB |
| fc > 2.5 GHz (n78) | Beam 1 | Input 1+2: V/H = 0 dB | 0.5064 dB |
|  | Beam 2 | Input 3+4: V/H = 0 dB | 0.6414 dB |

#### 6.2.2.5 Power validation

The Power validation results of UMa CDL-C for bands n41 and n78 are presented in Table 6.2.2.5-1~3.

Table 6.2.2.5-1: Lab 1: Power validation results for CDL-C UMa, bands n41 and n78 (Unit: dBm/30kHz)

| CDL-C UMa, n41, 2592.99 MHz |  |  |  |  |
| --- | --- | --- | --- | --- |
| Measured V power | Measured H power | Measured total power | Expected | Delta |
| -79.6244 | -79.5959 | -76.5999 | -77 | 0.4001 |
| CDL-C UMa, n78, 3549.99 MHz |  |  |  |  |
| Measured V power | Measured H power | Measured total power | Expected | Delta |
| -79.9988 | -79.6618 | -76.8168 | -77 | 0.1832 |

Table 6.2.2.5-2: Lab 4: Power validation results for CDL-C UMa, bands n41 and n78 (Unit: dBm/30kHz)

| UMa CDL-C 2450 MHz | Dipole | S=-50.6dBm |
| --- | --- | --- |
|  | Loop | S=-50.4dBm |
| UMa CDL-C 3600 MHz | Dipole | S=-45.6dBm |  |
|  | Loop | S=-46.68dBm |  |

Table 6.2.2.5-3 (a): Lab 6: Power validation results for CDL-C UMa, band n41 (Unit: dBm/30kHz)

Table 6.2.2.5-3 (b): Lab 6: Power validation results for CDL-C UMa, band n78 (Unit: dBm/30kHz)

6.3	Verification of Channel Model implementation of FR2

<Editor’s note: Detailed structure of the subclause is TBD. Include FR2 channel model validation results from labs. Refer to TS38.151 with channel model references, verification procedures and pass/fail limits.>

7	Lab alignment of FR1 MIMO OTA

7.1	General

<Editor’s note: Detailed structure of the subclause is TBD. >

7.2	Noise impact in MPAC on MIMO performance

<Editor’s note: This clause includes analysis of the impact of noise in MPAC on MIMO OTA performance.>

## 7.3	Lab alignment campaign for frequency bands > 1GHz

### 7.3.1	Framework

This clause defines the working procedure on how to proceed the lab alignment campaign for frequency bands > 1GHz. The purpose of Lab Alignment Campaign is to ensure there is no unexpected lab deviation and establish full trust and confidence on the results.

- Labs/companies volunteer to participate in the performance requirement part shall complete the lab alignment measurements and system validation measurements, results should be submitted to RAN4 for review.

- Using the testing conditions as defined in TS38.151.

- The test bands for lab alignment are n41 and n78. Three performance alignment devices (PADs) for each band should be tested to ensure the alignment of measurement results.

- TRMS value in SA mode will be used for alignment comparison.

- The reference value of each PAD should be the average of the PAD measurement results submitted on or before 12:00 UTC 30th April 2022, based on the condition at least 3 labs’ results collected. Submission with measurement data after 12:00 UTC 30th April can be considered for lab alignment, but will not change the reference TRMS value.

- The acceptance criteria for declaring alignment should be defined based on the preliminary MU value of MPAC system. The detailed criteria for accepting the outcome of the lab alignment activity are listed in Table 7.3.3-1.

Table 7.3.3-1: Requirements for lab alignment results (FR1 MIMO OTA)

| Band | Case | Acceptance criteria |
| --- | --- | --- |
| n41 | PAD_1 | The deviation between the measurement result and the reference value of each PAD shall be less than the pass/fail limit, i.e., TRMSaverage,70_mea−TRMSaverage,70_ref≦pass/fail limit |
|  | PAD_2 |  |
|  | PAD_3 |  |
| n78 | PAD_1 | The deviation between the measurement result and the reference value of each PAD shall be less than the pass/fail limit, i.e., TRMSaverage,70_mea−TRMSaverage,70_ref≦pass/fail limit |
|  | PAD_2 |  |
|  | PAD_3 |  |

Note:

- The PAD measurement results shall NOT be shared to anyone before submitting to RAN4 meetings or sharing in the NR MIMO OTA reflector. Comparison and alignment analyses should only be done in RAN4 meetings.

- Three PADs for each band are selected. Labs should submit PAD measurements results in an anonymous approach, i.e., the PADs for each band should be marked as PAD_1, PAD_2, and PAD_3, respectively. The mapping between the codename PAD_n and the actual PAD shall only be known among the labs participated in the alignment activity, and shall NOT be disclosed to any other companies.

### 7.3.2	Measurement results

RAN4 carried out the lab alignment campaign for frequency bands > 1GHz during Rel-17. Considering the test burden and urgent timeline, RAN4 decided to select three smartphones (named as PADs, Performance Alignment Devices) as reference DUTs, and select two bands, n41 and n78, to perform FR1 MIMO OTA lab alignment activity for frequencies bands > 1GHz.

There are 6 test labs participated in the lab alignment campaign, the reference value of each PAD at each band is derived based on linear average (with dBm) of PAD measurement results submitted by all labs. The summary of the lab alignment results is shown in Table 7.3.2-1 and Figure 7.3.2-1. The TRMS offset between each measurement result and reference value of FR1 MIMO OTA lab alignment is shown in Table 7.3.2-2 and Figure 7.3.2-1.

Table 7.3.2-1: Summary of FR1 MIMO OTA lab alignment results for bands > 1GHz

| Device | Band | TRMS measurement result [dBm/30kHz] |  |  |  |  |  | Average approach | Reference value | Max-Min deviation |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | Lab 1 | Lab 2 | Lab 3 | Lab 4 | Lab 5 | Lab 6 |  |  |  |
| PAD_n41_1 | n41 | -96.43 | -97.61 | -98.20 | -97.45 | -96.88 | -99.10 | Linear average | -97.61 | 2.67 |
| PAD_n41_2 | n41 | -99.30 | -97.80 | -100.02 | -99.96 | -99.62 | -101.43 |  | -99.69 | 3.63 |
| PAD_n41_3 | n41 | -96.31 | -97.39 | -97.81 | -96.53 | -96.74 | -98.59 |  | -97.23 | 2.28 |
| PAD_n78_1 | n78 | -96.02 | -96.54 | -96.44 | -96.10 | -96.53 | -98.38 |  | -96.67 | 2.36 |
| PAD_n78_2 | n78 | -95.42 | -95.95 | -96.11 | -96.48 | -96.66 | -98.97 |  | -96.60 | 3.55 |
| PAD_n78_3 | n78 | -99.06 | -97.42 | -99.53 | -99.08 | -99.54 | NA |  | -98.93 | 2.12 |

Table 7.3.2-2: FR1 MIMO OTA lab alignment outcome - TRMS offset (for bands > 1GHz)

| Device | Band | TRMS offset from reference [dBm/30kHz] |  |  |  |  |  | Pass/fail limit |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | Lab 1 | Lab 2 | Lab 3 | Lab 4 | Lab 5 | Lab 6 |  |  |  |
| PAD_n41_1 | n41 | 1.18 | 0.00 | -0.59 | 0.16 | 0.73 | -1.49 | +/- 0.75 MU, i.e., +/- 2.25 dB |  |  |
| PAD_n41_2 | n41 | 0.39 | 1.89 | -0.33 | -0.27 | 0.07 | -1.74 |  |  |  |
| PAD_n41_3 | n41 | 0.92 | -0.16 | -0.58 | 0.70 | 0.49 | -1.36 |  |  |  |
| PAD_n78_1 | n78 | 0.65 | 0.13 | 0.22 | 0.57 | 0.14 | -1.71 | +/- 0.75 MU, i.e., +/- 2.55 dB |  |  |
| PAD_n78_2 | n78 | 1.18 | 0.65 | 0.48 | 0.12 | -0.06 | -2.37 |  |  |  |
| PAD_n78_3 | n78 | -0.13 | 1.51 | -0.60 | -0.15 | -0.61 | NA |  |  |  |
| Lab alignment conclusion |  | Pass | Pass | Pass | Pass | Pass | Pass |  |  |  |

Figure 7.3.2-1: Measurement results of FR1 MIMO OTA lab alignment for bands > 1GHz

Figure 7.3.2-2: Deviation between each measurement result and reference value of FR1 MIMO OTA lab alignment for bands > 1GHz

### 7.3.3	Pass/fail limits

Based on the preliminary MU assessment of FR1 MPAC system in Annex A.2.5 of TS 38.151, and lab alignment measurement results in Clause 7.3.2, RAN4 decided the final pass/fail limits for FR1 MIMO OTA lab alignment activity for bands > 1GHz as 0.75* preliminary MU, i.e.,

-	< 3GHz: 2.25 dB

-	> 3GHz: 2.55 dB

### 7.3.4	Conclusion

RAN4 concluded the successful FR1 MIMO OTA lab alignment activity for bands > 1GHz, all the 6 test labs are aligned. FR1 MIMO OTA performance requirements for bands > 1GHz are specified based on the measurement results submitted by the above aligned test labs.

7.4	Lab alignment campaign for frequency bands < 1GHz

7.4.1	Framework

<Editor’s note: This clause defines the working procedure on how to proceed the lab alignment campaign for frequency bands <1GHz. >

7.4.2	Measurement results

<Editor’s note: This clause includes the measurement results of the reference DUTs in the selected frequency bands (<1 GHz). >

7.4.3	Pass/fail limits

<Editor’s note: This clause defines the pass/fail limits of lab alignment campaign for frequency bands < 1GHz. >

7.4.4	Conclusion

<Editor’s note: This clause describes the conclusion of lab alignment campaign for frequency bands < 1GHz.>

8	Lab alignment of FR2 MIMO OTA

8.1	Framework

<Editor’s note: This clause defines the working procedure on how to proceed the lab alignment campaign for FR2 MIMO OTA. >

8.2	Measurement results

<Editor’s note: This clause includes the measurement results of the reference DUTs in the selected frequency band (FR2). >

8.3	Pass/fail limits

<Editor’s note: This clause defines the pass/fail limits of lab alignment campaign for FR2 MIMO OTA. >

8.4	Conclusion

<Editor’s note: This clause describes the conclusion of lab alignment campaign for FR2 MIMO OTA.>

Annex <A>:
Change history

| Change history |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Date | Meeting | TDoc | CR | Rev | Cat | Subject/Comment | New version |
| 2023-10 | R4#108bis | R4-2316704 |  |  |  | Skeleton of TR38.761 on Measurement of NR MIMO OTA performance | 0.0.1 |
| 2023-11 | R4#109 | R4-2319164 |  |  |  | R4-2316946 TP to TR 38.761 on channel model validation for n41 n78 R4-2316968 TP to TR 38.761 on lab alignment for bands > 1GHz | 0.1.0 |