## Slide 1: Discussion on the Operation of the Probe Module Array for NFTF Measurement
Source: GIST, RRA, CSU
3GPP TSG-RAN WG4 Meeting #111
Fukuoka City, Fukuoka, Japan
May 20-24, 2024
R4-2407599

## Slide 2: Introduction
National Radio Research Agency (RRA), Chosun University, and Gwangju Institute of Science and Technology (GIST) have been developing rapid measurement systems that are based on NFTF technique with simultaneously activated probe modules.
Purpose: market surveillance
  requires rapid measurement of transmitter characteristics of the UE’s in the market
During RAN4 #110, we presented an NFTF system that measures the transmitter characteristics, such as total radiated power, frequency error, occupied bandwidth, spectrum emission mask, adjacent channel leakage ratio, spurious emissions, etc.
The system that consists of 4 arcs of receivers:
  Arc 1 with a radius of 0.8 m covering 3-20 GHz
  Arc 2 and 4 with a radius of 0.8 m covering 20-40 GHz
  Arc 3 with a radius of 0.5 m covering 40-60 GHz
In this contribution, the receiver structure and calibration procedure are discussed for both amplitude and phase measurements, without using a phase recovery unit (PRU)

## Slide 3: Receiver Module
Taking 3 – 20 GHz arc for discussion. 20 – 40 GHz and 40 – 60 GHz arc work with the same principle.
Receiver module = dual-pol ant + analog front-end + digital ckt board
Dual-pol ant takes v- and h-pol fields
RF (3 – 20 GHz)  down converted to IF (2.2 GHz with 400 MHz BW)  sampled in the digital ckt board
LO frequency, sampling rate, and sampling frame are synchronized by a common GPS-disciplined rubidium oscillator
dual-pol
antenna

## Slide 4: Analog Front-End
RF signals are filtered to have 2 GHz BW.
Filtered RF signals are down-converted to 2.2 GHz by mixing with a common LO signal.
  LO takes the 10 MHz output from the GPS disciplined rubidium oscillator as the reference source
IF signal is filtered to have 400 MHz BW.
The digital ckt board takes directly samples the IF signal
V-pol
3-20 GHz
SPnT
S/W
IF 2.2 GHz
SPnT
S/W
IF
RF
LO
BW 400 MHz
preselector filter bank
2 GHz BW
SPnT
S/W
IF 2.2 GHz
SPnT
S/W
IF
RF
LO
BW 400 MHz
LO synthesizer
5-18 GHz
H-pol
3-20 GHz
digital ckt board
digital ckt board

## Slide 5: Receiver Module Synchronization
A common rubidium oscillator outputs are distributed to each module:
  10 MHz output – references sources for LO and ADC sampling frequencies
  1pps output – subdivided to 1 ms and used as the frame sync signal
Synchronization
  SG generates RF signals based on Zadoff-Chu sequence
  Receiver modules down converts the RF and samples in IF
  Digital correlation gives the amount of delay in samples (CAZAC property of sequence)
  Sub-sample delay is determined from the phase spectrum
3-20 GHz
dual-pol.
receiver modules
rubidium clock
10 MHz
1 pps
Zadoff-Chu sequence
10 MHz
1 pps
GPS disciplining
signal generator

## Slide 6: Receiver Module Synchronization
Example:
  Signal 1 (probe module 1):
    time delay: 20.091 ns (10.23 samples)
    phase: -0.576 rad
  Delays more than a sampling interval is determined by correlation
  Phases and delays less than a sampling interval is compensated using the slope in the phase spectrum
  Signal 2 (probe module 2):
    time delay: 56.203 ns (28.45 samples)
    phase: 0.175 rad

## Slide 7: Calibration Procedure
Signal generator (SG) and a calibration antenna are used for calibration
The SG and all the receivers are frame synched with the the GPS-disciplined rubidium oscillator
Procedure
Calibration antenna is directed toward a receiver antenna.
SG generates calibration waveform
The receiver samples the waveform and divide it to subcarrier signals
Phase and amplitude for each subcarrier are recorded
Steps 1-4 are repeated for all receivers for both polarizations.
synchronized
receiver modules
calibration antenna
SG

## Slide 8: Measurement Procedure
All receiver modules and the call simulator are frame synched
The call simulator and the EUT have completed the call setup procedures
The receiver modules in the arc are triggered simultaneously, covering an arc of a spherical surface.
The EUT is rotated by a preset amount in the azimuth direction, and Step 1 is repeated until the entire spherical surface is covered.
The calibration data is applied to the measured data.
This procedure requires one rotation of the EUT
By employing the goniometer, arbitrary sampling on the spherical surface can be achieved.
synchronized
receiver modules
EUT
link
antenna
elevation (goniometer)
azimuth
Call
Simulator

## Slide 9: Conclusion
Key aspects presented:
  Structure of the receiver modules
  Synchronization and calibration procedure
  Measurement procedure utilizing a call simulator
The system can measure both the amplitude and phase information, which are necessary for NFTF processing.
Phase information is acquired without the requirement of a phase recovery unit (PRU).
synchronized
receiver modules
EUT
link
antenna
elevation (goniometer)
azimuth
Call
Simulator