3GPP TSG-RAN WG4 Meeting #79	R4-164197

# Nanjing, China, 23 – 27 May, 2016

Source:	R&S, Keysight Technologies, Bluetest, CTTC

Title:	Harmonization Project Plan

Agenda Item:	7.15.3 Harmonization

Document for:	Approval

# Introduction

This contribution outlines a harmonization project plan incorporating the previously agreed principle of statistical significance so that the first phase of harmonization (two low bands, two high bands) can be performed at a single lab thus allowing most of the agreements reducing uncertainty from the previous harmonization campaign to apply.

This document has been made in co-operation with EMITE, a manufacturer of MIMO OTA Test Equipment, and Félix Belzunce, Professor of Statistical Theory at the University of Murcia, in Spain.

# Goal: Single Lab

The main goal for this proposal is to utilize a single lab for the initial harmonization efforts to reduce the MU, to avoid multiple alignment exercises, and to spend the least amount of time on this effort.

Results from UEs tested as part of the harmonization efforts shall be included in the MIMO OTA performance campaign.

# Statistical Significance

The agreement [1] is to perform harmonization on 30 devices but also allowed for less devices to be tested before declaring harmonization when taking statistical significance into account: “Include statistical analysis to determine when to stop testing more devices for harmonization” [1, subclause 10.4.5 open issue 5]. It is therefore proposed to start the harmonization effort with a total of 5 devices which should allow this effort to be carried out in a single lab supporting all three methodologies (MPAC, RTS, RC+CE). All devices need to be able to support to the ATF for this effort.

In addition to the 5 devices, it is proposed to add an additional 3 devices from the MPAC performance efforts to the harmonization campaign (provided that the devices can support the ATF). Preferably, these additional devices are known good or bad outliers and need to be provided within 6 weeks of the testing of the initial set of harmonization devices; otherwise, different sets of devices will be used to augment the initial set of devices before the harmonization decision based on statistical significance is made.

The following sets of proposals aim to declare harmonization per method based on testing a subset of the original 30 devices while taking the statistical significance concept from Annex A into account:

- Harmonization per method is achieved if the cost of harmonization after testing 8 devices is less than 79% of the target cost of harmonization.

- If harmonization cannot be achieved after the first 8 devices, 7 additional devices are tested and if the cost of harmonization is less than 98% of the target cost of harmonization, harmonization is achieved. Preferably, out of those 7 devices, at least 2 known good or bad outliers should be included which need to be provided within 6 weeks of the testing of the last set of harmonization devices; otherwise, different sets of devices will be used to augment the initial set of devices before the harmonization decision based on statistical significance is made.

- If harmonization cannot be achieved after 15 devices, the full set of the previously agreed 30 devices are tested and harmonization is achieved if the cost of harmonization is less than or equal to the target cost of harmonization.

# Cost of Harmonization

Based on the previous harmonization campaign in [1, subclause 10.3], the residual errors between a pair of test methodologies correspond to the maximum difference of RS_EPRE at the given KPI between the respective test methodologies. To the results from each methodology, one offset per band could be applied to reduce the residual errors. One of the set of methodologies was always MPAC. The harmonized MU term, h, was determined by adding the residual error between methods to the maximum MU of the considered methodologies. Since the largest MU was MPAC and MPAC was always one of the methodologies, m1=mMPAC. The cost of harmonization was later defined as the difference of the harmonized MU term h and the MU of MPAC, i.e., the cost of harmonization was defined to be the residual error between methods. This definition is illustrated in Figure 1.

Figure 1: Definition of the cost of harmonization from the previous harmonization campaign

After the previous harmonization campaign, MPAC was declared as the reference method which affects the cost of harmonization per method since by definition, r = 0 for the reference method, and the maximum uncertainty in device testing will be the reference method’s MU, mR=mMPAC.

If measurements are now made with a second method, the observed spread will be that method’s MU plus any residual error r between that method and the reference method. This is illustrated in Figure 2.

Figure 2: Proposed definition of the cost of harmonization after MPAC has been selected as the reference methodology.

The actual cost of harmonization is defined as the difference between the harmonized MU term, h, and the MU of the reference methodology, mR., i.e., the cost of harmonization is slightly smaller than the residual error since the MUs of the other methodologies (RTS and RC+CE) are smaller than the MU of MPAC.

It shall be noted that even though use of a single harmonization lab helps to minimize the impact of MU on the calculation of r, uncertainties still exist due to the repeatability of the test devices, differences in test systems, calibration procedures, and test equipment such as channel emulators and base station emulators. An estimate of these residual uncertainties in a single lab must be taken into account, either by reducing the observed residuals by an amount corresponding to an estimate of these uncertainties or by adjusting the target cost to a higher level. For example, test device stability/repeatability cannot be allowed to consume any of the margin allowed for harmonized methods.

# Offset resolution

As agreed in [1], each methodology is allowed to apply offsets to shifts results to align with MPAC. It is proposed to define the offsets in multiples of 0.01dB.

# Target cost of harmonization

It is proposed to set the final target for the cost of harmonization to 2 dB. This value takes into account the previous guidance from the operator proposal [2], as well any residual uncertainty when testing in a single lab, e.g., test device repeatability, different measurement systems, channel emulators, base station emulators, etc.

# Proposals

# Start the harmonization campaign in a single lab with a limited set of devices supporting the ATF

  - 5 devices chosen by the harmonization methodology proponents

  - additional 3 devices (good/bad outliers preferred) from the MPAC performance efforts to be provided within 6 weeks of completion of first set of 5 devices, else 3 devices will be alternatively sourced and used to augment the initial set of devices

  - all devices need to support the ATF for RTS

- Results from UEs tested as part of the harmonization efforts shall be included in the MIMO OTA performance campaign.

- Take the following statistical significance into account for the limited set of devices:

  - Harmonization per method is achieved if the cost of harmonization after testing 8 devices is less than or equal to 79% of the target cost of harmonization.

  - If harmonization cannot be achieved after the first 8 devices, 7 additional devices are tested and if the cost of harmonization is less than or equal to 98% of the target cost of harmonization, harmonization is achieved. Preferably, out of those 7 devices, at least 2 known good or bad outliers should be included which need to be provided within 6 weeks of the testing of the last set of harmonization devices; otherwise, different sets of devices will be used to augment the initial set of devices before the harmonization decision based on statistical significance is made.

  - If harmonization cannot be achieved after [15] devices, the full set of the previously agreed 30 devices are tested and harmonization is achieved if the cost of harmonization is less than or equal to the target cost of harmonization.

- The cost of harmonization calculated by taking into account that MPAC is chosen as reference methodology, i.e., C= max(mR,m2+rR,2) - mR

- The offsets are defined in multiples of 0.01dB

- The final target for the cost of harmonization is 2dB

# References

- 3GPP TR 37.977, “Verification of radiated multi-antenna reception performance of User Equipment (UE)”.

- R4-161492, “WF on MIMO OTA” Vodafone et al.

- David, H.A. and Nagaraja, H.N. (2003). Order Statistics, Wiley-Interscience, Hoboken, NJ.

Annex A: Statistical Significance: Downselection of m outliers from a set of n devices

Let us consider that we have a set of observations of size n in method 1 (n averaged results obtained from testing n different devices) and the same number of observations, n, in method 2, giving different results. Let us consider that both methods follow a normal distribution, with different standard deviation. Assuming normality, under the central limit theorem, and given that the difference of the mean values of the two methods are equal to zero by using offsets, the expected maximum value can be obtained from [3]

fn=n σ001F-1xxn-1dx

where F-1 is the quantile function of a typical normal distribution and σ0 =(σm1+ σm2)½.

If we now want to consider only a limited m (m<n) number of observations, but those from outliers (extremes), then the expected value of those outliers can be found by [3]

gm=m σ001G-1xxm-1dx

where

G-1(x)=F-1x m/n  if x<0.5F-1xmn+1-m/n  if x≥0.5

We can now find the relationship in % between these two quantities, taking n=30 as reference and using only m outliers from 2 to 29

We can see that using 5 outliers out of 30 devices seems sufficient to properly identify the maximum.

| m (outliers) | % |
| --- | --- |
| 2 | 30.5955 |
| 3 | 79.3659 |
| 4 | 92.6687 |
| 5 | 97.5896 |
| 6 | 99.5693 |
| 7 | 99.6402 |
| 8 | 99.3557 |
| 9 | 99.2878 |
| 10 | 99.3095 |
| 11 | 99.3647 |
| 12 | 99.4292 |
| 13 | 99.4928 |
| 14 | 99.5519 |
| 15 | 99.6053 |
| 16 | 99.6531 |
| 17 | 99.696 |
| 18 | 99.7344 |
| 19 | 99.769 |
| 20 | 99.8002 |
| 21 | 99.8285 |
| 22 | 99.8544 |
| 23 | 99.878 |
| 24 | 99.8997 |
| 25 | 99.9197 |
| 26 | 99.9382 |
| 27 | 99.9553 |
| 28 | 99.9713 |
| 29 | 99.9861 |