3GPP TSG-RAN4 Meeting #83	R4-1704620

# Hangzhou, China, 15th May 2017 - 19th May 2017

| CR-Form-v11.2.1 |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CHANGE REQUEST |  |  |  |  |  |  |  |  |
|  | 37.977 | CR | 0056 | rev | - | Current version: | 14.3.0 |  |
| For HELP on using this form: comprehensive instructions can be found at http://www.3gpp.org/Change-Requests. |  |  |  |  |  |  |  |  |

| Proposed change affects: | UICC apps |  | ME | X | Radio Access Network |  | Core Network |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |

| Title: | CR on Channel model validation procedures for RTS |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Source to WG: | ROHDE & SCHWARZ, Keysight Technologies |  |  |  |  |  |  |  |  |  |
| Source to TSG: |  |  |  |  |  |  |  |  |  |  |
| Work item code: | LTE_MIMO_OTA |  |  |  |  |  | Date: |  |  | 2017-05-03 |
| Category: | F |  |  |  |  |  | Release: |  |  | Rel-14 |
|  | Use one of the following categories: F (correction) A (mirror corresponding to a change in an earlier release) B (addition of feature), C (functional modification of feature) D (editorial modification) Detailed explanations of the above categories can be found in 3GPP TR 21.900. |  |  |  |  |  |  |  | Use one of the following releases: Rel-8 (Release 8) Rel-9 (Release 9) Rel-10 (Release 10) Rel-11 (Release 11) Rel-12 (Release 12) Rel-13 (Release 13) Rel-14 (Release 14) Rel-15 (Release 15) Rel-16 (Release 16) |  |
| Reason for change: |  | Channel model validation procedure have not been defined in CR |  |  |  |  |  |  |  |  |
| Summary of change: |  | Include channel model validation procedures for RTS |  |  |  |  |  |  |  |  |
| Consequences if not approved: |  | Channel models for RTS cannot be validated |  |  |  |  |  |  |  |  |
| Clauses affected: |  | 8.3.1.2 to 8.3.2.4 |  |  |  |  |  |  |  |  |
|  |  | Y | N |  |  |  |  |  |  |  |
| Other specs |  |  | X | Other core specifications |  |  |  | TS/TR ... CR ... |  |  |
| affected: |  |  | X | Test specifications |  |  |  | TS/TR ... CR ... |  |  |
| (show related CRs) |  |  | X | O&M Specifications |  |  |  | TS/TR ... CR ... |  |  |
| Other comments: |  |  |  |  |  |  |  |  |  |  |

<Beginning of proposed text>

## 8.3	Verification of Channel Model implementations

Channel Models have been specified in Clause 8.2. 
This clause describes the MIMO OTA validation measurements, in order to ensure that the channel models are correctly implemented and hence capable of generating the propagation environment, as described by the model, within a test area, Measurements are done mainly with a Vector Network Analyser (VNA) and a spectrum analyzer.

Note that the term "OTA antenna configuration" used in this subclause refers only to the physical antennas used to perform the validation. The BS antenna assumptions defined as part of the channel model in subclause 8.5 apply in addition to the physical antenna configuration.

### 8.3.1	Measurement instruments and setup

The measurement setup includes the following equipment:

# Table 8.3.1-1: Measurement equipment list for the verification procedure

| Item | Quantity | Item |
| --- | --- | --- |
| 1 | 1 | Channel Emulator |
| 2 | 1 | Signal Generator |
| 3 | 1 | Spectrum Analyzer |
| 4 | 1 | VNA |
| 5 | 1 | Magnetic Dipole |
| 6 | 1 | Sleeve Dipole |

#### 8.3.1.1	Vector Network Analyzer (VNA) setup

Most of the measurements are performed with a VNA. An example set of equipment required for this set-up is shown in Figure 8.3.1.1-1. VNA transmits frequency sweep signals thorough the MIMO OTA test system. 
A test antenna, within the test area, receives the signal and VNA analyzes the frequency response of the system. 
A number of traces (frequency responses) are measured and recorded by VNA and analyzed by a post processing SW, e.g., Matlab. Special care has to be taken into account to keep the fading conditions unchanged, i.e. frozen, during the short period of time of a single trace measurement. The fading may proceed only in between traces. This setup can be used to measure PDP, Spatial Correlation and Polarization of the Channel models defined in clause 8.2.

# Figure 8.3.1.1-1: Setup for VNA measurements

#### 8.3.1.2	Spectrum Analyzer (SA) setup

The Doppler spectrum is measured with a spectrum analyzer as shown in Figure 8.3.1.2-1. In this case a signal generator transmits CW signal through the MIMO OTA test system. The signal is received by a test antenna within the test area. Finally the signal is analyzed by a spectrum analyzer and the measured spectrum is compared to the target spectrum. This setup can be used to measure Doppler Spectrum of the Channel models defined in Clause 8.2.

# Figure 8.3.1.2-1: Setup for VNA measurements

The RTS channel model validation setup consists of a vector signal generator and a spectrum analyser with signal analysis capabilities or scope as depicted in Figure 8.3.1.2-2 and the provided validation procedures apply for both, correlation-based and geometry-based channel model implementations.

### 8.3.2	Validation measurements

#### 8.3.2.1	Power Delay Profile (PDP)

This measurement checks that the resulting Power Delay Profile (PDP) is like defined in the channel model.

# Method of measurement:

Step the emulation and store traces from VNA. I.e. run the emulation to CIR number 1, pause, measure VNA trace, run the emulation to CIR number 10, pause, measure VNA trace. Continue until 1000 VNA traces are measured.

# VNA settings:

# Table 8.3.2.1-1: VNA settings for PDP

| Item | Unit | Value |
| --- | --- | --- |
| Center frequency | MHz | Downlink center frequency in 3GPP TS 36.508 [19] as required per band |
| Span | MHz | 200 [TDB] |
| RF output level | dBm | -15 |
| Number of traces |  | 1000 |
| Distance between traces in channel model | wavelength (Note) | > 2 |
| Number of points |  | 1101 |
| Averaging |  | 1 |
| NOTE: Time [s] = distance [] / MS speed [/s] MS speed [/s] = MS speed [m/s] / Speed of light [m/s] * Center frequency [Hz] |  |  |

# Channel model specification:

# Table 8.3.2.1-2: Channel model specification for PDP

| Item | Unit | Value |
| --- | --- | --- |
| Center frequency | MHz | Downlink center frequency in 3GPP TS 36.508 [19] as required per band |
| Channel model samples | wavelength | > 2000 |
| Channel model |  | As specified in Clause 8.2 |

# Method of measurement result analysis:

Measured VNA traces (frequency responses H(t,f)) are saved into a hard drive. The data is read into, e.g., Matlab. 
The analysis is performed by taking the Fourier transform of each FR. The resulting impulse responses h(t,tau) are averaged in power over time:

Finally the resulting PDP is shifted in delay, such that the first tap is on delay zero. The reference PDP plots from Table 8.2-1 and Table 8.2-2 are shown in Figure 8.3.2.1-1.

OTA antenna configuration: 		For e.g. 1 full ring (or single cluster configuration) of V polarized elements.

Measurement antenna: 			For e.g. Vertically oriented sleeve dipole.

Figure 8.3.2.1-1: Reference PDP values for SCME Urban Macro / SCME Urban Micro
 plotted from Table 8.2-1 and Table 8.2-2

Method of measurement for RTS:

The vector signal generator repeatedly generates the Zadoff-Chu sequence as defined in Table 8.3.2.1-3. After a given time interval, the spectrum analyser or scope is triggered by the vector signal generator at the beginning of the next Zadoff-Chu sequence. The time interval needs to exceed the required minimum distance between two PDPs as defined in Table 8.3.2.1-4 and depends on the centre frequency and the mobile speed.

After receiving the trigger signal, the spectrum analyser or scope starts with the acquisition of the IQ data. The number of recorded IQ samples corresponds to the length of the used Zadoff-Chu sequence. This step is repeated until at least 1000 blocks of IQ samples are recorded.

Vector signal generator settings:

Table 8.3.2.1-3: PDP measurement parameters for signal generator

| Item | Unit | Value |
| --- | --- | --- |
| Centre frequency | MHz | Downlink centre frequency in 3GPP TS 36.508 [19]as required per band |
| Min. signal bandwidth | MHz | 100 |
| RF output level | dBm | -15 |
| Sequence type |  | Zadoff-Chu or similar waveform with good autocorrelation properties |
| Min. sequence length | Samples | 1023 |

Signal and spectrum analyser settings:

Table 8.3.2.1-4: PDP measurement parameters for spectrum analyser

| Item | Unit | Value |
| --- | --- | --- |
| Centre frequency | MHz | Downlink centre frequency in 3GPP TS 36.508 [19] as required per band |
| Min. sampling rate | MHz | 100 |
| Min. number of PDPs for averaging |  | 1000 |
| Distance between PDPs in channel model | Wavelength (Note) | > 2 |
| NOTE: Time [s] = distance [] / MS speed [/s] MS speed [/s] = MS speed [m/s] / Speed of light [m/s] * Centre frequency [Hz] |  |  |

The channel model specifications match those from Table 8.3.2.1-2.

Method of measurement result analysis:

The measured blocks of IQ samples are saved and later analysed by a post processing SW, e.g., Matlab. The analysis is performed by applying the circular cross-correlation between each recorded block of IQ samples and the IQ samples of the original Zadoff-Chu sequence. The resulting impulse responses h(t,tau) are averaged in power over time:

Finally the resulting PDP is shifted in delay and normalized in power, such that the first tap is on delay zero and has a relative power of 0 dB.

Antenna configurations: 		For e.g. a pair of vertically oriented sleeve dipoles or vertically polarized horn antennas.

#### 8.3.2.2	Doppler/Temporal correlation

This measurement checks the Doppler/temporal correlation.

# Method of measurement:

Sine wave (CW, carrier wave) signal is transmitted from the signal generator. The signal is connected from the signal generator to fading emulator via cables. The fading emulator output signals are connected to power amplifier boxes via cables. The amplified signals are then transferred via cables to the probe antennas. The probe antennas radiate the signals over the air to the test antenna The Doppler spectrum is measured by the spectrum analyzer and the trace is saved.

# Signal generator settings:

# Table 8.3.2.2-1: Signal generator settings for Doppler/Temporal correlation

| Item | Unit | Value |
| --- | --- | --- |
| Center frequency | MHz | Downlink center frequency in 3GPP TS 36.508 [19] as required per band |
| Output level | dBm | -15 |
| Modulation |  | OFF |

# Spectrum analyzer settings:

# Table 8.3.2.2-2: Spectrum analyzer settings for Doppler/Temporal correlation

| Item | Unit | Value |
| --- | --- | --- |
| Center frequency | MHz | Downlink center frequency in 3GPP TS 36.508 [19] as required per band |
| Minimum Span | Hz | max(4 x MaxDoppler, 400 Hz) |
| RBW | Hz | 1 |
| VBW | Hz | 1 or use FFT |
| Number of points |  | 8001 |
| Averaging |  | 100 |

# Channel model specification:

# Table 8.3.2.2-3: Channel model specification for Doppler/Temporal correlation

| Item | Unit | Value |
| --- | --- | --- |
| Center frequency | MHz | Downlink center frequency in 3GPP TS 36.508 [19] as required per band |
| Channel model |  | As specified in Clause 8.2 |
| Mobile speed | km/h | 100 |

Method of measurement result analysis: Measurement data file (Doppler power spectrum) is saved into hard drive. 
The data is read into, e.g., Matlab. 
The analysis is performed by taking the Fourier transformation of the Doppler spectrum. 
The resulting temporal correlation function   is normalized such that . 
Then the function values left from the maximum is cut out. 
Further on the function values after, e.g. seven periods is cut out. 
The reference temporal correlation plots from Table 8.2-1 and 8.2-2 are shown in Figure 8.3.2.2-1 and Figure 8.3.2.2-2.

OTA antenna configuration: 		For e.g. 1 full ring (or single cluster configuration) of V polarized elements.

Measurement antenna: 			For e.g. vertically oriented dipole.

Figure 8.3.2.2-1: Reference Temporal Correlation Functions for SCME Urban Macro (left)
 and SCME Urban Micro (right) plotted from Table 8.2-1 and Table 8.2-2

Figure 8.3.2.2-2: Reference Temporal Correlation Function for correlation implementation of
 SCME UMa and UMi with Jake's Doppler spectrum plotted from Table 8.2-1 and Table 8.2-2

#### 8.3.2.3	Spatial correlation

This measurement checks whether the measured correlation curve follows the theoretical curve.

# Method of measurement:

Step the emulation and store traces from VNA. I.e. run the emulation to CIR number 1, pause, measure VNA traces in 11 different DUT positions, run the emulation to CIR number 10, pause, measure VNA traces in 11 different DUT positions, … etc. Continue until frequency response of 1000 CIRs in 11 positions (=1000*11 VNA traces) are measured.

11 test antenna positions sample a segment of line of length 1 wavelength with sampling interval of 0.1 wavelengths. Antenna spacing (wave lengths): -0.5 to +0.5 step of 0.1.

# Figure 8.3.2.3-1: Test antenna positions

# VNA settings:

# Table 8.3.2.3-1: VNA settings for spatial correlation

| Item | Unit | Value |
| --- | --- | --- |
| Center frequency | MHz | Downlink center frequency in 3GPP TS 36.508 [19] as required per band |
| Span | MHz | 10 |
| RF output level | dBm | -15 |
| Number of traces |  | 1000 |
| Distance between traces in channel model | Wavelength (Note) | > 2 |
| Number of points |  | 1 (or the smallest possible) |
| Averaging |  | 1 |
| NOTE: Time in seconds = distance [] / MS speed [/s] MS speed [/s] = MS speed [m /s] / Speed of light [m/s] * Center frequency [Hz] |  |  |

# Channel model specification:

# Table 8.3.2.3-2: Channel model specification for spatial correlation

| Item | Unit | Value |
| --- | --- | --- |
| Center frequency | MHz | Downlink center frequency in 3GPP TS 36.508 [19] as required per band |
| Channel model samples | Wavelength | > 2000 |
| Channel model |  | As specified in Clause 8.2 |
| Mobile speed | km/h | 30 |

# Measurement Procedure

# CALIBRATE

# OPEN corrVNATrace trace file

FOR EACH gridPoint IN [test zone grid set]

MOVE measurement antenna to gridPoint

FOR EACH chanIRNumber IN [0:SD:1000*SD]

Measure Freq Resp with VNA

Save freq resp trace to trace file

# END

# END

# CLOSE corrVNATrace_<calibMethod>_<polarization> trace file

# Method of Measurement Results Analysis

Calculate correlation of 1000 x 11 matrix H(f) of frequency response samples. The procedure is to correlate sixth column (the trace measured at the centre of chamber) with the 10 other columns as follows (Matlab example)

for ind = 1:11;

Corr(: , : , ind) = abs(corrcoef(H(: , 1),H(: , ind)));

end

Correlation = squeeze(Corr(1, 2, :));

The reference spatial correlation plots from Table 8.2-1 and 8.2-2 are shown in Figure 8.3.2.3-2.

OTA antenna configuration: 		For e.g. 1 full ring (or a single cluster configuration) of V polarized elements.

Measurement antenna: 			For e.g. Sleeve dipole.

Figure 8.3.2.3-2: Reference Spatial Correlation Functions for 
SCME Urban Macro / SCME Urban Micro plotted from Table 8.2-1 and Table 8.2-2

For RTS, spatial correlation is verified using the setup in Figure 8.3.1.2-2. The channel emulator is configured for a MIMO channel model with two transmit and two receive antenna patterns. The transmit antennas are assumed to be dual-polarized equal power elements with a fixed 0λ separation, 45 degrees slanted, Clause 8.5, while the two receiver antennas are configured as vertically polarized dipoles with varying antenna separation. The antenna positions sample a line of length 1λ with a sampling interval of 0.1λ.

Vector signal generator settings:

Table 8.3.2.3-3: Signal generator settings for spatial correlation for RTS

| Item | Unit | Value |
| --- | --- | --- |
| Centre frequency | MHz | Downlink centre frequency in 3GPP TS 36.508 [19] as required per band |
| Min. signal bandwidth | kHz | 10 |
| RF output level | dBm | -15 |
| Sequence type |  | Zadoff-Chu or similar waveform with good autocorrelation properties |
| Min. sequence length | Samples | 1023 |

Signal and spectrum analyser settings:

Table 8.3.2.3-4: Spectrum analyser settings for spatial correlation for RTS

| Item | Unit | Value |
| --- | --- | --- |
| Centre frequency | MHz | Downlink centre frequency in 3GPP TS 36.508 [19] as required per band |
| Min. sampling rate | kHz | 10 |
| Min. measurement duration | Wavelength (Note) | 2000 |
| NOTE: Time [s] = distance [] / MS speed [/s] MS speed [/s] = MS speed [m/s] / Speed of light [m/s] * Centre frequency [Hz] |  |  |

The channel model specifications match those from Table 8.3.2.3-2.

Measurement Process:

For each antenna position, the signal generator continuously sends the Zadoff-Chu sequence to the channel emulator which applies the required channel model as specified in Table 8.3.2.3-2. The output RF signal from the channel emulator is connected to the TX antenna and received by the RX antenna which is connected to the RF input of the spectrum analyser or scope, Figure 8.3.1.2-2.

The RF signals received by the RX antenna for different simulated antenna positions/separations are captured by the spectrum analyser or the scope. The spatial correlation analysis is performed by performing the cross-correlation operation directly on the received signals captured for different antenna positions/separations.

#### 8.3.2.4	Cross-polarization

This measurement checks how well the measured vertically or horizontally polarized power levels follow expected values.

Method of measurement: 		Step the emulation and store traces from VNA.

# VNA settings:

# Table 8.3.2.4-1: VNA settings for cross-polarization

| Item | Unit | Value |
| --- | --- | --- |
| Center frequency | MHz | Downlink Center Frequency in 3GPP TS 36.508 [19] as required per band |
| Span | MHz | 10 |
| RF output level | dBm | -15 |
| Number of traces |  | 1000 |
| Distance between traces in channel model | Wavelength (Note) | > 2 |
| Number of points |  | 201 |
| Averaging |  | 1 |
| NOTE: Time [s] = distance [] / MS speed [/s] MS speed [/s] = MS speed [m /s] / Speed of light [m/s] * Center frequency [Hz] |  |  |

# Channel model specification:

# Table 8.3.2.4-2: Channel model specification for cross-polarization.

| Item | Unit | Value |
| --- | --- | --- |
| Center frequency | MHz | Downlink center frequency in 3GPP TS 36.508 [19] as required per band |
| Channel model samples | wavelength | > 2000 |
| Channel model |  | As specified in Clause 8.2 |
| Mobile speed | km/h | 30 |

# Measurement Procedure

1.	Play or step through the channel model -> SCME UMi, or UMa X Corr.

2.	Measure the absolute power received at the center of the array, averaged over a statistically significant number of fades.

a.	Use a vertically polarized sleeve dipole to measure the V component.

b.	Use a horizontally polarized (vertically oriented) magnetic loop dipole, or a horizontally polarized sleeve dipole measured in two orthogonal horizontal positions and summed to measure the H component.

3.	Calculate the V/H ratio.

4.	Compare it with the theory -> 0.83dB for UMi, and 8.13dB for UMa.

# Expected measurement results

V/H ratio (composite, i.e. all 6 paths combined) of the 3GPP SCME Umicro model is 0.83 dB and for Umacro 8.13 dB. 
The BS antennas are isotropic dipoles with +/- 45 degrees slant and subject to a foreshortening of the slanted radiating element. See channel model details specified in clause 8.2.

For RTS, the cross-polarization is verified using the setup in Figure 8.3.1.2-2. The channel emulator is configured for a MIMO channel model with two transmit and two receive antenna patterns. The transmit antenna patterns are assumed to be dual-polarized equal power elements with a fixed 0λ separation, 45 degrees slanted, Clause 8.5. One simulated receive antenna is configured as an ideal dipole with vertical polarization, and the other one is configured as ideal dipole with horizontal polarization. Both antennas have a fixed 0λ separation.

This cross polarization measurement verifies the correct power ratio of the different polarizations in the channel model. The measurement can be accomplished by either using a physical transmit and receive antenna with a single polarization while alternately changing the RX antenna polarization configuration in the fading emulator or by using physical dual-polarized antennas with high isolation between the two polarizations of the transmit and receive antennas. Both approaches guarantee that only the cross polarization of the channel model is verified.

# Channel model specification:

The channel model specifications match those from Table 8.3.2.4-2.

Measurement Procedure

- Play the channel model defined in Table 8.3.2.4-2.

- Measure the absolute power received at the RX antenna, averaged over a statistically sufficient number of fades (defined as channel model samples in Table 8.3.2.4-2).

  - Measure the power intended for the vertically polarized receive antenna to measure the V component.

  - Measure the power intended for the horizontally polarized receive antenna to measure the H component.

- Calculate the V/H ratio.

Theoretical and expected measurement results are the same as MPAC.

<End of proposed text>