3GPP TSG-RAN WG4 #86bis	R4-1804990

Melbourne, Australia 16th April –20th April 2018

Source:	Ericsson

Title:	TP to draft CR 37.843: CATR Test Method Measurement Uncertainty for EVM

Agenda item:	6.26.3.1.2

Document for:	Approval

# Introduction

This contribution presents the measurement uncertainty framework and assessment for EVM measurement in a Compact Antenna Test Range which is a method for conformance testing in [2].

# Discussion

Modulation quality is defined by the difference between the measured carrier signal and a reference signal. Modulation quality can be expressed as Error Vector Magnitude (EVM) for E-UTRA.

The EVM requirement however is important only between the BS and its intended UE. Hence the EVM is only of importance in directions where the BS intends to communicate with the UE’s. It is therefore in a set of specific directions from the AAS BS which is of interest rather than the average or total EVM over the entire sphere.

The range of directions which the AAS BS meets the EVM requirement is declared by the manufacturer as an EVM directions range.  As described in the core requirement, and also captured in the description portion of section 10 in TR 37.843 the conformance is required by measuring EVM by configuring the eNB to point a beam in one of the 5 conformance directions and repeat for all 5 directions.  The OTA test is only required to be met in the particular direction during the time of test, as the test is repeated for different conformance points it is not longer required that the EVM requirement is needed to be met at the previous conformance point.  In other words, the EVM requirement is only mandated to be met at that spatial point during that specific beamforming setting and not during other beam forming settings.

Figure 1: Example of OTA coverage range and 5 conformance points

To reflect those points, the following is a proposed measurement uncertainty for a compact antenna test range for measuring OTA EVM. It should be noted that the mechanism by which physical effects lead to uncertainty may differ between EIR and EVM. For EVM, receiving a power level different than expected at the input to the range antenna is not of huge importance.  However, if the beam point direction is not precisely on the OTA EVM requirement conformance point and the EVM varies spatially, then the wrong EVM level would be captured.  Or as another example, the quality of the QZ would play a part in the effect of multipath interference which could effect the measured EVM value, although we do not expect this to be large, the uncertainty of this would need to be captured. Thus care needs to be taken not just to quantify the uncertainties but to understand the mechanism by which they impact the EVM measurement.

# Conclusion

This contribution presents the MU budget for EVM.  Many aspects are still FFS, but as a starting point we propose to have the following TP to be considered.  The values have been derived from using the EIRP MU from TR 37.842 and contributors that have been removed due to the negligible residual value from the fact that EVM is a ratio measurement requirement.

How the values should be expressed, either as % or first in dB then converted to % in the final value, and the understanding of the mechanism by which the uncertainties impact EVM is also for FFS.

# References

[1]	TS 37.105, Active Antenna System (AAS) Base Station (BS) transmission and reception

[2] 	TS 37.145-2, Active Antenna System (AAS) Base Station (BS) conformance testing; Part 2: radiated conformance testing

[3] 	TR 37.842, Evolved Universal Terrestrial Radio Access (E-UTRA) and Universal Terrestrial Radio Access (UTRA; Radio Frequency (RF) requirement background for Active Antenna System (AAS) Base Station (BS)

[Start of Text Proposal]

10.2.2.1.1.4	Uncertainty Budget Format

# Table 10.2.2.1.1.4-1: CATR uncertainty budget format for EVM measurement

| UID | Description of uncertainty contribution | Details in annex |
| --- | --- | --- |
| Stage 2: DUT measurement |  |  |
| 1 | Misalignment DUT & pointing error | B2-1 |
| 2 | RF power measurement equipment | E |
| 3 | Standing wave between DUT and test range antenna | B2-3 |
| 4 | RF leakage (SGH connector terminated & test range antenna connector cable terminated) | B2-4 |
| 5 | QZ ripple DUT | B2-5 |
| 19 | Miscellaneous uncertainty | B2-14 |

10.2.2.1.1.5	Uncertainty assessment

# Table 10.2.2.1.1.5-1: CATR uncertainty assessment for EVM measurement

| UID | Uncertainty Source | Uncertainty value f ≦ 3GHz | Uncertainty value 3 GHz < f ≦ 4.2 GHz | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB] f ≦ 3GHz | Standard uncertainty ui [dB] 3 GHz < f ≦ 4.2 GHz |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |  |
| 1 | Misalignment DUT & pointing error | FFS | FFS | Exp. normal | 2 | 1 | FFS | FFS |
| 2 | Spectrum Analyzer | [0.33%] | [0.33%] | Gaussian | 1 | 1 | [0.33%] | [0.33%] |
| 3 | Standing wave between DUT and test range antenna | FFS | FFS | U-shaped | √2 | 1 | FFS | FFS |
| 4 | RF leakage (SGH connector terminated & test range antenna connector cable terminated) | FFS | FFS | Normal | 1 | 1 | FFS | FFS |
| 5 | QZ ripple DUT | FFS | FFS | Normal | 1 | 1 | FFS | FFS |
| 19 | [Miscellaneous uncertainty] | FFS | FFS |  |  | 1 | FFS | FFS |
| Combined standard uncertainty (1σ) [dB] |  |  |  |  |  |  |  |  |
| Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  |  |  |  |

[End of Text Proposal]