3GPP TSG-RAN WG4 #86-bis	R4-1805112

Melbourne, Australia, 16th – 20th April 2018

Source:	MVG Industries

Title:	TP to Draft CR 37.843 - OTA EVM measurement in a Near Field Test Range

Agenda item:	6.26.3.1.2

Document for:	Approval

# Introduction

During 3GPP RAN4 #86, a contribution highlighting the test procedures for EVM type of measurements in Near Field Test Range was provided [1].

This contribution is revising the test procedure for EVM measurement in Near Field test Range.

# Discussion

Near Field is a testing methodology which can be used for radiated transmit power. Basically, the test would require to measure EVM in some manufacturers declared directions (OTA coverage range). The requirement needs to be defined so that all transmitter units are active and the system is operating at the declared maximum rated total radiated power. For conformance testing purposes, the EVM is tested at the maximum and minimum power settings, in order to comply with the total power dynamic range requirement (refer to the Total power dynamic range requirement for E-UTRA in section 5.3).

EVM is a directional requirement, the calibration of the system setup is not taken into account.

# Proposal

<< Unchanged sections omitted >>

----------------------------------------------------------------------------------------------------------------------------------------------------

<< start of proposal >>

## 10.2.2.2 OTA EVM

## 10.2.2.2.3 Near Field

### 10.2.2.2.3.1 Description

The near field measurement techniques consist in measuring amplitude and phase of the modulated signal at the AAS BS under test, on some specific surfaces such as planar, cylindrical, and spherical. Each of the near field test methods can be implemented by one or more mechanical rotations of the probe and/or AAS BS under test as shown in figure 10.3.2.1.4.1-1. One or more mechanical movement can be substitute by a probe array. All the scanning methods will need an RF transmit and receive system equipped with an automated scanning, a data collection and control system, and computerized analysis ability.

# Figure 10.2.2.2.3.1-1: Probe/scanner near field systems: spherical, cylindrical and planar

NOTE:	Although there are three methods available for obtaining the near field data, the spherical method is used as a working example.

In case the radiated field is sampled on a sphere surrounding completely the AAS BS under test, the 3D full sphere EIS values can be measured in near field when the AAS BS is receiving a defined modulated signal. The near field measurement techniques would imply the use of mathematical artefact, NF to FF transform in order to have the EIS value the in far field. The near field to far field transform is based on the well-known Huygens-Fresnel principle. The spherical modal wave expansion is the implicit application of the Huygens principle. A direct solution of the Helmholtz equations is found by applying boundary conditions on the surface S at infinity. From the tangential fields over the surface, the modal coefficients can be determined using the orthogonality of the modal expansion.  Based on this formulation, the near field sampling criteria does play an important role.

### 10.2.2.2.3.2 Procedure

Stage 1 – Calibration:

For directional requirement, it is not needed to calibrate the system setup.

Stage 2 - Measurement:

The testing procedure consists of the following steps:

- Measure the Near Field pattern -> same as step 1 for the radiated transmit power [2]

- Compute the NFtoFF transform -> EIRP pattern in Far Field

  - Calibration would be applied to the Near Field pattern -> this process would convert dB in dBm [2]

- Record the direction to where EVM shall be measured (either beam peak or one point in the main beam).

- Move the AAS BS in order to have it aligned with (Theta;Phi) angles corresponding to the directions to be measured (step 3)

- Measure EVM in Near Field by using a Spectrum Analyzer.

The direction is declared by the AAS BS manufacturer.

For conformance tests, EVM shall be measured at maximum and minimum power settings.

### 10.2.2.2.3.3 Uncertainty budget format

Table 10.2.2.2.3.3-1: Near Field Test Range Uncertainty contributors for eAAS OTA EVM measurement

| UID | Description of uncertainty contribution | Details in paragraph |
| --- | --- | --- |
| 1 | Axes Intersection | 10.2.2.2.3.3-1 |
| 2 | Axes Orthogonality | 10.2.2.2.3.3-2 |
| 3 | Horizontal Pointing | 10.2.2.2.3.3-3 |
| 4 | Probe Vertical Position | 10.2.2.2.3.3-4 |
| 5 | Probe H/V pointing | 10.2.2.2.3.3-5 |
| 6 | Measurement Distance | 10.2.2.2.3.3-6 |
| 7 | Amplitude and Phase Drift | 10.2.2.2.3.3-7 |
| 8 | Amplitude and Phase Noise | 10.2.2.2.3.3-8 |
| 9 | Leakage and Crosstalk | 10.2.2.2.3.3-9 |
| 10 | Amplitude Non-Linearity | 10.2.2.2.3.3-10 |
| 11 | Amplitude and Phase Shift in rotary joints | 10.2.2.2.3.3-11 |
| 12 | Channel Balance Amplitude and Phase | 10.2.2.2.3.3-12 |
| 13 | Probe Polarization Amplitude and Phase | 10.2.2.2.3.3-13 |
| 14 | Probe Pattern Knowledge | 10.2.2.2.3.3-14 |
| 15 | Multiple Reflections | 10.2.2.2.3.3-15 |
| 16 | Room Scattering | 10.2.2.2.3.3-16 |
| 17 | DUT support Scattering | 10.2.2.2.3.3-17 |
| 18 | Scan Area Truncation | 10.2.2.2.3.3-18 |
| 19 | Sampling Point Offset | 10.2.2.2.3.3-19 |
| 20 | Spherical Mode Truncation | 10.2.2.2.3.3-20 |
| 21 | Positioning | 10.2.2.2.3.3-21 |
| 22 | Probe Array Uniformity | 10.2.2.2.3.3-22 |
| 23 | Mismatch of receiver chain (i.e. between receiving antenna and measurement receiver) | 10.2.2.2.3.3-23 |
| 24 | Insertion loss of receiver chain | 10.2.2.2.3.3-24 |
| 25 | Uncertainty of the absolute gain of the probe antenna | 10.2.2.2.3.3-25 |
| 26 | RF Measurement Equipment | 10.2.2.2.3.3-26 |
| 27 | Measurement repeatability - Positioning Repeatability | 10.2.2.2.3.3-27 |
| 28 | Phase Curvature | 10.2.2.2.3.3-28 |
| 29 | Phase Recovery Uncertainty | 10.2.2.2.3.3-29 |

# 10.2.2.2.3.3-1	Axes Intersection

This is a mechanical uncertainty term and aim to find the uncertainty related with the lateral displacement between the horizontal and vertical axes of the AAS BS positioner. This can result in sampling the field on a non-ideal sphere. This uncertainty is assumed to have a Gaussian distribution.

# 10.2.2.2.3.3-2	Axes Orthogonality

The difference from 90° of the angle between the horizontal and vertical axes also results in sampling the field on a non-ideal sphere. This uncertainty is assumed to have a Gaussian distribution.

# 10.2.2.2.3.3-3	Horizontal Pointing

The horizontal mispointing of the horizontal axis to the probe reference point for Theta=0° also results in sampling the field on a non-ideal sphere. This uncertainty is assumed to have a Gaussian distribution.

# 10.2.2.2.3.3-4	Probe Vertical position

The vertical displacement of the probe reference point from the horizontal axis results in sampling the field on a non-ideal sphere. This uncertainty is assumed to have a Gaussian distribution.

# 10.2.2.2.3.3-5	Probe Horizontal/Vertical pointing

The horizontal or vertical mispointing of the probe z-axis from the intersection point of the horizontal/vertical axis. This uncertainty is assumed to have a Gaussian distribution.

# 10.2.2.2.3.3-6	Measurement distance

This is the knowledge of the distance between the intersection point of the horizontal and vertical axis and probe reference point. This uncertainty is assumed to have a Gaussian distribution.

# 10.2.2.2.3.3-7	Amplitude and phase drift

The system drift due to temperature variations the signal at AAS BS location to drift in amplitude and phase. This uncertainty is assumed to have a Gaussian distribution.

# 10.2.2.2.3.3-8	Amplitude and phase noise

This uncertainty is due to the noise level of the test range so that the S/N ratio should be determined or measured at the AAS BS location. The noise level is usually measured with a Spectrum Analyzer. This uncertainty is assumed to have a Gaussian distribution.

# 10.2.2.2.3.3-9	Leakage and Crosstalk

This uncertainty can be addressed by measurements on the actual system setup. The leakage and crosstalk cannot be separated from the random amplitude and phase errors so that the relative importance should be determined. This uncertainty is assumed to have a Gaussian distribution.

# 10.2.2.2.3.3-10	Amplitude non-Linearity

This uncertainty is the linearity of the receiver used for the measurement. It can be taken from the data sheet of the receiver.

# 10.2.2.2.3.3-11	Amplitude and phase shift in rotary joint

This uncertainty is due to the variation of the rotary joint. It can be measured and is assumed to have a Gaussian distribution.

# 10.2.2.2.3.3-12	Channel balance amplitude and phase

This uncertainty is relevant for systems which are using dual polarized probes and polarization switches. The amplitude and phase difference between two signal channels of the receiver includes the difference between the probe ports, difference between the channels of the polarization switch, connecting cables and reflection coefficients. This uncertainty is assumed to have a Gaussian distribution.

# 10.2.2.2.3.3-13	Probe polarization amplitude and phase

The amplitude and phase of the probe polarization coefficients should be measured. This uncertainty is assumed to have a Gaussian distribution.

# 10.2.2.2.3.3-14	Probe pattern knowledge

The probe(s) pattern(s) is assumed to be known so that the AAS BS measurement in near field can be corrected when performing the near field to far field transform. There is no direct dependence between the AAS BS pattern and the probe pattern in near field measurements. This uncertainty is assumed to have a Gaussian distribution.

# 10.2.2.2.3.3-15	Multiple reflections

The multiple reflections occur when a portion of the transmitted signal is reflected form the receiving antenna back to the transmitting antenna and re-reflected by the transmitting antenna back to the receiving antenna. This uncertainty can be determined by multiple measurements of the AAS BS when at different distance from the probes. This uncertainty is assumed to have a Gaussian distribution.

# 10.2.2.2.3.3-16	Room scattering

As for the multiple reflections, a portion of the transmitted signal is reflected by either the absorbers or other structures in the measurement anechoic chamber before being received by the receiving antenna. This effect can be isolated from the multiple reflections by testing the AAS BS in different positions, separated by λ/4 with respect to the anechoic chamber and comparing these measurements with the reference. This uncertainty is assumed to have a Gaussian distribution.

# 10.2.2.2.3.3-17	DUT support scattering

This is the uncertainty due to the AAS BS supporting structure on the signal level. This uncertainty is assumed to have a Gaussian distribution.

# 10.2.2.2.3.3-18	Scan area truncation

This uncertainty does affect this near field measurement. It can be addressed by comparing the measurement result when scanning the full area. This uncertainty is assumed to have a Gaussian distribution.

# 10.2.2.2.3.3-19	Sampling point offset

This uncertainty has an influence in near field and far field. It is assumed to have a Gaussian distribution.

# 10.2.2.2.3.3-20	Mode truncation

The measured near field is expanded using a finite set of spherical modes. The number of modes is linked to number of samples. The filtering effect generated by the finite number of modes can improve measurement results by removing signals from outside the physical area of the AAS BS. Care should be taken in order to make sure the removed signals are not from the AAS BSitself. This uncertainty is usually negligible.

# 10.2.2.2.3.3-21	Positioning

The relative position of the probe array is not ideal. This uncertainty is assumed to have a rectangular distribution.

# 10.2.2.2.3.3-22	Probe array uniformity

This is the uncertainty due to the fact that different probes are used for each physical position. Different probes have different radiation patterns. Generally, the probe array is calibrated so that the uniformity of the probes is achieved.

# 10.2.2.2.3.3-23	Mismatch of receiver chain

If the same chain configuration (including the measurement receiver; the probe antenna and other elements) is used in both stages, the uncertainty is considered systematic and constant  0.00 dB value.

If it is not the case, this uncertainty contribution has to be taken into account and should be measured or determined by the method described in [14]. This uncertainty is assumed to have a U-shaped distribution

# 10.2.2.2.3.3-24	Insertion loss of receiver chain

It is composed of the following:

-	Insertion loss of the probe antenna cable.

-	Insertion loss of the probe antenna attenuator (if used).

-	Insertion loss of RF relays (if used).

If the same chain configuration is used for measurement and calibration, the uncertainty due to the above components is considered systematic and constant  0.00 dB value. This uncertainty is assumed to have a Gaussian distribution.

# 10.2.2.2.3.3-25	Uncertainty of the absolute gain of the probe antenna

This uncertainty appears in the both stages and it is thus considered systematic and constant 0.00 dB value.

# 10.2.2.2.3.3-26	RF Measurement Equipment

# This uncertainty is related to the test equipment which receives the signal.  All uncertainties that pertain to the measurement being performed shall be included.

# 10.2.2.2.3.3-27	Measurement Repeatability - Positioning Repeatability

This uncertainty is due to the repositioning of the AAS BS in the test setup. It can be addressed by repeating the corresponding measurement 5 times. Calculate the standard deviation of the metric obtained and use that as the measurement uncertainty. For tests that require multiple setups, the worst-case standard deviation is used. This uncertainty is assumed to have a Gaussian distribution.

# 10.2.2.2.3.3-28	Phase Recovery Non-Linearity over signal bandwidth

This uncertainty originates from the non-linearity of the phase recovery for wide band signal. The phase recovery can be due to either phase non-linearity of the receiver and/or the DUT itself. The method to quantify the non-linarites is FFS.

# 10.2.2.2.3.3-29	Phase Curvature

This contribution originates from the finite far field measurement distance, which causes phase curvature across the antenna of UE/reference antenna. At a measurement distance of 2D2/lambda the phase curvature is 22.5 degrees.  The impact of this factor is FFS.

### 10.2.2.2.3.4 Uncertainty assessment

Table 10.2.2.2.3.4-1: Near Field Test Range Uncertainty assessment for eAAS OTA EVM measurement

| UID | Uncertainty source |  | Uncertainty value f ≦ 3GHz | Uncertainty value 3GHz < f ≦ 4.2 GHz | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB] f ≦ 3GHz | Standard uncertainty ui [dB] 3GHz < f ≦ 4.2 GHz |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | Axes Intersection |  | 0 | 0 | Gaussian | 1.00 | 1 | 0.00 | 0.00 |
| 2 | Axes Orthogonality |  | 0 | 0 | Gaussian | 1.00 | 1 | 0.00 | 0.00 |
| 3 | Horizontal Pointing |  | 0 | 0 | Gaussian | 1.00 | 1 | 0.00 | 0.00 |
| 4 | Probe Vertical Position |  | 0 | 0 | Gaussian | 1.00 | 1 | 0.00 | 0.00 |
| 5 | Probe H/V pointing |  | 0 | 0 | Gaussian | 1.00 | 1 | 0.00 | 0.00 |
| 6 | Measurement Distance |  | 0 | 0 | Gaussian | 1.00 | 1 | 0.00 | 0.00 |
| 7 | Amplitude and Phase Drift |  | 0 | 0 | Gaussian | 1.00 | 1 | 0.00 | 0.00 |
| 8 | Amplitude and Phase Noise |  | 0.02 | 0.02 | Gaussian | 1.00 | 1 | 0.02 | 0.02 |
| 9 | Leakage and Crosstalk |  | 0 | 0 | Gaussian | 1.00 | 1 | 0.00 | 0.00 |
| 10 | Amplitude Non-Linearity |  | 0.04 | 0.04 | Gaussian | 1.00 | 1 | 0.04 | 0.04 |
| 11 | Amplitude and Phase Shift in rotary joints |  | 0 | 0 | Gaussian | 1.00 | 1 | 0.00 | 0.00 |
| 12 | Channel Balance Amplitude and Phase |  | 0 | 0 | Gaussian | 1.00 | 1 | 0.00 | 0.00 |
| 13 | Probe Polarization Amplitude and Phase |  | 0.0001 | 0.0001 | Gaussian | 1.00 | 1 | 0.00 | 0.00 |
| 14 | Probe Pattern Knowledge |  | 0 | 0 | Gaussian | 1.00 | 1 | 0.00 | 0.00 |
| 15 | Multiple Reflections |  | 0 | 0 | Gaussian | 1.00 | 1 | 0.00 | 0.00 |
| 16 | Room Scattering |  | 0.09 | 0.09 | Gaussian | 1.00 | 1 | 0.09 | 0.09 |
| 17 | DUT support Scattering |  | 0 | 0 | Gaussian | 1.00 | 1 | 0.00 | 0.00 |
| 18 | Scan Area Truncation |  | 0.003 | 0.03 | Gaussian | 1.00 | 1 | 0.00 | 0.00 |
| 19 | Sampling Point Offset |  | 0.0058 | 0.0058 | Gaussian | 1.00 | 1 | 0.01 | 0.01 |
| 20 | Spherical Mode Truncation |  | 0.015 | 0.015 | Gaussian | 1.00 | 1 | 0.02 | 0.02 |
| 21 | Positioning |  | 0.03 | 0.03 | Rectangular | 1.73 | 1 | 0.02 | 0.02 |
| 22 | Probe Array Uniformity |  | 0.055 | 0.055 | Gaussian | 1.00 | 1 | 0.06 | 0.06 |
| 23 | Mismatch of receiver chain |  | 0.284 | 0.284 | U-Shaped | 1.41 | 1 | 0.20 | 0.20 |
| 24 | Insertion loss of receiver chain |  | 0 | 0 | Gaussian | 1.00 | 1 | 0.00 | 0.00 |
| 25 | Uncertainty of the absolute gain of the probe antenna |  | 0 | 0 | Gaussian | 1.00 | 1 | 0.00 | 0.00 |
| 26 | RF measurement equipment |  | TBD | TBD | Rectangular | 1.73 | 1 | TBD | TBD |
| 27 | Measurement repeatability – positioning repeatability |  | 0.15 | 0.15 | Gaussian | 1.00 | 1 | 0.15 | 0.15 |
| 28 | Phase Recovery Non-Linearity over signal bandwidth |  | TBD | TBD | TBD | TBD | 1 | TBD | TBD |
| 29 | Phase Curvature |  | TBD | TBD | U-shaped | 1.41 | 1 | TBD | TBD |
|  |  | Combined standard uncertainty (1σ) [dB] |  |  |  |  |  | TBD | TBD |
|  |  | Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  | TBD | TBD |

<< end of proposal >>

# Conclusion

Text proposal to TR 37.843 is presented in order to add the test procedure and MU for OTA EVM measurement in a Near Field test range.

# Reference

[1] R4-1803039, “Test Procedure for OTA EVM measurement in a Near Field Test Range”, Feb 2018

[2] 3GPP TR 37.842 v 13.2.0, March 2017