## Slide 1: Way Forward on Measurement Grids for non sparse antenna arrays
Rohde&Schwarz, Anritsu, Intel, Keysight Technologies
3GPP TSG-RAN WG4 Meeting #86bis Melbourne, Australia, 16th – 20th April, 2018
Agenda item: 7.12
R4-1805995

## Slide 2: Background
A total of three measurement grids are currently considered
  Beam Peak Search Grid: using this grid, the TX and RX beam peak direction will be determined using 3D EIRP (for TX beam peak direction) and 3D RSRP measurements (for RX beam peak directions). A sample procedure is outlined in [R4-1804444]
  CDF/spherical coverage Grid: using this grid, the CDF of EIRP values in 3D is calculated to determine the spherical coverage performance. A review of CDF/spherical coverage grid types and requirements are provided in [R4-1804464]
  TRP Measurement Grid: using this grid, the total power radiated by the DUT is determined by integrating the EIRP measurements taken on the sampling grid. An overview of TRP measurement grids is provided in [R4-1805086] and [R4-1804463]

## Slide 3: Measurement Grid Types
Two measurement grid types have been studied in more detail
  Constant Step Size: the azimuth and elevation angles are uniformly distributed
  Constant Density: measurement points are evenly distributed on the surface of the sphere with a constant density
Proposal 1: Focus on the above two measurement grid types for the permitted UE RF methods to determine the minimum number of grid points.
Proposal 2: The choice and implementation of measurement grid types is left up to the system vendor. It is not precluded to utilize additional measurement points on top of those two grid types.

## Slide 4: Beam Peak Search Grid
Proposal 3: Consider TX or RX beam peak directions only for beam steering directions towards DL signals presented to the UE

## Slide 5: Beam Peak Search Grid
A very fine measurement grid is required to determine the beam peak direction in which a large amount of TX and RX conformance test cases are performed at. The consequences of measuring in a direction other than the true beam peak direction could be significant in terms of test time and cost.
In order to guarantee that labs determine the same beam peak directions, it is proposed to define a minimum number of measurement points sufficient to capture the correct beam peak direction
Proposal 4: For the beam peak search grid for smartphone UEs, assume the [8x2] single panel antenna array outlined in [R4-1804463] to determine the minimum number of measurements points for each grid type. Antenna array assumptions for other DUT types are TBD
Proposal 5: The minimum number of measurements points shall guarantee that the EIRP deviation between the beam peak and the 4 closest, neighboring measurement grid points is at most [0.25dB] for all DUT types. The simulations shall be performed with the beam peak steered towards the equator for constant step size grids
Proposal 6: A measurement grid with more points than the minimum number of measurement points is not precluded
beam peak
4 closest, neighboring grid points to beam peak

## Slide 6: CDF/spherical coverage Grid
Since the TX beam peak search procedure captures a very large number of EIRP measurements in 3D to determine the TX beam peak direction used for a large variety of TX conformance test cases, it is proposed to use those measurements for the spherical coverage tests, i.e., perform the EIRP CDF analysis on the EIRP measurements used to determine the TX beam peak direction.
Proposal 7: Re-use the EIRPs measured using the agreed measurement grid types (Proposal 1) during the TX beam peak direction search procedure for the spherical coverage test.
For adaptive measurement grids, any additional measurement points on top of the previously agreed measurement grid types (Proposal 1) shall not be considered for the CDF calculation
Proposal 8: When using constant step size measurement grids, a theta-dependent correction shall be applied, i.e., the PDF probability contribution for each measurement point is scaled by sin(theta).
Proposal 9: Measurement points added on top of the two grid types shall not be used for the EIRP CDF

## Slide 7: TRP Measurement Grid
Proposal 10: The uncertainty due to the TRP sampling grid shall be added to the MU budget for TRP measurements
Proposal 11: The standard deviation for the TRP measurement grid shall not exceed 0.25dB for all DUT types
Proposal 12: The framework to determine the standard deviation for each sampling grid shall be based on the framework introduced in [R4-1804463], i.e., the standard deviation is derived from [10000] random orientations of the simulated antenna array relative to the measurement grid
When the beam peak is directed to the equator of a constant step size measurement grid, the standard deviation/MU is minimized. It is therefore proposed:
Proposal 13: For constant step-size measurement grids, the beam peak shall be directed towards the equator (by repositioning the UE or coordinate system transformation) to minimize the offset between mean TRP and expected TRP unless a correction is identified for the TRP mean error caused by the sin(theta) term in the TRP equation for the constant step size grid

## Slide 8: TRP Measurement Grid
To determine the minimum number of measurement points for each measurement grid type and for the max TRP measurement grid standard deviation (proposal 11), an assumption for antenna array architecture (in terms of directivity and HPBW) needs to be made. Similar to proposal 4, it is proposed:
Proposal 14: For the TRP measurement grid for smartphone UEs, assume the [8x2] antenna array outlined in [R4-1804463] to determine the minimum number of measurements points for each grid type. Antenna array assumptions for other DUT types are TBD
Proposal 15: Define the minimum number of points per grid type based on maximum TRP measurement grid standard deviation (proposal 11) and the assumed antenna array architecture (proposal 14)
Proposal 16: A measurement grid with more points than the minimum number of measurement points is not precluded

## Slide 9: Supporting Information
Antenna Array Simulation Assumptions [R4-1804463]