3GPP TSG-RAN WG4 Meeting #88	R4-1811136

Gothenburg, Sweden, 20th – 24th August 2018

Source:	Nokia, Nokia Shanghai Bell

Title:	On measurement uncertainty for OTA AAS BS maximum output power

Agenda item:	6.1.3.3.2

Document for:	Discussion

# 1	Introduction

During the RAN4#AH-1807 meeting, the way forward on TRP systematic error and measurement uncertainty budget [1] was approved, which captured some agreements and a list of open issues to be addressed at the forthcoming meetings.

This document discusses measurement uncertainty contributions for OTA AAS BS maximum output power test requirement in CATR chambers. It introduces new uncertainty contributors in addition to the Rel-13 EIRP measurement uncertainty and provides a complete list of uncertainty budget. Finally, a value for the total measurement uncertainty is proposed.

# 2	Discussion

## 2.1 	Uncertainty budget format

From the way forward [1], the following agreements were made:

Proposal 3: adopt Rel-13 EIRP MU as the MU per point for in band TRP test requirements (OTA BS maximum output power) for frequency ranges, f ≤ 3 GHz and 3 GHz < f ≤ 4.2 GHz

Proposal 5: in addition to the MU per point, consider to adopt the following uncertainty contributor and revise the existing CATR contributor “Misalignment DUT & pointing error”:

| Uncertainty contributor | Description |
| --- | --- |
| Test system frequency flatness | This uncertainty is coming from frequency interpolation error caused by a finite frequency resolution during the calibration stage. |

| TRP uncertainty budget |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| UID | Uncertainty Source | Uncertainty value f ≦ 3GHz | Uncertainty value 3GHz ≦ f < 4.2 GHz | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB] f ≦ 3GHz | Standard uncertainty ui [dB] 3GHz < f ≦ 4.2 GHz |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |
| 1 | Misalignment DUT & pointing error | [0.3] | [0.3] | Rectangular | √3 | 1 | [0.174] | [0.174] |
| 9 | Test system frequency flatness | [0.25] | [0.25] | Gaussian | 1 | 1 | [0.25] | [0.25] |

Herein, we explain the rationale behind the above Proposal 5.

Misalignment DUT & pointing error

The following geometric equation describes the expected error contribution resulting from polarization mismatch that this uncertainty needs to address

10log10cos2(offset_angle)

A typical variation in polarization purity occurs at offset_angle = ±15° and the uncertainty is assumed to be rectangularly distributed.

Test system frequency flatness

This uncertainty is coming from frequency interpolation error caused by using a finite frequency resolution during the calibration stage. Figure 1 shows the frequency response of two TCH traffic channels in the n77 band; the bandwidth of each channel is 100 MHz. As can be observed, the response is not flat across the entire bandwidth of the two TCH channels as compared with the ideal frequency response indicated by black lines. The green dot is the centre frequency of TCH channels and the red dot is the corresponding measured OTA path loss value which is normalized to the nominal reference. The difference between the two dots is the error in the OTA test system frequency flatness, which is dependent on the accuracy of the calibration stage and quality of the OTA system quiet zone.

TCH Channel

TCH Channel

Normalized Channel Response (dB)

Frequency (MHz)

Figure 1: Frequency response

If one single discrete frequency point (i.e., the middle frequency of the TCH) is used to compensate for the measured power value, then there is a power error measured at frequencies other than the middle frequency but it is dependent on the selected points. An infinite number of discrete frequency points are needed to obtain a power error ≈ 0 dB in OTA test system frequency flatness, which is not practical. Simulations were carried out to study the effect of the number of frequency points on error variations. Table 1 shows the simulation results with different discrete frequency points. As shown in the table, the maximum and minimum deviations with resepect to the ideal frequency response are 0.64 dB and 0.38 dB, respectively. The average error variation is approximately 0.5 dB and the median is 0.49 dB. It can be observed that 50% of the error variations are within ±0.5 dB for all the of the N=51 frequency points. Therefore, an error of 0.25 dB is assumed. This error is converted to a standard measurement uncertainty contribution using a Gaussion distribution.

Table 1: OTA Test system frequency flatness versus the number of frequency discrete points

| Number of discrete frequency points during OTA calibration phase (N) | n77 frequency band with 2 TCH channels |  |
| --- | --- | --- |
|  | OTA Test system frequency flatness |  |
|  | Bottom TCH Channel | Middle TCH Channel |
| 1 (centre frequency) | 0.44 | 0.46 |
| 3 (30 MHz intervals) | 0.51 | 0.59 |
| 6 (20 MHz intervals) | 0.38 | 0.60 |
| 11 (10 MHz intervals) | 0.42 | 0.58 |
| 23 (4 MHz intervals) | 0.39 | 0.64 |
| 51 (2 MHz intervals) | 0.38 | 0.64 |

The uncertainty budget format for Rel-13 EIRP measurement uncertainty for CATR is outlined in TR 37.842, Section 10.3.1.1.3. Referring to the agreed Proposal 3 in [1], the uncertainty budget for OTA AAS BS maximum output power test requirement includes all Rel-13 EIRP uncertainty contributors.

With reference to the agreed Proposal 5 in [1], we propose to add the “Test system frequency flatness” contributor to the uncertainty budget for the OTA AAS BS maximum output power as shown in Table 2.

Table 2: CATR uncertainty budget format for OTA AAS BS maximum output power

| UID | Description of uncertainty contribution | Details in annex |
| --- | --- | --- |
| Stage 2: DUT measurement |  |  |
| 1 | Misalignment DUT & pointing error | B2-1 |
| 2 | RF power measurement equipment | E |
| 3 | Standing wave between DUT and test range antenna | B2-3 |
| 4 | RF leakage (SGH connector terminated & test range antenna connector cable terminated) | B2-4 |
| 5 | QZ ripple DUT | B2-5 |
| 18 | Test system frequency flatness | TBD |
| 19 | Miscellaneous uncertainty | B2-14 |
| Stage 1: Calibration measurement |  |  |
| 6 | Uncertainty of network analyser | E |
| 7 | Mismatch of receiver chain | B2-7 |
| 8 | Insertion loss variation of receiver chain | B2-8 |
| 9 | RF leakage, (SGH connector terminated & test range antenna connector cable terminated) | B2-4 |
| 10 | Influence of the calibration antenna feed cable: a) Flexing cables, adapters, attenuators, connector repeatability | B2-9 |
| 11 | Uncertainty of the absolute gain of the calibration antenna | E |
| 12 | Misalignment positioning system | B2-11 |
| 13 | Misalignment of calibration antenna and test range antenna | B2-1 |
| 14 | Rotary Joints | B2-12 |
| 15 | Standing wave between reference calibration antenna and test range antenna | B2-3 |
| 16 | Quality of quiet zone | B2-5 |
| 17 | Switching uncertainty | B2-15 |

## 2.2	Uncertainty calculation

The uncertainty calculation for OTA BS maximum output power is given in Table 3. All of the uncertainty contributors except “Test system frequency flatness” are adopted from Rel-13 EIRP measurements for CATR outlined in TR 37.842, Section 10.3.1.1.3. The value of the “Misalignment DUT & pointing error” will be updated based on the outcome of the above analysis.

Table 3: CATR uncertainty calculation for OTA AAS BS maximum output power

| OTA AAS BS maximum output power uncertainty budget |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| UID | Uncertainty Source | Uncertainty value f ≦ 3GHz | Uncertainty value 3GHz ≦ f < 4.2 GHz | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB] f ≦ 3GHz | Standard uncertainty ui [dB] 3GHz < f ≦ 4.2 GHz |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |
| 1 | Misalignment DUT & pointing error | 0.3 | 0.3 | Rectangular | √3 | 1 | 0.174 | 0.174 |
| 2 | RF power measurement equipment (e.g. spectrum analyzer, power meter) | 0.14 | 0.26 | Gaussian | 1 | 1 | 0.14 | 0.26 |
| 3 | Standing wave between DUT and test range antenna | 0.21 | 0.21 | U-shaped | √2 | 1 | 0.15 | 0.15 |
| 4 | RF leakage, test range antenna cable connector terminated. | 0.0012 | 0.0012 | Normal | 1 | 1 | 0.0012 | 0.0012 |
| 5 | QZ ripple with DUT | 0.0928 | 0.0928 | Normal | 1 | 1 | 0.0928 | 0.0928 |
| TBD | Test system frequency flatness | 0.25 | 0.25 | Gaussian | 1 | 1 | 0.25 | 0.25 |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |  |
| 6 | Network Analyzer | 0.13 | 0.20 | Normal | 1 | 1 | 0.13 | 0.20 |
| 7 | Uncertainty of return loss (S11) measurement of SGH and test receiver (VNA) ports | 0.127 | 0.325 | U-shaped | √2 | 1 | 0.09 | 0.23 |
| 8 | Insertion loss variation in receiver chain | 0.18 | 0.18 | Rectangular | √3 | 1 | 0.10 | 0.10 |
| 9 | RF leakage, test range antenna cable connector terminated. | 0.0012 | 0.0012 | Normal | 1 | 1 | 0.0012 | 0.0012 |
| 10 | Influence of the calibration antenna feed cable | 0.022 | 0.022 | U-shaped | √2 | 1 | 0.015 | 0.015 |
| 11 | SGH Calibration uncertainty | 0.50 | 0.433 | Rectangular | √3 | 1 | 0.29 | 0.25 |
| 12 | Misalignment positioning system | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 |
| 13 | Misalignment SGH and pointing error | 0.5 | 0.5 | Exp. normal | 2 | 1 | 0.25 | 0.25 |
| 14 | Rotary joints | 0.048 | 0.048 | U-shaped | √2 | 1 | 0.034 | 0.034 |
| 15 | Standing wave between SGH and test range antenna | 0.09 | 0.09 | U-shaped | √2 | 1 | 0.06 | 0.06 |
| 16 | QZ ripple with SGH | 0.009 | 0.009 | Normal | 1 | 1 | 0.009 | 0.009 |
| 17 | Switching uncertainty | 0.26 | 0.26 | Rectangular | √3 | 1 | 0.15 | 0.15 |
| Combined standard uncertainty (1σ) [dB] |  |  |  |  |  |  | 0.59 | 0.67 |

According to the agreed Proposal 4 in [1], the total uncertainty MUtotal is calculated as

MUtotal=MUper-point2+TRPsyserr2, where

MUper-point= uc which is given in Table 3 and TRPsyserr = 0.75 dB.

The total uncertainty values are given in Table 4.

Table 4: CATR total uncertainty calculation for OTA AAS BS maximum output power

|  | Uncertainty value [dB] f ≤ 3 GHz | Uncertainty value [dB] 3 GHz < f ≤ 4.2 GHz |
| --- | --- | --- |
| MUper-point, uc | 0.59 | 0.67 |
| TRPsyserr | 0.75 | 0.75 |
| Total uncertainty, MUtotal=uc2+TRPsyserr2 [dB] | 0.96 | 1.00 |
| 95% confidence interval, total uncertainty MUtotal=1.96MUper-point2+TRPsyserr2 [dB] | 1.88 | 1.96 |

# 3	Conclusions

The document has discussed measurement uncertainty contributions for OTA AAS BS maximum output power test requirement in CATR chambers. In addition to the Rel-13 EIRP measurement uncertainty contributors, one new uncertainty contributor so-called “Test system frequency flatness” is added to the list of uncertainty budget. The total measurement uncertainty values are as follows:

|  | Uncertainty value [dB] f ≤ 3 GHz | Uncertainty value [dB] 3 GHz < f ≤ 4.2 GHz |
| --- | --- | --- |
| MUper-point, uc | 0.59 | 0.67 |
| TRPsyserr | 0.75 | 0.75 |
| Total uncertainty, MUtotal=uc2+TRPsyserr2 [dB] | 0.96 | 1.00 |
| 95% confidence interval, total uncertainty MUtotal=1.96MUper-point2+TRPsyserr2 [dB] | 1.88 | 1.96 |

# References

- R4-1809436, Way forward on TRP systematic error and measurement uncertainty budget, Nokia, Nokia Shanghai Bell, NTT DoCoMo

- TR 37.843 v15.0.0, RF requirement background for AAS BS radiated requirements (Release 15)