## Slide 1: Way Forward on Measurement Grids for Beam Peak Search & Spherical Coverage for PC3 UEs
Keysight Technologies, Rohde&Schwarz, Apple, EMITE Ing.
3GPP TSG-RAN WG4 Meeting #88bis
Chengdu, China, 8 – 12 October 2018 Agenda Item: 10.1.2.1
R4-1816262

## Slide 2: Background
Two measurement grids are considered
  Beam Peak Search Grid: using this grid, the TX and RX beam peak direction will be determined using 3D EIRP (for TX beam peak direction) and 3D RSRP/Throughput/EIS measurements (for RX beam peak directions).
  Spherical Coverage Grid: using this grid, the CDF of EIRP/EIS values in 3D is calculated to determine the spherical coverage performance.

## Slide 3: MU Elements: Normal Distribution
For a normal distribution,
  determine the mean error, μ, from the reference value and the standard deviation, σ.
  Proposal 1: In the MU budget for the respective test cases, add the
    mean error as a systematic error
    Standard deviation as an MU element with ‘actual’ distribution of the probability (with divisor 1) in the DUT Measurement Stage

## Slide 4: MU Element: Half-Normal Distribution
For a half-normal distribution
  Determine the offset from the reference that contains 95% of the PDF distribution/at which the CDF is 5%, Offset5%CDF
  Proposal 2: In the MU budget for the respective test cases, add the
    Offset5%CDF as a systematic error
Offset from Reference

## Slide 5: Simulation Assumptions for Beam Peak Search
The previously agreed approach to determine the minimum number of grid points for the beam peak search agreed in [R4-1805995] based on the maximum deviation between closest, neighboring grid points is flawed as outlined in [R4-1814420]
  The minimum number of grid points were overestimated by a factor of ~4
  Instead, the MU for Beam Peak Search shall be based on a statistical approach
  Proposal 3: MU for Beam Peak Search shall be based on a statistical approach
For the simulations, the 8x2 reference antenna array is used. A statistical significant number of relative orientations between the simulated antenna array and the beam peak search grid was altered randomly. The simulation assumptions of the rotations shall be the same as those outlined in Annex G.1.1 of [TR38.810].

## Slide 6: Distribution of Beam Peak Error
The distribution of the Beam Peak Error was found to be a half-normal distribution
  The Offset5%CDF needs to be determined and added to the MU budget for EIRP MOP and EIS REFSENS test cases as systematic error

## Slide 7: Beam Peak Search MUs
Simulations for beam peak search MUs were presented in [R4-1814420, revised in R4-1816261] for constant-step size and constant density (charged particle) grids
Proposal 4: Consider beam peak search grids with a systematic error of “Beam Peak Search” of 0.5dB, i.e.,
  Constant density grid with at least 800 grid points
  Constant step size grid with at least 1106 grid points
Systematic Error of ‘Beam Peak Search’: Offset from Beam Peak at which CDF is 5% | Minimum Number of Unique Grid Points for Constant Step Size Grid | Minimum Number of Unique Grid Points for Constant Density Grid
0.2dB | 2522 (5o step size) | 2000
0.3dB | 1742 (6o step size) | 1500
0.4dB | N/A | 1000
0.5dB | 1106 (7.5ostep size) | 800
0.6dB | N/A | 750
0.7dB | 762 (9o step size) | 600

## Slide 8: Simulation Assumptions for Spherical Coverage
The beamformer assumptions were captured in [R4-1813581]
Orientations/Rotations:
  10000 random relative orientations between the simulated UE and the respective measurement grids will be analysed
  The rotations of UE/grid will be along q and f as well as around the beam peak
  The rotations along q will utilize a sin(q) weighting to assume a uniform sampling on the surface
When using constant step size measurement grids, a theta-dependent correction shall be applied, i.e., the PDF probability contribution for each measurement point is scaled by sin(theta)
Proposal 5: These simulation assumptions need to be added to Annex G of [38.810]

## Slide 9: Definition of EIRP50%CDF
The definition of the EIRP at the target CDF (for PC3: 50%) was agreed in [R4-1814419] for cases where CDF is flat over EIRP ranges
From [R4-1814419]:
Proposal: For the case when the target CDF is not met with any EIRP values, it is proposed to determine the min. EIRP at the target CDF based on an interpolation of the CDF curve between the raising edges located right above the CDF target and right below the target. For the case where the target CDF is met with one more or EIRP value, define the min. EIRP at the target CDF as the min. EIRP value that meets the CDF target.

## Slide 10: EIRP Spherical Coverage Results
EIRP Spherical Coverage Simulations were presented in [R4-1814813], [R4-1814837], and [R4-1815394].
The distribution of EIRP at the 50% CDF is a normal distribution
  The Mean Error (with respect to the reference CDF) and the standard deviation of the EIRP distribution shall be added to the MU budget as systematic error and as MU element in DUT measurement stage, respectively, for the EIRP spherical coverage test cases

## Slide 11: EIRP Spherical Coverage Results
EIRP Spherical Coverage Simulations presented in [R4-1814813], [R4-1814837] show that the EIRP beam peak does not have to be aligned on the spherical coverage grid
The following agreement was captured in the chairman notes:
  Agreement: For PC3 UE, allow the EIRP spherical coverage measurement to be performed without having to have the beam peak having to be placed on a grid point, e.g., for coarse grids of beam peak searches.

## Slide 12: EIRP Spherical Coverage MU
Proposal 6: Require EIRP spherical coverage grids with at least 200 unique measurement points, i.e.,
  constant density grid with at least 200 grid points: STD of 0.11dB and 0dB Mean Error
  constant step size grid with at least 266 grid points: STD of 0.12dB and 0dB Mean Error

## Slide 13: EIS Spherical Coverage Results
EIS Spherical Coverage Simulations were presented in [R4-1814838]
The distribution of EIS at the 50% CDF is a normal distribution
  The Mean Error (with respect to the reference CDF) and the standard deviation of the EIS distribution shall be added to the MU budget as systematic error and MU element in DUT measurement stage, respectively, for the EIS spherical coverage test cases

## Slide 14: EIS Spherical Coverage Results
EIS Spherical Coverage Simulations presented in [R4-1814838] show that the EIS beam peak does not have to be aligned on the spherical coverage grid
The following agreement was captured in the chairman notes:
  Agreement: For PC3 UE, allow the EIS spherical coverage measurement to be performed without having to have the beam peak having to be placed on a grid point, e.g., for coarse grids of beam peak searches.

## Slide 15: EIS Spherical Coverage Results
EIS Spherical Coverage Simulations presented in [R4-1814838] show that the Mean Error corresponds to the DL power step size of the final EIS search procedure

## Slide 16: EIS Spherical Coverage MU
Proposal 7: Require EIS spherical coverage grids with at least 200 unique measurement points, i.e.,
  constant density grid with at least 200 grid points: STD of 0.11dB and [DL power step size] Mean Error
  constant step size grid with at least 266 grid points: STD of 0.12dB and [DL power step size] Mean Error

## Slide 17: Combined Beam Peak and Spherical Coverage Analyses
Since the beam peak does not have to be aligned on the spherical coverage grid for spherical coverage analyses, the EIRP/EIS results from the EIRP/EIS beam peak search could be re-used for EIRP/EIS spherical coverage which could further reduce the EIRP/EIS spherical coverage MU
Proposal 8: The EIRP/EIS results from the EIRP/EIS beam peak search with the measurement grid proposed in Proposal 4 could be re-used for EIRP/EIS spherical coverage
Coarse EIRP/EIS beam peak searches can be used for EIRP/EIS spherical coverage analyses as long as they meet the minimum number of measurement points (200 grid points for constant density and 266 for constant step size)
  Subsequent fine search approaches using constant step size grids used to determine the EIRP/EIS beam peaks require a minimum step size of 7.5o
  Subsequent fine search approaches using constant density grids used to determine the EIRP/EIS beam peaks require at least 6 points around the peak identified with the coarse grid, with a spacing corresponding to a constant density over the whole sphere of at least 800 points.
Proposal 9: Coarse&fine beam peak search approaches could be used to further reduce test time
Proposal 10: The EIRP/EIS spherical coverage performance can be determined from the corresponding EIRP/EIS coarse beam peak search

## Slide 18: Coarse & Fine Measurement Grids
Coarse & Fine EIRP/EIS beam peak measurement grid proposals have been discussed in [R4-1813582], [R4-1813583], [R4-1813273]
  The coarse grid has fewer number of points than proposed in Proposal 4
  The fine grid has at a minimum the number of points proposed in Proposal 4
The metrics for coarse & fine, single fine measurement grids have been captured in the AH meeting notes:
    Metrics used for the Tx beam peak search:
      Single fine grid approach: EIRP is used
      Coarse & Fine grid approach: EIRP is used for both grids
    Metrics used for the Rx beam peak search:
      Single fine grid approach: EIS metric is used
      Coarse & Fine grid approach:
        EIS/Throughput metric is used for Coarse grids;
        EIS metric is used for Fine grids

## Slide 19: Illustration: Coarse& Fine Constant Step Size Grids
Coarse Grid
Fine Grid
Used Points of Fine Grid

## Slide 20: Illustration: Coarse& Fine Constant Density Grids
Coarse Grid
Fine Grid
Used Points of Fine Grid