3GPP TSG-RAN4 Meeting #92	R4-1910434

# Ljubljana, Slovenia, 26th Aug 2019 - 30th Aug 2019

| CR-Form-v12.0 |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CHANGE REQUEST |  |  |  |  |  |  |  |  |
|  | 37.843 | CR | 0029 | rev | 1 | Current version: | 15.4.0 |  |
| For HELP on using this form: comprehensive instructions can be found at http://www.3gpp.org/Change-Requests. |  |  |  |  |  |  |  |  |

| Proposed change affects: | UICC apps |  | ME |  | Radio Access Network |  | Core Network |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |

| Title: | CR to TR 37.843: Addition of 2D Compact Range. MU, calibration and test procedure for EIRP and EIS. |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Source to WG: | Rohde & Schwarz, CAICT, MVG |  |  |  |  |  |  |  |  |  |
| Source to TSG: | R4 |  |  |  |  |  |  |  |  |  |
| Work item code: | AASenh_BS_LTE_UTRA-Perf |  |  |  |  |  | Date: |  |  | 2019-08-29 |
| Category: | F |  |  |  |  |  | Release: |  |  | Rel-15 |
|  | Use one of the following categories: F (correction) A (mirror corresponding to a change in an earlier release) B (addition of feature), C (functional modification of feature) D (editorial modification) Detailed explanations of the above categories can be found in 3GPP TR 21.900. |  |  |  |  |  |  |  | Use one of the following releases: Rel-8 (Release 8) Rel-9 (Release 9) Rel-10 (Release 10) Rel-11 (Release 11) Rel-12 (Release 12) Rel-13 (Release 13) Rel-14 (Release 14) Rel-15 (Release 15) Rel-16 (Release 16) |  |
| Reason for change: |  | The addition of the Plane Wave Synthesizer as accepted methodology for radiated testig of AAS BS requires modifications of some clauses. |  |  |  |  |  |  |  |  |
| Summary of change: |  | Plane Wave Synthesizer system description, together with the calibration and test procedure added to the corresponding clauses for EIRP and EIS test cases. The corresponding MU evaluation and MU terms description added. New reference added to clause 2. |  |  |  |  |  |  |  |  |
| Consequences if not approved: |  | Plane Wave Synthesizer could not be used as methodology for AAS BS conformance testing. |  |  |  |  |  |  |  |  |
| Clauses affected: |  | Clauses 2, 10.2.2, 10.3.2 and Annex E |  |  |  |  |  |  |  |  |
|  |  | Y | N |  |  |  |  |  |  |  |
| Other specs |  |  | X | Other core specifications |  |  |  | TS/TR ... CR ... |  |  |
| affected: |  |  | X | Test specifications |  |  |  | TS/TR ... CR ... |  |  |
| (show related CRs) |  |  | X | O&M Specifications |  |  |  | TS/TR ... CR ... |  |  |
| Other comments: |  | - |  |  |  |  |  |  |  |  |

| This CR's revision history: | - Changed name of the methodology to Plane Wave Synthesizer. - Editorial updates to generalize the description of the method, calibration procedure and measurement procedures. - Figures were updated to generalize the test setup description. - Added wording in clause 10.2.2.1 to clarify the addition of Plane Wave Synthesizer as test method. - Added wording in clause 10.3.2 to clarify the addition of Plane Wave Synthesizer as test method. - Updated the list of MU terms in tables 10.2.2.3.4.1-1, 10.2.2.3.4.2-1, 10.3.2.2.3.1-1, 10.3.2.2.3.2-1. - Updated the list of MU terms description in clause E.6. - Removed changes in clause 2. |
| --- | --- |

< Unchanged Text Deleted >

< Beginning of Changes >

### 10.2.2	EIRP accuracy

#### 10.2.2.1	General

The EIRP accuracy uncertainty budget was re-assessed and adjusted slightly according as detailed in the following tables. The MU budget for the CATR is shown as it is this MU that determines the agreed MU and TT.

Plane Wave Synthesizer (PWS) description and MU budget is also shown as an additional methodology to those presented in TR 37.842 [4].

< End of Changes >

< Unchanged Text Deleted >

< Beginning of Changes >

#### 10.2.2.3	Plane Wave Synthesizer

##### 10.2.2.3.1	Description

This method is a natural extension of a compact antenna test range (CATR) where the reflector is replaced by an antenna array, or Plane Wave Synthesizer (PWS), capable of approximating a plane wave within a specified quiet zone in front of the antenna in a similar manner as the reflector in a CATR system.

The principle of the Plane Wave Synthesizer is shown in figure 10.2.2.3.1-1:

Figure 10.2.2.3.1-1: PWS test system diagram

The PWS typically consists of two main components: the radiating elements arranged in the structure of an array and the feeding system. The architecture of the later depends on the adopted technology, e.g. it could be fed by a standard Beam Forming Network (BFN) or by an active system.

Using similar technique of setting the phase and amplitude of the array signals, a quiet zone containing planar waves as a linear superposition of the array elements' spherical radiation waves can be created within the near field of the array.

The PWS can achieve far-field testing conditions in a Quiet Zone (QZ) as in the case of Compact Antenna Test Ranges (CATRs) it enables direct measurements of far-field AUT performance in a controlled indoor environment as an alternative to CATR.

##### 10.2.2.3.2	Procedure

# Stage 1 - Calibration:

The calibration measurement is done by using a reference antenna with known efficiency or gain values (e.g. SGH). In the calibration measurement the reference antenna is measured in the same place as the DUT, and the attenuation of the complete transmission path (C↔A, see figure 10.2.2.3.2-1) from the DUT to the measurement receiver is calibrated out.

Figure 10.2.2.3.2-1: Tx range path loss calibration

1) Path loss calibration E→A:

a)	Measure SGH (or other calibrated reference antenna) reflection coefficient separately at the antenna's connector with a network analyzer (or equivalent measurement equipment) to obtain ΓSGH.

b)	Measure cable loss from point E to input of SGH, call this LE↔SGH which is the equivalent of 20log|S21| from the use of a network analyzer.

c)	Calculate the combined total path loss from E→A by using the following expression:

-	LSGHcal = LE,SGH + 10log(1 - |ΓSGH|2) - GSGH;

-	where10log1-ΓSGH2  10log(1 - |ΓSGH|2)  is the compensation for SGH connector return loss, GSGH  is the known gain of the reference SGH.

2)	Connect SGH and D↔E cable.

3)	Measure path loss C→E with network analyzer LC→E = 20log|S21|.

4)	Calculate the test path loss compensation factor.  This is the total path loss between C↔A using the results from step 1c and 3. L = LSGHcal LSGHcal  - LC→E.

Where ΓSGH is the reflection coefficient (or mismatch) seen at the SGH connector (S11 with a network analyzer).

# Stage 2 - Measurement:

1)	Install the AAS BS as DUT with the manufacturer declared coordinate system reference point in the same place as the phase center of the reference antenna A as shown in figure 10.2.2.3.2-2. The manufacturer declared coordinate system orientation of the AAS BS is set to be aligned with the testing system.

2)	Set the AAS BS to generate the tested beam with the appropriate test model with the beam peak direction intended to be the same as the testing direction.

3)	Measure mean power (Pmeas) at the measurement equipment (such as a spectrum analyzer or power meter) denoted in figure 10.2.2.3.2-2.

4)	Calculate EIRP, where EIRP = Pmeas + LC→A.

5)	Repeat steps 2-4 for all conformance test beam direction pairs and test conditions.

Figure 10.2.2.3.2-2: PWS setup for EIRP measurements

##### 10.2.2.3.4	MU assessment

###### 10.2.2.3.4.1	MU Budget

Table 10.2.2.3.4.1-1: Plane Wave Synthesizer uncertainty contributions for EIRP measurements

| UID | Uncertainty Source | Details in annex |
| --- | --- | --- |
| Stage 2: DUT measurement |  |  |
| 1 | Misalignment DUT & pointing error | E6-1 |
| 2 | RF power measurement equipment | F.1 |
| 3 | Longitudinal position uncertainty (i.e. standing wave and imperfect field synthesis) for DUT antenna | E6-2 |
| 4 | RF leakage (calibration antenna connector terminated) | E6-3 |
| 5 | QZ ripple with DUT | E6-4 |
| 6 | Miscellaneous Uncertainty | E6-5 |
| 20 | System non-linearity | E6-14 |
| Stage 1: Calibration measurement |  |  |
| 7 | Uncertainty of network analyzer | F.1 |
| 8 | Mismatch (i.e. reference antenna, network analyser and reference cable) | E6-6 |
| 9 | Insertion loss variation | E6-7 |
| 4 | RF leakage (calibration antenna connector terminated) | E6-3 |
| 10 | Influence of the calibration antenna feed cable | E6-8 |
| 11 | Uncertainty of the absolute gain of the calibration antenna | F.1 |
| 12 | Misalignment of positioning system | E6-9 |
| 13 | Misalignment of calibration antenna & pointing error | E6-1 |
| 14 | Rotary joints | E6-10 |
| 15 | Longitudinal position uncertainty (i.e. standing wave and imperfect field synthesis) for calibration antenna | E6-2 |
| 16 | QZ ripple with calibration antenna | E6-4 |
| 17 | Switching uncertainty | E6-11 |
| 18 | Field repeatability | E6-12 |
| 19 | Frequency flatness | E6-13 |

###### 10.2.2.3.4.2	MU Value

Table 10.2.2.3.4.2-1: Plane Wave Synthesizer uncertainty assessment for EIRP measurements

| UID | Uncertainty Source | Uncertainty Value |  | Distribution of the probability | Divisor based on distribution shapre | ci | Std. uncertainty ui [dB] |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | f ≦ 3GHz | 3GHz ≦ f < 4.2GHz |  |  |  | f ≦ 3GHz | 3GHz ≦ f < 4.2GHz |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |
| 1 | Misalignment DUT & pointing error | 0.10 | 0.10 | Rectangular | √3 | 1 | 0.06 | 0.06 |
| 2 | RF power measurement equipment | 0.14 | 0.26 | Gaussian | 1 | 1 | 0.14 | 0.26 |
| 3 | Longitudinal position uncertainty (i.e. standing wave and imperfect field synthesis) for DUT antenna | 0.05 | 0.14 | Rectangular | √3 | 1 | 0.03 | 0.08 |
| 4 | RF leakage (calibration antenna connector terminated) | 0.086 | 0.086 | Normal | 1 | 1 | 0.09 | 0.09 |
| 5 | QZ ripple with DUT | [0.4] | [0.4] | Rectangular | √3 | 1 | [0.23] | [0.23] |
| 6 | Miscellaneous Uncertainty | 0.00 | 0.00 | Normal | 1 | 1 | 0.00 | 0.00 |
| 20 | System non-linearity | [0.062] | [0.062] | Rectangular | √3 | 1 | [0.04] | [0.04] |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |  |
| 7 | Uncertainty of network analyzer | 0.13 | 0.20 | Normal | 1 | 1 | 0.13 | 0.20 |
| 8 | Mismatch (i.e. reference antenna, network analyser and reference cable) | 0.127 | 0.325 | U-shaped | √2 | 1 | 0.09 | 0.23 |
| 9 | Insertion loss variation | 0.18 | 0.18 | Rectangular | √3 | 1 | 0.10 | 0.10 |
| 4 | RF leakage (calibration antenna connector terminated) | 0.086 | 0.086 | Normal | 1 | 1 | 0.09 | 0.09 |
| 10 | Influence of the calibration antenna feed cable | 0.103 | 0.104 | Rectangular | √3 | 1 | 0.06 | 0.06 |
| 11 | Uncertainty of the absolute gain of the calibration antenna | 0.50 | 0.43 | Rectangular | √3 | 1 | 0.29 | 0.25 |
| 12 | Misalignment of positioning system | 0.00 | 0.00 | Exp. normal | 2 | 1 | 0.00 | 0.00 |
| 13 | Misalignment of calibration antenna & pointing error | 0.05 | 0.05 | Rectangular | √3 | 1 | 0.03 | 0.03 |
| 14 | Rotary joints | 0.00 | 0.00 | U-shaped | √2 | 1 | 0.00 | 0.00 |
| 15 | Longitudinal position uncertainty (i.e. standing wave and imperfect field synthesis) for calibration antenna | 0.12 | 0.12 | Rectangular | √3 | 1 | 0.07 | 0.07 |
| 16 | QZ ripple with calibration antenna | [0.2] | [0.2] | Rectangular | √3 | 1 | [0.12] | [0.12] |
| 17 | Switching uncertainty | 0.02 | 0.02 | Rectangular | √3 | 1 | 0.01 | 0.01 |
| 18 | Field repeatability | [0.003] | [0.003] | Normal | 1 | 1 | [0.003] | [0.003] |
| 19 | Frequency flatness | 0.13 | 0.13 | Rectangular | √3 | 1 | 0.08 | 0.08 |
| Combined standard uncertainty (1σ) [dB] |  |  |  |  |  |  | [0.49] | [0.58] |
| Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  |  | [0.96] | [1.14] |

< End of Changes >

< Unchanged Text Deleted >

< Beginning of Changes >

### 10.3.2	OTA sensitivity

#### 10.3.2.1	General

The OTA sensitivity MU was derived in TR 37.842 [4] as shown in Table 10.3.2-1. Plane Wave Synthesizer description and MU budget is also shown as an additional methodology to those presented in TR 37.842 [4].

Table 10.3.2-1: Test system specific measurement uncertainty values for the EIS test

|  | Expanded uncertainty ue [dB] |  |
| --- | --- | --- |
|  | f ≦ 3GHz | 3GHz < f ≦ 4.2 GHz |
| Indoor Anechoic Chamber | 1.22 | 1.25 |
| Compact Antenna Test Range | 1.33 | 1.40 |
| One Dimensional Compact Range Chamber | 1.29 | 1.43 |
| Near Field Test Range | 1.24 | 1.24 |
| Plane Wave Synthesizer | [1.29] | [1.37] |
| Common maximum accepted test system uncertainty | 1.3 | 1.4 |

#### 10.3.2.2	Plane Wave Synthesizer

##### 10.3.2.2.1	Procedure

The calibration measurement is done by using a reference antenna with known efficiency or gain values (e.g. SGH). In the calibration measurement the reference antenna is measured in the same place as the DUT, and the attenuation of the complete transmission path (C↔A, see figure 10.3.2.2.1-1) from the DUT to the measurement receiver is calibrated out.

Figure 10.3.2.2.1-1: Rx range path loss calibration

1)	Path loss calibration E→A:

a)	Measure SGH (or other calibrated reference antenna) reflection coefficient separately at the antenna's connector with a network analyzer (or equivalent measurement equipment) to obtain ΓSGH ΓSGH. .

b)	Measure cable loss from point E to input of SGH, call this LC↔SGH LE↔SGH  which is the equivalent of 20log|S21| from the use of a network analyzer.

c)	Calculate the combined total path loss from E→A by using the following expression

-	LSGHcal = LE,SGH + 10log(1 - |ΓSGH|2) - GSGH;

-	where10log1-ΓSGH2  10log(1 - |ΓSGH|2)  is the compensation for SGH connector return loss, GSGH  is the known gain of the reference SGH.

2)	Connect SGH and D↔E cable.

3)	Measure path loss C→E with network analyzer LC→E = 20log|S21|.

4)	Calculate the test path loss compensation factor. This is the total path loss between C↔A using the results from step 1c and 3. L = LSGHcal LSGHcal  - LC→E.

Where ΓSGH is the reflection coefficient (or mismatch) seen at the SGH connector (S11 with a network analyzer).

# Stage 2 - Measurement:

1)	Install AAS BS as DUT with the manufacturer declared coordinate system reference point in the same place as the phase center of the reference antenna A as shown in figure 10.3.2.2.1-2. The manufacturer declared coordinate system orientation of the AAS BS is set to be aligned with the testing system.

2)	Configure signal generator, one port and one carrier at a time according to maximum power requirement.

3)	Calculate EIS at this point with EIS = PBER - LC→A.

4)	Repeat steps 2-3 for all conformance test beam direction pairs and test conditions.

Figure 10.3.2.2.1-2: PWS setup for Sensitivity measurements

##### 10.3.2.2.3	MU assessment

###### 10.3.2.2.3.1	MU Budget

Table 10.3.2.2.3.1-1: Plane Wave Synthesizer uncertainty contributions for EIS measurements

| UID | Uncertainty Source | Details in annex |
| --- | --- | --- |
| Stage 2: DUT measurement |  |  |
| 1 | Misalignment DUT & pointing error | E6-1 |
| 2 | RF signal generator | F.1 |
| 3 | Longitudinal position uncertainty (i.e. standing wave and imperfect field synthesis) for DUT antenna | E6-2 |
| 4 | RF leakage (calibration antenna connector terminated) | E6-3 |
| 5 | QZ ripple with DUT | E6-4 |
| 6 | Miscellaneous Uncertainty | E6-5 |
| 20 | System non-linearity | E6-14 |
| Stage 1: Calibration measurement |  |  |
| 7 | Uncertainty of network analyzer | F.1 |
| 8 | Mismatch (i.e. reference antenna, network analyzer and reference cable) | E6-6 |
| 9 | Insertion loss variation | E6-7 |
| 4 | RF leakage (calibration antenna connector terminated) | E6-3 |
| 10 | Influence of the calibration antenna feed cable | E6-8 |
| 11 | Uncertainty of the absolute gain of the calibration antenna | F.1 |
| 12 | Misalignment of positioning system | E6-9 |
| 13 | Misalignment of calibration antenna & pointing error | E6-1 |
| 14 | Rotary joints | E6-10 |
| 15 | Longitudinal position uncertainty (i.e. standing wave and imperfect field synthesis) for calibration antenna | E6-2 |
| 16 | QZ ripple with calibration antenna | E6-4 |
| 17 | Switching uncertainty | E6-11 |
| 18 | Field repeatability | E6-12 |
| 19 | Frequency flatness | E6-13 |

###### 10.3.2.2.3.2	MU Value

Table 10.3.2.2.3.2-1: Plane Wave Synthesizer uncertainty assessment for EIS measurements

| UID | Uncertainty Source | Uncertainty Value |  | Distribution of the probability | Divisor based on distribution shapre | ci | Std. uncertainty ui [dB] |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | f ≦ 3GHz | 3GHz ≦ f < 4.2GHz |  |  |  | f ≦ 3GHz | 3GHz ≦ f < 4.2GHz |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |
| 1 | Misalignment DUT & pointing error | 0.10 | 0.10 | Rectangular | √3 | 1 | 0.06 | 0.06 |
| 2 | RF signal generator | 0.46 | 0.46 | Gaussian | 1 | 1 | 0.46 | 0.46 |
| 3 | Longitudinal position uncertainty (i.e. standing wave and imperfect field synthesis) for DUT antenna | 0.05 | 0.14 | Rectangular | √3 | 1 | 0.03 | 0.08 |
| 4 | RF leakage (measurement antenna connector terminated) | 0.086 | 0.086 | Normal | 1 | 1 | 0.09 | 0.09 |
| 5 | QZ ripple with DUT | [0.4] | [0.4] | Rectangular | √3 | 1 | [0.23] | [0.23] |
| 6 | Miscellaneous Uncertainty | 0.00 | 0.00 | Normal | 1 | 1 | 0.00 | 0.00 |
| 20 | System non-linearity | [0.062] | [0.062] | Rectangular | √3 | 1 | [0.04] | [0.04] |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |  |
| 7 | Uncertainty of network analyzer | 0.13 | 0.20 | Normal | 1 | 1 | 0.13 | 0.20 |
| 8 | Mismatch (i.e.reference antenna, network analyser and reference cable) | 0.127 | 0.325 | U-shaped | √2 | 1 | 0.09 | 0.23 |
| 9 | Insertion loss variation | 0.18 | 0.18 | Rectangular | √3 | 1 | 0.10 | 0.10 |
| 4 | RF leakage (calibration antenna connector terminated) | 0.086 | 0.086 | Normal | 1 | 1 | 0.09 | 0.09 |
| 10 | Influence of the calibration antenna feed cable | 0.103 | 0.104 | Rectangular | √3 | 1 | 0.06 | 0.06 |
| 11 | Uncertainty of the absolute gain of the calibration antenna | 0.50 | 0.43 | Rectangular | √3 | 1 | 0.29 | 0.25 |
| 12 | Misalignment of positioning system | 0.00 | 0.00 | Exp. normal | 2 | 1 | 0.00 | 0.00 |
| 13 | Misalignment of calibration antenna & pointing error | 0.05 | 0.05 | Rectangular | √3 | 1 | 0.03 | 0.03 |
| 14 | Rotary joints | 0.00 | 0.00 | U-shaped | √2 | 1 | 0.00 | 0.00 |
| 15 | Longitudinal position uncertainty (i.e. standing wave and imperfect field synthesis) for calibration antenna | 0.12 | 0.12 | Rectangular | √3 | 1 | 0.07 | 0.07 |
| 16 | QZ ripple with calibration antenna | [0.2] | [0.2] | Rectangular | √3 | 1 | [0.12] | [0.12] |
| 17 | Switching uncertainty | 0.02 | 0.02 | Rectangular | √3 | 1 | 0.01 | 0.01 |
| 18 | Field repeatability | [0.003] | [0.003] | Normal | 1 | 1 | [0.003] | [0.003] |
| 19 | Frequency flatness | 0.13 | 0.13 | Rectangular | √3 | 1 | 0.08 | 0.08 |
| Combined standard uncertainty (1σ) [dB] |  |  |  |  |  |  | [0.66] | [0.70] |
| Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  |  | [1.29] | [1.37] |

< End of Changes >

< Unchanged Text Deleted >

< Beginning of Changes >

# E.6	Plane Wave Synthesizer

E6-1 Misalignment DUT/calibration antenna & pointing error

This contribution denotes uncertainty in DUT/calibration antenna alignment and DUT/calibration antenna pointing error.  In this measurement the DUT/calibration antenna is aligned to maximum, also allowing for a zero contribution for polarization mismatch uncertainty. By adjusting for maximums to align, this contribution can be a small contribution. The calibration antenna's phase centre and polarization purity changes slightly according to the frequency. Therefore, there should be some uncertainty reserved for this. To ensure that the point error is at a minimal, this contribution should be captured using the antenna pattern cut which is broadest (in the case of the DUT this would most likely be in the azimuth domain).

E6-2 Longitudinal position uncertainty (i.e. standing wave and imperfect field synthesis) for DUT/calibration antenna

This value covers the effect of standing wave between DUT or calibration antenna and the test range antenna, but also counts for the PWS imperfect field synthesis over distance. This value can be captured by moving the DUT or calibration antenna towards the test range antenna.

E6-3 RF leakage

This contribution denotes noise leaking in to connector and cable(s) between test range antenna and receiving equipment.  The contribution also includes the noise leakage between the connector and cable(s) between reference antenna and transmitting equipment.

E6-4 QZ ripple with DUT/calibration antenna

This is the quiet zone (QZ) ripple experienced by the DUT/reference antenna during the measurement phase. The purpose of this component is to capture the contributions that the reflections from the walls, roof and floor that will add to measurements. The sum of all these reflections from the walls, roof and floor will give the overall value for the QZ ripple. In other words, the uncertainty component from the wall will not be separated from the roof or the floor. The purpose of this uncertainty component is to capture the overall reflections from the chamber walls experienced by the DUT/reference antenna.

E6-5 Miscellaneous Uncertainty

The term 'miscellaneous uncertainty' is used to define all the unknown, unquantifiable, etc. uncertainties associated with EIRP measurements. This term should include truly random effects as well as systematic uncertainties, such as that arising from dissimilarity between the patterns of the reference antenna and the DUT.

E6-6 	Mismatch

This uncertainty is the residual uncertainty contribution coming from multiple reflections between the receiving antenna and the test receiver equipment. This value can be captured through measurement by measuring the S11 towards the receive antenna and also towards the test receiver. The mismatch between the antenna reflection and the receiver reflection can also be calculated. If the same cable is used for calibration Stage 1, this can be considered systematic and negligible.

E6-7 	Insertion loss variation

This uncertainty is the residual uncertainty contribution coming from introducing an antenna at the end of the cable.  If this cable does not change/move between the calibration Stage 1 and the measurement Stage 2, the uncertainty is assumed to be systematic and negligible during the measurement stage. Alternatively, the insertion loss can also be calculated by taking the measurement of the cable where port 2 is the end of the cable connected to the AAS BS or calibration antenna.

IL = -20log10|S21| dB

E6-8 	Influence of the calibration antenna feed cable

a)	Flexing cables, adapters, attenuators, extra pathloss cable & connector repeatability.

During the calibration phase this cable is used to feed the calibration antenna and any influence it may have upon the measurements is captured. This is assessed by repeated measurements while flexing the cables and rotary joints. The largest difference between the results is recorded as the uncertainty.

E6-9 Misalignment of positioning system

This contribution originates from uncertainty in sliding position and turn table angle accuracy. If the calibration antenna is aligned to maximum this contribution can be considered negligible and therefore set to zero.

E6-10 Rotary Joints

If applicable, this uncertainty term corresponds to the accuracy in changing from azimuth to vertical measurements.

E6-11 Switching Uncertainty

The purpose of the switching unit is to switch electromechanically different RF path to different measurement instruments of different measurement modes. The electromechanical switching clearly reduces the errors arising from manual switching work. Switching is also used to measure the path loss values of each polarization component.  Even though the electromechanical switching is preferable during path loss and antenna performance measurements, some minor uncertainties can occur when the switch states are programmed to change their polarity.

E6-12 Field repeatability

[Each execution of field calibration of the measurement antenna array to find the PWS settings provides a slightly different set of settings for the RF components for each antenna path. This results in variation of the synthesized plane wave in the QZ and variation of PWS antenna to reference antenna coupling. This variation is described by field repeatability term. ]

E6-13 Frequency flatness

This uncertainty contribution to account for the frequency interpolation error caused by a finite frequency resolution during the calibration stage.

E6-14 System non-linearity

[This uncertainty term is calculated as RSS of the following items, assuming a rectangular distribution:

- System non-linearity in time. This is assessed by repeated measurements over a period of time (e.g. 60 minutes) for the same reference power transmitted by the reference antenna. The largest difference between the results is recorded as the uncertainty.

- System non-linearity in power. This is assessed by repeated measurements over a range of transmitted powers. The largest delta between the increments are the transmitting and receiving sides is recorded as the uncertainty.]

< End of Changes >