3GPP TSG-RAN4 Meeting #92-Bis	R4-1911849

Chongqing, China, 14th – 18th Oct, 2019

Agenda Item:		5.1.3.5

Source: 		Rohde & Schwarz

Title: 		Plane Wave Synthesizer – Pending MU terms

Document for:		Approval

# Introduction

Following the outcome from previous meeting where Plane Wave Synthesizer (PWS) was introduced as accepted method in TR 37.843 [4], this contribution presents an analysis of the MU terms that required further discussion.

# MU budgets

The following are the MU budget tables applicable to PWS as agreed in [2], both EIRP and EIS measurements, where the MU terms that required further analysis are highlighted:

##### 10.2.2.3.4	MU assessment

###### 10.2.2.3.4.1	MU Budget

Table 10.2.2.3.4.1-1: Plane Wave Synthesizer uncertainty contributions for EIRP measurements

| UID | Uncertainty Source | Details in annex |
| --- | --- | --- |
| Stage 2: DUT measurement |  |  |
| 1 | Misalignment DUT & pointing error | E6-1 |
| 2 | RF power measurement equipment | F.1 |
| 3 | Longitudinal position uncertainty (i.e. standing wave and imperfect field synthesis) for DUT antenna | E6-2 |
| 4 | RF leakage (calibration antenna connector terminated) | E6-3 |
| 5 | QZ ripple with DUT | E6-4 |
| 6 | Miscellaneous Uncertainty | E6-5 |
| 20 | System non-linearity | E6-14 |
| Stage 1: Calibration measurement |  |  |
| 7 | Uncertainty of network analyzer | F.1 |
| 8 | Mismatch (i.e. reference antenna, network analyser and reference cable) | E6-6 |
| 9 | Insertion loss variation | E6-7 |
| 4 | RF leakage (calibration antenna connector terminated) | E6-3 |
| 10 | Influence of the calibration antenna feed cable | E6-8 |
| 11 | Uncertainty of the absolute gain of the calibration antenna | F.1 |
| 12 | Misalignment of positioning system | E6-9 |
| 13 | Misalignment of calibration antenna & pointing error | E6-1 |
| 14 | Rotary joints | E6-10 |
| 15 | Longitudinal position uncertainty (i.e. standing wave and imperfect field synthesis) for calibration antenna | E6-2 |
| 16 | QZ ripple with calibration antenna | E6-4 |
| 17 | Switching uncertainty | E6-11 |
| 18 | Field repeatability | E6-12 |
| 19 | Frequency flatness | E6-13 |

##### 10.3.2.2.3	MU assessment

###### 10.3.2.2.3.1	MU Budget

Table 10.3.2.2.3.1-1: Plane Wave Synthesizer uncertainty contributions for EIS measurements

| UID | Uncertainty Source | Details in annex |
| --- | --- | --- |
| Stage 2: DUT measurement |  |  |
| 1 | Misalignment DUT & pointing error | E6-1 |
| 2 | RF signal generator | F.1 |
| 3 | Longitudinal position uncertainty (i.e. standing wave and imperfect field synthesis) for DUT antenna | E6-2 |
| 4 | RF leakage (calibration antenna connector terminated) | E6-3 |
| 5 | QZ ripple with DUT | E6-4 |
| 6 | Miscellaneous Uncertainty | E6-5 |
| 20 | System non-linearity | E6-14 |
| Stage 1: Calibration measurement |  |  |
| 7 | Uncertainty of network analyzer | F.1 |
| 8 | Mismatch (i.e. reference antenna, network analyzer and reference cable) | E6-6 |
| 9 | Insertion loss variation | E6-7 |
| 4 | RF leakage (calibration antenna connector terminated) | E6-3 |
| 10 | Influence of the calibration antenna feed cable | E6-8 |
| 11 | Uncertainty of the absolute gain of the calibration antenna | F.1 |
| 12 | Misalignment of positioning system | E6-9 |
| 13 | Misalignment of calibration antenna & pointing error | E6-1 |
| 14 | Rotary joints | E6-10 |
| 15 | Longitudinal position uncertainty (i.e. standing wave and imperfect field synthesis) for calibration antenna | E6-2 |
| 16 | QZ ripple with calibration antenna | E6-4 |
| 17 | Switching uncertainty | E6-11 |
| 18 | Field repeatability | E6-12 |
| 19 | Frequency flatness | E6-13 |

# Error contribution descriptions

The following sections describe in detail the corresponding evaluation, based on measurements on the actual R&S implementation of PWS, for each of the MU terms highlighted in previous tables, although QZ ripple and system non-linearity values are not presented in this contribution since the testing campaign is still being performed.

## Field repeatability

Each execution of field calibration of the measurement antenna array to find the PWS settings provides a slightly different set of settings for the RF components for each antenna path. This results in variation of the synthesized plane wave in the QZ and variation of PWS antenna to reference antenna coupling. This variation u is described by field calibration repeatability term.

In other to determine this term, the following step by step process is followed:

- Perform field calibration.

- Enable PWS setting.

- Measure S21 between the PWS and calibration reference antenna at the PWS setting frequency.

- Repeat field calibration and S21 measurement for all PWC setting frequencies at least I=5 times. A Standard Deviation is calculated for the results per frequency.

Error is determined as the maximum among all test frequencies for the frequency range being considered.

| UID | Uncertainty Source | Value f <= 3 GHz | Value f > 3 GHz | Unit | Distr. | Divisor | Sens. | ui (dB) f <= 3 GHz | ui (dB) f > 3 GHz |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 18 | Field calibration repeatability | 0.06 | 0.12 | dB | Normal | 1 | 1 | 0.06 | 0.12 |

## System non-linearity

This uncertainty term is calculated as RSS of the following items, assuming a rectangular distribution:

- System non-linearity in time. This is assessed by repeated measurements over a period of time (e.g. 60 minutes) for the same reference power transmitted by the reference antenna. The largest difference between the results is recorded as the uncertainty.

- System non-linearity in power. This is assessed by repeated measurements over a range of transmitted powers. The largest delta between the increments on the receiving side versus the transmitting side is recorded as the uncertainty.

This contribution is assumed to have a rectangular distribution.

# MU values

As an outcome of previous sections, the following tables 4-1 and 4-2 show the final MU value tables applicable to the PWS methodology after included the declared values above for the Field repeatability term.

| UID | Uncertainty Source | Uncertainty Value |  | Distribution of the probability | Divisor based on distribution shapre | ci | Std. uncertainty ui [dB] |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | f ≦ 3GHz | 3GHz ≦ f < 4.2GHz |  |  |  | f ≦ 3GHz | 3GHz ≦ f < 4.2GHz |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |
| 1 | Misalignment DUT & pointing error | 0.10 | 0.10 | Rectangular | √3 | 1 | 0.06 | 0.06 |
| 2 | RF power measurement equipment | 0.14 | 0.26 | Gaussian | 1 | 1 | 0.14 | 0.26 |
| 3 | Longitudinal position uncertainty (i.e. standing wave and imperfect field synthesis) for DUT antenna | 0.05 | 0.14 | Rectangular | √3 | 1 | 0.03 | 0.08 |
| 4 | RF leakage (calibration antenna connector terminated) | 0.086 | 0.086 | Normal | 1 | 1 | 0.09 | 0.09 |
| 5 | QZ ripple with DUT | [0.4] | [0.4] | Rectangular | √3 | 1 | [0.23] | [0.23] |
| 6 | Miscellaneous Uncertainty | 0.00 | 0.00 | Normal | 1 | 1 | 0.00 | 0.00 |
| 20 | System non-linearity | [0.062] | [0.062] | Rectangular | √3 | 1 | [0.04] | [0.04] |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |  |
| 7 | Uncertainty of network analyzer | 0.13 | 0.20 | Normal | 1 | 1 | 0.13 | 0.20 |
| 8 | Mismatch (i.e. reference antenna, network analyser and reference cable) | 0.127 | 0.325 | U-shaped | √2 | 1 | 0.09 | 0.23 |
| 9 | Insertion loss variation | 0.18 | 0.18 | Rectangular | √3 | 1 | 0.10 | 0.10 |
| 4 | RF leakage (calibration antenna connector terminated) | 0.086 | 0.086 | Normal | 1 | 1 | 0.09 | 0.09 |
| 10 | Influence of the calibration antenna feed cable | 0.103 | 0.104 | Rectangular | √3 | 1 | 0.06 | 0.06 |
| 11 | Uncertainty of the absolute gain of the calibration antenna | 0.50 | 0.43 | Rectangular | √3 | 1 | 0.29 | 0.25 |
| 12 | Misalignment of positioning system | 0.00 | 0.00 | Exp. normal | 2 | 1 | 0.00 | 0.00 |
| 13 | Misalignment of calibration antenna & pointing error | 0.05 | 0.05 | Rectangular | √3 | 1 | 0.03 | 0.03 |
| 14 | Rotary joints | 0.00 | 0.00 | U-shaped | √2 | 1 | 0.00 | 0.00 |
| 15 | Longitudinal position uncertainty (i.e. standing wave and imperfect field synthesis) for calibration antenna | 0.12 | 0.12 | Rectangular | √3 | 1 | 0.07 | 0.07 |
| 16 | QZ ripple with calibration antenna | [0.2] | [0.2] | Rectangular | √3 | 1 | [0.12] | [0.12] |
| 17 | Switching uncertainty | 0.02 | 0.02 | Rectangular | √3 | 1 | 0.01 | 0.01 |
| 18 | Field repeatability | 0.06 | 0.12 | Normal | 1 | 1 | 0.06 | 0.12 |
| 19 | Frequency flatness | 0.13 | 0.13 | Rectangular | √3 | 1 | 0.08 | 0.08 |
| Combined standard uncertainty (1σ) [dB] |  |  |  |  |  |  | [0.49] | [0.60] |
| Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  |  | [0.97] | [1.17] |

Table 4-1: PWS uncertainty values for EIRP measurements

| UID | Uncertainty Source | Uncertainty Value |  | Distribution of the probability | Divisor based on distribution shapre | ci | Std. uncertainty ui [dB] |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | f ≦ 3GHz | 3GHz ≦ f < 4.2GHz |  |  |  | f ≦ 3GHz | 3GHz ≦ f < 4.2GHz |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |
| 1 | Misalignment DUT & pointing error | 0.10 | 0.10 | Rectangular | √3 | 1 | 0.06 | 0.06 |
| 2 | RF signal generator | 0.46 | 0.46 | Gaussian | 1 | 1 | 0.46 | 0.46 |
| 3 | Longitudinal position uncertainty (i.e. standing wave and imperfect field synthesis) for DUT antenna | 0.05 | 0.14 | Rectangular | √3 | 1 | 0.03 | 0.08 |
| 4 | RF leakage (measurement antenna connector terminated) | 0.086 | 0.086 | Normal | 1 | 1 | 0.09 | 0.09 |
| 5 | QZ ripple with DUT | [0.4] | [0.4] | Rectangular | √3 | 1 | [0.23] | [0.23] |
| 6 | Miscellaneous Uncertainty | 0.00 | 0.00 | Normal | 1 | 1 | 0.00 | 0.00 |
| 20 | System non-linearity | [0.062] | [0.062] | Rectangular | √3 | 1 | [0.04] | [0.04] |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |  |
| 7 | Uncertainty of network analyzer | 0.13 | 0.20 | Normal | 1 | 1 | 0.13 | 0.20 |
| 8 | Mismatch (i.e.reference antenna, network analyser and reference cable) | 0.127 | 0.325 | U-shaped | √2 | 1 | 0.09 | 0.23 |
| 9 | Insertion loss variation | 0.18 | 0.18 | Rectangular | √3 | 1 | 0.10 | 0.10 |
| 4 | RF leakage (calibration antenna connector terminated) | 0.086 | 0.086 | Normal | 1 | 1 | 0.09 | 0.09 |
| 10 | Influence of the calibration antenna feed cable | 0.103 | 0.104 | Rectangular | √3 | 1 | 0.06 | 0.06 |
| 11 | Uncertainty of the absolute gain of the calibration antenna | 0.50 | 0.43 | Rectangular | √3 | 1 | 0.29 | 0.25 |
| 12 | Misalignment of positioning system | 0.00 | 0.00 | Exp. normal | 2 | 1 | 0.00 | 0.00 |
| 13 | Misalignment of calibration antenna & pointing error | 0.05 | 0.05 | Rectangular | √3 | 1 | 0.03 | 0.03 |
| 14 | Rotary joints | 0.00 | 0.00 | U-shaped | √2 | 1 | 0.00 | 0.00 |
| 15 | Longitudinal position uncertainty (i.e. standing wave and imperfect field synthesis) for calibration antenna | 0.12 | 0.12 | Rectangular | √3 | 1 | 0.07 | 0.07 |
| 16 | QZ ripple with calibration antenna | [0.2] | [0.2] | Rectangular | √3 | 1 | [0.12] | [0.12] |
| 17 | Switching uncertainty | 0.02 | 0.02 | Rectangular | √3 | 1 | 0.01 | 0.01 |
| 18 | Field repeatability | 0.06 | 0.12 | Normal | 1 | 1 | 0.06 | 0.12 |
| 19 | Frequency flatness | 0.13 | 0.13 | Rectangular | √3 | 1 | 0.08 | 0.08 |
| Combined standard uncertainty (1σ) [dB] |  |  |  |  |  |  | [0.66] | [0.71] |
| Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  |  | [1.29] | [1.38] |

Table 4-2: PWS uncertainty values for EIS measurements

# Conclusion

In this contribution, a description of the pending MU terms for the Plane Wave Synthesizer has been provided. It is proposed to incorporate the Field repeatability MU final values as detailed in  Tables 4-1 and 4-2 into TR 37.843.

Proposal: incorporate Field repeatability MU values from Tables 4-1 and 4-2 into the corresponding tables for PWS in TR 37.843.

# References

- R4-1806605, “2D Compact Range for Testing of AAS Basestations”, Rohde & Schwarz, RAN4#87, May 2018

- R4-1910434, “CR to TR 37.843: Addition of 2D Compact Range. MU, calibration and test procedure for EIRP and EIS”,  Rohde & Schwarz, CAICT, MVG, RAN4#92, August 2019

- 3GPP TR 37.843: “Radio Frequency (RF) requirement background for Active Antenna System (AAS) Base Station (BS) radiated requirements”.

- R.  Haupt,  “Generating  a  Plane  Wave  in  the  Near  Field  with  a  Planar Array Antenna”, Microwave Journal, September 2003

- C. Rowell, A. Tankielun, “Plane Wave Converter for 5G Massive MIMO Basestation Measurements”, the 12th European Conference on Antennas and Propagation (EuCAP2018), London, UK, 9 - 13 April, 2018

- B. Derat, “Creating Far-field Condition at a Tenth of Far-field Distance: An Innovative Technique for 5G OTA measurement”, Electronic Design Innovation Conference (EDICON), Beijing, China, 20 – 22 March, 2018