3GPP TSG-RAN4 Meeting #92	R4-1912006

# Ljubljana, Slovenia, 26th Aug 2019 - 30th Aug 2019

| CR-Form-v12.0 |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CHANGE REQUEST |  |  |  |  |  |  |  |  |
|  | 37.843 | CR | 0032 | rev | - | Current version: | 15.5.0 |  |
| For HELP on using this form: comprehensive instructions can be found at http://www.3gpp.org/Change-Requests. |  |  |  |  |  |  |  |  |

| Proposed change affects: | UICC apps |  | ME |  | Radio Access Network |  | Core Network |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |

| Title: | CR to TR 37.843: Completion of MU terms for PWS. |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Source to WG: | Rohde & Schwarz |  |  |  |  |  |  |  |  |  |
| Source to TSG: |  |  |  |  |  |  |  |  |  |  |
| Work item code: | AASenh_BS_LTE_UTRA-Perf |  |  |  |  |  | Date: |  |  | 2019-10-04 |
| Category: | F |  |  |  |  |  | Release: |  |  | Rel-15 |
|  | Use one of the following categories: F (correction) A (mirror corresponding to a change in an earlier release) B (addition of feature), C (functional modification of feature) D (editorial modification) Detailed explanations of the above categories can be found in 3GPP TR 21.900. |  |  |  |  |  |  |  | Use one of the following releases: Rel-8 (Release 8) Rel-9 (Release 9) Rel-10 (Release 10) Rel-11 (Release 11) Rel-12 (Release 12) Rel-13 (Release 13) Rel-14 (Release 14) Rel-15 (Release 15) Rel-16 (Release 16) |  |
| Reason for change: |  | There are still MU values and MU descriptions between brackets that requires completion. |  |  |  |  |  |  |  |  |
| Summary of change: |  | - New values for Field Repeatibility MU term have been included on tables 10.2.2.3.3.2-1 and 10.3.2.2.2.2-1, and corresponding square brackets have been removed. - MU description for System non-linearity in clause E.6 have been reworded. - Square brackets have been removed from MU descriptions for Field repeatibility and System non-linearity terms in clause E.6. |  |  |  |  |  |  |  |  |
| Consequences if not approved: |  | Plane Wave Synthesizer could not be used as methodology for AAS BS conformance testing. |  |  |  |  |  |  |  |  |
| Clauses affected: |  | 10.2.2.3.3.2, 10.3.2.2.2.2 and E.6 |  |  |  |  |  |  |  |  |
|  |  | Y | N |  |  |  |  |  |  |  |
| Other specs |  |  | X | Other core specifications |  |  |  | TS/TR ... CR ... |  |  |
| affected: |  |  | X | Test specifications |  |  |  | TS/TR ... CR ... |  |  |
| (show related CRs) |  |  | X | O&M Specifications |  |  |  | TS/TR ... CR ... |  |  |
| Other comments: |  | - |  |  |  |  |  |  |  |  |

| This CR's revision history: | - |
| --- | --- |

< Unchanged Text Deleted >

< Beginning of Changes >

###### 10.2.2.3.3.2	MU value

Table 10.2.2.3.3.2-1: Plane Wave Synthesizer uncertainty assessment for EIRP measurements

| UID | Uncertainty Source | Uncertainty Value |  | Distribution of the probability | Divisor based on distribution shapre | ci | Std. uncertainty ui [dB] |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | f ≦ 3GHz | 3GHz ≦ f < 4.2GHz |  |  |  | f ≦ 3GHz | 3GHz ≦ f < 4.2GHz |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |
| 1 | Misalignment DUT & pointing error | 0.10 | 0.10 | Rectangular | √3 | 1 | 0.06 | 0.06 |
| 2 | RF power measurement equipment | 0.14 | 0.26 | Gaussian | 1 | 1 | 0.14 | 0.26 |
| 3 | Longitudinal position uncertainty (i.e. standing wave and imperfect field synthesis) for DUT antenna | 0.05 | 0.14 | Rectangular | √3 | 1 | 0.03 | 0.08 |
| 4 | RF leakage (calibration antenna connector terminated) | 0.086 | 0.086 | Normal | 1 | 1 | 0.09 | 0.09 |
| 5 | QZ ripple with DUT | [0.4] | [0.4] | Rectangular | √3 | 1 | [0.23] | [0.23] |
| 6 | Miscellaneous Uncertainty | 0.00 | 0.00 | Normal | 1 | 1 | 0.00 | 0.00 |
| 20 | System non-linearity | [0.062] | [0.062] | Rectangular | √3 | 1 | [0.04] | [0.04] |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |  |
| 7 | Uncertainty of network analyzer | 0.13 | 0.20 | Normal | 1 | 1 | 0.13 | 0.20 |
| 8 | Mismatch (i.e. reference antenna, network analyser and reference cable) | 0.127 | 0.325 | U-shaped | √2 | 1 | 0.09 | 0.23 |
| 9 | Insertion loss variation | 0.18 | 0.18 | Rectangular | √3 | 1 | 0.10 | 0.10 |
| 4 | RF leakage (calibration antenna connector terminated) | 0.086 | 0.086 | Normal | 1 | 1 | 0.09 | 0.09 |
| 10 | Influence of the calibration antenna feed cable | 0.103 | 0.104 | Rectangular | √3 | 1 | 0.06 | 0.06 |
| 11 | Uncertainty of the absolute gain of the calibration antenna | 0.50 | 0.43 | Rectangular | √3 | 1 | 0.29 | 0.25 |
| 12 | Misalignment of positioning system | 0.00 | 0.00 | Exp. normal | 2 | 1 | 0.00 | 0.00 |
| 13 | Misalignment of calibration antenna & pointing error | 0.05 | 0.05 | Rectangular | √3 | 1 | 0.03 | 0.03 |
| 14 | Rotary joints | 0.00 | 0.00 | U-shaped | √2 | 1 | 0.00 | 0.00 |
| 15 | Longitudinal position uncertainty (i.e. standing wave and imperfect field synthesis) for calibration antenna | 0.12 | 0.12 | Rectangular | √3 | 1 | 0.07 | 0.07 |
| 16 | QZ ripple with calibration antenna | [0.2] | [0.2] | Rectangular | √3 | 1 | [0.12] | [0.12] |
| 17 | Switching uncertainty | 0.02 | 0.02 | Rectangular | √3 | 1 | 0.01 | 0.01 |
| 18 | Field repeatability | 0.06 | 0.12 | Normal | 1 | 1 | 0.06 | 0.12 |
| 19 | Frequency flatness | 0.13 | 0.13 | Rectangular | √3 | 1 | 0.08 | 0.08 |
| Combined standard uncertainty (1σ) [dB] |  |  |  |  |  |  | [0.49] | [0.60] |
| Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  |  | [0.97] | [1.17] |

< End of Changes >

< Unchanged Text Deleted >

< Beginning of Changes >

###### 10.3.2.2.2.2	MU value

Table 10.3.2.2.2.2-1: Plane Wave Synthesizer uncertainty assessment for EIS measurements

| UID | Uncertainty Source | Uncertainty Value |  | Distribution of the probability | Divisor based on distribution shapre | ci | Std. uncertainty ui [dB] |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | f ≦ 3GHz | 3GHz ≦ f < 4.2GHz |  |  |  | f ≦ 3GHz | 3GHz ≦ f < 4.2GHz |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |
| 1 | Misalignment DUT & pointing error | 0.10 | 0.10 | Rectangular | √3 | 1 | 0.06 | 0.06 |
| 2 | RF signal generator | 0.46 | 0.46 | Gaussian | 1 | 1 | 0.46 | 0.46 |
| 3 | Longitudinal position uncertainty (i.e. standing wave and imperfect field synthesis) for DUT antenna | 0.05 | 0.14 | Rectangular | √3 | 1 | 0.03 | 0.08 |
| 4 | RF leakage (measurement antenna connector terminated) | 0.086 | 0.086 | Normal | 1 | 1 | 0.09 | 0.09 |
| 5 | QZ ripple with DUT | [0.4] | [0.4] | Rectangular | √3 | 1 | [0.23] | [0.23] |
| 6 | Miscellaneous Uncertainty | 0.00 | 0.00 | Normal | 1 | 1 | 0.00 | 0.00 |
| 20 | System non-linearity | [0.062] | [0.062] | Rectangular | √3 | 1 | [0.04] | [0.04] |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |  |
| 7 | Uncertainty of network analyzer | 0.13 | 0.20 | Normal | 1 | 1 | 0.13 | 0.20 |
| 8 | Mismatch (i.e.reference antenna, network analyser and reference cable) | 0.127 | 0.325 | U-shaped | √2 | 1 | 0.09 | 0.23 |
| 9 | Insertion loss variation | 0.18 | 0.18 | Rectangular | √3 | 1 | 0.10 | 0.10 |
| 4 | RF leakage (calibration antenna connector terminated) | 0.086 | 0.086 | Normal | 1 | 1 | 0.09 | 0.09 |
| 10 | Influence of the calibration antenna feed cable | 0.103 | 0.104 | Rectangular | √3 | 1 | 0.06 | 0.06 |
| 11 | Uncertainty of the absolute gain of the calibration antenna | 0.50 | 0.43 | Rectangular | √3 | 1 | 0.29 | 0.25 |
| 12 | Misalignment of positioning system | 0.00 | 0.00 | Exp. normal | 2 | 1 | 0.00 | 0.00 |
| 13 | Misalignment of calibration antenna & pointing error | 0.05 | 0.05 | Rectangular | √3 | 1 | 0.03 | 0.03 |
| 14 | Rotary joints | 0.00 | 0.00 | U-shaped | √2 | 1 | 0.00 | 0.00 |
| 15 | Longitudinal position uncertainty (i.e. standing wave and imperfect field synthesis) for calibration antenna | 0.12 | 0.12 | Rectangular | √3 | 1 | 0.07 | 0.07 |
| 16 | QZ ripple with calibration antenna | [0.2] | [0.2] | Rectangular | √3 | 1 | [0.12] | [0.12] |
| 17 | Switching uncertainty | 0.02 | 0.02 | Rectangular | √3 | 1 | 0.01 | 0.01 |
| 18 | Field repeatability | 0.06 | 0.12 | Normal | 1 | 1 | 0.06 | 0.12 |
| 19 | Frequency flatness | 0.13 | 0.13 | Rectangular | √3 | 1 | 0.08 | 0.08 |
| Combined standard uncertainty (1σ) [dB] |  |  |  |  |  |  | [0.66] | [0.71] |
| Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  |  | [1.29] | [1.38] |

< End of Changes >

< Unchanged Text Deleted >

< Beginning of Changes >

# E.6	Plane Wave Synthesizer

E6-1 Misalignment DUT/calibration antenna & pointing error

This contribution denotes uncertainty in DUT/calibration antenna alignment and DUT/calibration antenna pointing error.  In this measurement the DUT/calibration antenna is aligned to maximum, also allowing for a zero contribution for polarization mismatch uncertainty. By adjusting for maximums to align, this contribution can be a small contribution. The calibration antenna's phase centre and polarization purity changes slightly according to the frequency. Therefore, there should be some uncertainty reserved for this. To ensure that the point error is at a minimal, this contribution should be captured using the antenna pattern cut which is broadest (in the case of the DUT this would most likely be in the azimuth domain).

E6-2 Longitudinal position uncertainty (i.e. standing wave and imperfect field synthesis) for DUT/calibration antenna

This value covers the effect of standing wave between DUT or calibration antenna and the test range antenna, but also counts for the PWS imperfect field synthesis over distance. This value can be captured by moving the DUT or calibration antenna towards the test range antenna.

E6-3 RF leakage

This contribution denotes noise leaking in to connector and cable(s) between test range antenna and receiving equipment.  The contribution also includes the noise leakage between the connector and cable(s) between reference antenna and transmitting equipment.

E6-4 QZ ripple with DUT/calibration antenna

This is the quiet zone (QZ) ripple experienced by the DUT/reference antenna during the measurement phase. The purpose of this component is to capture the contributions that the reflections from the walls, roof and floor that will add to measurements. The sum of all these reflections from the walls, roof and floor will give the overall value for the QZ ripple. In other words, the uncertainty component from the wall will not be separated from the roof or the floor. The purpose of this uncertainty component is to capture the overall reflections from the chamber walls experienced by the DUT/reference antenna.

E6-5 Miscellaneous Uncertainty

The term 'miscellaneous uncertainty' is used to define all the unknown, unquantifiable, etc. uncertainties associated with EIRP measurements. This term should include truly random effects as well as systematic uncertainties, such as that arising from dissimilarity between the patterns of the reference antenna and the DUT.

E6-6 	Mismatch

This uncertainty is the residual uncertainty contribution coming from multiple reflections between the receiving antenna and the test receiver equipment. This value can be captured through measurement by measuring the S11 towards the receive antenna and also towards the test receiver. The mismatch between the antenna reflection and the receiver reflection can also be calculated. If the same cable is used for calibration Stage 1, this can be considered systematic and negligible.

E6-7 	Insertion loss variation

This uncertainty is the residual uncertainty contribution coming from introducing an antenna at the end of the cable.  If this cable does not change/move between the calibration Stage 1 and the measurement Stage 2, the uncertainty is assumed to be systematic and negligible during the measurement stage. Alternatively, the insertion loss can also be calculated by taking the measurement of the cable where port 2 is the end of the cable connected to the AAS BS or calibration antenna.

IL = -20log10|S21| dB

E6-8 	Influence of the calibration antenna feed cable

a)	Flexing cables, adapters, attenuators, extra pathloss cable & connector repeatability.

During the calibration phase this cable is used to feed the calibration antenna and any influence it may have upon the measurements is captured. This is assessed by repeated measurements while flexing the cables and rotary joints. The largest difference between the results is recorded as the uncertainty.

E6-9 Misalignment of positioning system

This contribution originates from uncertainty in sliding position and turn table angle accuracy. If the calibration antenna is aligned to maximum this contribution can be considered negligible and therefore set to zero.

E6-10 Rotary Joints

If applicable, this uncertainty term corresponds to the accuracy in changing from azimuth to vertical measurements.

E6-11 Switching Uncertainty

The purpose of the switching unit is to switch electromechanically different RF path to different measurement instruments of different measurement modes. The electromechanical switching clearly reduces the errors arising from manual switching work. Switching is also used to measure the path loss values of each polarization component.  Even though the electromechanical switching is preferable during path loss and antenna performance measurements, some minor uncertainties can occur when the switch states are programmed to change their polarity.

E6-12 Field repeatability

Each execution of field calibration of the measurement antenna array to find the PWS settings provides a slightly different set of settings for the RF components for each antenna path. This results in variation of the synthesized plane wave in the QZ and variation of PWS antenna to reference antenna coupling. This variation is described by field repeatability term.

E6-13 Frequency flatness

This uncertainty contribution to account for the frequency interpolation error caused by a finite frequency resolution during the calibration stage.

E6-14 System non-linearity

This uncertainty term is calculated as RSS of the following items, assuming a rectangular distribution:

- System non-linearity in time. This is assessed by repeated measurements over a period of time (e.g. 60 minutes) for the same reference power transmitted by the reference antenna. The largest difference between the results is recorded as the uncertainty.

- System non-linearity in power. This is assessed by repeated measurements over a range of transmitted powers. The largest delta between the increments on the receiving side versus the transmitting side is recorded as the uncertainty.

< End of Changes >