3GPP TSG-RAN4 Meeting #92-Bis	R4- 1912936

# Chongqing, China, 14th Oct 2019 - 18th Oct 2019

| CR-Form-v12.0 |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CHANGE REQUEST |  |  |  |  |  |  |  |  |
|  | 37.843 | CR | 2540 | rev | 1 | Current version: | 15.5.0 |  |
| For HELP on using this form: comprehensive instructions can be found at http://www.3gpp.org/Change-Requests. |  |  |  |  |  |  |  |  |

| Proposed change affects: | UICC apps |  | ME |  | Radio Access Network | X | Core Network |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |

| Title: | Revised CR to TR 37.843: Addition of Plane Wave Synthesizer. MU, calibration and test procedure for In-band TRP requirements. |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Source to WG: | CAICT SRTC |  |  |  |  |  |  |  |  |  |
| Source to TSG: | R4 |  |  |  |  |  |  |  |  |  |
| Work item code: | AASenh_BS_LTE_UTRA-Perf |  |  |  |  |  | Date: |  |  | 2019-10-17 |
| Category: | F |  |  |  |  |  | Release: |  |  | Rel-15 |
|  | Use one of the following categories: F (correction) A (mirror corresponding to a change in an earlier release) B (addition of feature), C (functional modification of feature) D (editorial modification) Detailed explanations of the above categories can be found in 3GPP TR 21.900. |  |  |  |  |  |  |  | Use one of the following releases: Rel-8 (Release 8) Rel-9 (Release 9) Rel-10 (Release 10) Rel-11 (Release 11) Rel-12 (Release 12) Rel-13 (Release 13) Rel-14 (Release 14) Rel-15 (Release 15) Rel-16 (Release 16) |  |
| Reason for change: |  | -Contribute to complete the PWS for OTA. -Some error need to be change,and add some discri |  |  |  |  |  |  |  |  |
| Summary of change: |  | Add the PWS methord in clauses 10.4. |  |  |  |  |  |  |  |  |
| Consequences if not approved: |  | The PWS methord will not be complete. |  |  |  |  |  |  |  |  |
| Clauses affected: |  | Clauses 10.4.1, 10.4.2 ,10.4.3 and 10.4.4. |  |  |  |  |  |  |  |  |
|  |  | Y | N |  |  |  |  |  |  |  |
| Other specs |  |  | X | Other core specifications |  |  |  | TS/TR ... CR ... |  |  |
| affected: |  |  | X | Test specifications |  |  |  | TS/TR ... CR ... |  |  |
| (show related CRs) |  |  | X | O&M Specifications |  |  |  | TS/TR ... CR ... |  |  |
| Other comments: |  | - |  |  |  |  |  |  |  |  |

| This CR's revision history: | - Modify the MU value of every sections - Add the limitation description of PWS. - Change the subcluse number of every sections. |
| --- | --- |

< Unchanged Text Deleted >

< Beginning of Changes >

#### 10.4.1.4B	Plane Wave Synthesizer

10.4.1.4B.1	General

The PWS is described in section 10.2.2.3.

PWS has limitation on max EIRP base station power. Different PWS systems might have different limits on max power or other items.

10.4.1.4B.2	Calibration

Calibration should be carried out using the same procedure as in 10.2.2.3

##### 10.4.1.4B.3	Procedure

1)	Install the AAS BS as DUT with the manufacturer declared coordinate system reference point in the same place as the phase center of the reference antenna A as shown in figure 10.4.1.4B.3-1. The manufacturer declared coordinate system orientation of the AAS BS is set to be aligned with the testing system.

Figure 10.4.1.4B.3-1: PWS setup for TRP measurements

2)	Set the AAS BS to generate the tested beam with the appropriate test model with the beam peak direction intended to be the same as the testing direction.

3)	Measure mean power (Pmeas) at the measurement equipment (such as a spectrum analyzer or power meter) denoted in figure 10.4.1.4B.3-1.

4)	Measure the appropriate test parameter PR_AAS_desired,C.

5)	Repeat the above steps 3)~5) at a number of discrete directions around the sphere according to the chosen measurement grid, see subclause 10.8.3.

6)	Calculate TRPs from power density, as shown in subclause 10.8.3.

Figure 10.4.1.5.3-1: PWS setup for TRP measurements

The appropriate test parameter in step 5 is PR_AAS_desired, C, measured mean power within the desired signal channel bandwidth for each carrier arriving at the measurement equipment connector C in figure 10.4.1.4B.3-1, and  calculation of power powerc using the following equitation:

PowerC = PR_AAS_desired, C+ L

NOTE: If the test facility only supports single polarization, then measure power with the test facility's test antenna/probe polarization matched to the AAS BS, then measure and sum the power on both polarizations. If the test facility supports dual polarization then measure total power for two orthogonal polarizations and calculate total radiated transmit power as the sum over both polarizations.

##### 10.4.1.4B.4	MU assessment

10.4.1.4B.4.1	MU Budget

Table 10.4.1.4B.4.1-1: Plane Wave Synthesizer uncertainty contributions for OTA Base station output power measurement

| UID | Uncertainty Source | Details in annex |
| --- | --- | --- |
| Stage 2: DUT measurement |  |  |
| 1 | Misalignment DUT & pointing error | E6-1 |
| 2 | RF power measurement equipment | F.1 |
| 3 | Longitudinal position uncertainty (i.e. standing wave and imperfect field synthesis) for DUT antenna | E6-2 |
| 4 | RF leakage (calibration antenna connector terminated) | E6-3 |
| 5 | QZ ripple with DUT | E6-4 |
| 6 | Miscellaneous Uncertainty | E6-5 |
| 20 | System non-linearity | E6-14 |
| Stage 1: Calibration measurement |  |  |
| 7 | Uncertainty of network analyzer | F.1 |
| 8 | Mismatch (i.e. reference antenna, network analyser and reference cable) | E6-6 |
| 9 | Insertion loss variation | E6-7 |
| 4 | RF leakage (calibration antenna connector terminated) | E6-3 |
| 10 | Influence of the calibration antenna feed cable | E6-8 |
| 11 | Uncertainty of the absolute gain of the calibration antenna | F.1 |
| 12 | Misalignment of positioning system | E6-9 |
| 13 | Misalignment of calibration antenna & pointing error | E6-1 |
| 14 | Rotary joints | E6-10 |
| 15 | Longitudinal position uncertainty (i.e. standing wave and imperfect field synthesis) for calibration antenna | E6-2 |
| 16 | QZ ripple with calibration antenna | E6-4 |
| 17 | Switching uncertainty | E6-11 |
| 18 | Field repeatability | E6-12 |
| 19 | Frequency flatness | E6-13 |

10.4.1.4B.4.2	MU Value

Table 10.4.1.4B.4.2-1: Plane Wave Synthesizer uncertainty assessment for OTA base station output power

| UID | Uncertainty Source | Uncertainty Value |  | Distribution of the probability | Divisor based on distribution shapre | ci | Std. uncertainty ui [dB] |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | f ≦ 3GHz | 3GHz ≦ f < 4.2GHz |  |  |  | f ≦ 3GHz | 3GHz ≦ f < 4.2GHz |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |
| 1 | Misalignment DUT & pointing error | 0.10 | 0.10 | Rectangular | √3 | 1 | 0.06 | 0.06 |
| 2 | RF power measurement equipment | 0.14 | 0.26 | Gaussian | 1 | 1 | 0.14 | 0.26 |
| 3 | Longitudinal position uncertainty (i.e. standing wave and imperfect field synthesis) for DUT antenna | 0.05 | 0.14 | Rectangular | √3 | 1 | 0.03 | 0.08 |
| 4 | RF leakage (calibration antenna connector terminated) | 0.086 | 0.086 | Normal | 1 | 1 | 0.09 | 0.09 |
| 5 | QZ ripple with DUT | [0.4] | [0.4] | Rectangular | √3 | 1 | [0.23] | [0.23] |
| 6 | Miscellaneous Uncertainty | 0.00 | 0.00 | Normal | 1 | 1 | 0.00 | 0.00 |
| 20 | System non-linearity | [0.062] | [0.062] | Rectangular | √3 | 1 | [0.04] | [0.04] |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |  |
| 7 | Uncertainty of network analyzer | 0.13 | 0.20 | Normal | 1 | 1 | 0.13 | 0.20 |
| 8 | Mismatch (i.e. reference antenna, network analyser and reference cable) | 0.127 | 0.325 | U-shaped | √2 | 1 | 0.09 | 0.23 |
| 9 | Insertion loss variation | 0.18 | 0.18 | Rectangular | √3 | 1 | 0.10 | 0.10 |
| 4 | RF leakage (calibration antenna connector terminated) | 0.086 | 0.086 | Normal | 1 | 1 | 0.09 | 0.09 |
| 10 | Influence of the calibration antenna feed cable | 0.103 | 0.104 | Rectangular | √3 | 1 | 0.06 | 0.06 |
| 11 | Uncertainty of the absolute gain of the calibration antenna | 0.50 | 0.43 | Rectangular | √3 | 1 | 0.29 | 0.25 |
| 12 | Misalignment of positioning system | 0.00 | 0.00 | Exp. normal | 2 | 1 | 0.00 | 0.00 |
| 13 | Misalignment of calibration antenna & pointing error | 0.05 | 0.05 | Rectangular | √3 | 1 | 0.03 | 0.03 |
| 14 | Rotary joints | 0.00 | 0.00 | U-shaped | √2 | 1 | 0.00 | 0.00 |
| 15 | Longitudinal position uncertainty (i.e. standing wave and imperfect field synthesis) for calibration antenna | 0.12 | 0.12 | Rectangular | √3 | 1 | 0.07 | 0.07 |
| 16 | QZ ripple with calibration antenna | [0.2] | [0.2] | Rectangular | √3 | 1 | [0.12] | [0.12] |
| 17 | Switching uncertainty | 0.02 | 0.02 | Rectangular | √3 | 1 | 0.01 | 0.01 |
| 18 | Field repeatability | 0.06 | 0.12 | Normal | 1 | 1 | [0.06] | [0.12] |
| 19 | Frequency flatness | 0.13 | 0.13 | Rectangular | √3 | 1 | 0.08 | 0.08 |
| Combined standard uncertainty (1σ) [dB] |  |  |  |  |  |  | [0.49] | [0.60] |
| Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  |  | [0.97] | [1.17] |

The agreed summation error (SE) of 0.75 dB (see subclause 10.8) is then root square sum combined with the per point values to give the following result (with 95% confidence level):

Total uncertainty f ≦ 3GHz (95% confidence level): [1.76dB].

Total uncertainty 3GHz < f ≦ 4.2 GHz (95% confidence level): [1.88dB].

10.4.1.5	Summary

Table 10.4.1.5-1: Test system specific measurement uncertainty values for the OTA BS output power test

|  | Expanded uncertainty ue [dB] |  |
| --- | --- | --- |
|  | f ≦ 3GHz | 3GHz < f ≦ 4.2 GHz |
| Indoor Anechoic Chamber | 1,1 | 1,3 |
| Compact Antenna Test Range | 1,4 | 1,5 |
| Near Field chamber |  |  |
| Reverberation chamber | 1.4 | 1.46 |
| Plane Wave Synthesizer | [1.76] | [1.88] |
| Common maximum accepted test system uncertainty | 1.4 | 1.5 |

< End of Changes >

< Unchanged Text Deleted >

< Beginning of Changes >

#### 10.4.2.4B	Plane Wave Synthesizer

10.4.2.4B.1	General

The PWS is described in section 10.4.1.4B.1.

10.4.2.4B.2	Calibration

[Note: This stage may be omitted provided calibration stage has been performed already during output power measurement]

Calibration of the PWS test setup should be carried out in the same manner as described for the TRP procedure in subclause 10.4.1.4B.2.

##### 10.4.2.4B.3	Procedure

1)	Align DUT with boresight of the range antenna.

2)	Configure carrier at a power level according to the manufacturer's declared rated output power and test configuration.

3)	Measure wanted and adjacent channel power for the frequency offsets both side of carrier frequency considering both polarizations of the range antenna. In multiple carrier case only offset frequencies below the lowest and above the offsets highest carrier frequency used shall be measured for ACLR; offsets in between carriers may be subject to CACLR according to the rules in 37.105.

4)	Repeat step 4 for additional points for all necessary points needed for full TRP for the wanted signal and adjacent channel emissions.

5)	Calculate ACLR or CACLR from the wanted signal TRP and the adjacent channel emissions TRP,

The appropriate parameters in step 5 are:

PR_AAS_desired, C: Measured mean power within the desired signal channel bandwidth for each carrier at the measurement equipment connector at C in figure 10.4.1.4B.3-1.

PR_AAS_emission, C: Measured mean emission power in the neighbouring channel bandwidth for each carrier at the measurement equipment connector at C in figure 10.4.1.4B.3-1.

Calculation of Powers Powerd and Powere using following formula:

Powerc = PR_AAS_desired, c+ L

Powere = PR_AAS_emission, C+ L

After calculation of TRP from power as shown in subclause 10.8, calculate ACLR.

##### 10.4.2.4B.4	MU assessment

###### 10.4.2.4B.4.1	MU Budget

The MU budget is the same as in subclause 10.4.1.4B.4.1.

###### 10.4.2.4B.4.2	MU Value

Table 10.4.2.5.4.2-1: Plane Wave Synthesizer uncertainty assessment for absolute ACLR measurements

The MU Value is the same as in subclause 10.4.1.4B.4.2.

The agreed summation error (SE) of 0.75 dB (see subclause 10.8) is then root square sum combined with the per point values to give the following result (with 95% confidence level):

Total uncertainty f ≦ 3GHz (95% confidence level): [1.76dB]

Total uncertainty 3GHz < f ≦ 4.2 GHz (95% confidence level): [1.86dB]

Table 10.4.2.4B.4.1-3: PWS uncertainty assessment for relative ACLR measurement

| UID | Uncertainty Source | Uncertainty value f ≦ 3GHz | Uncertainty value 3GHz ≦ f < 4.2 GHz | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB] f ≦ 3GHz | Standard uncertainty ui [dB] 3GHz < f ≦ 4.2 GHz |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |
| 1 | Misalignment DUT & pointing error | 0.10 | 0.10 | Rectangular | √3 | 1 | 0.06 | 0.06 |
| 2 | RF power measurement equipment (e.g. spectrum analyzer, power meter) | 0.14 | 0.26 | Gaussian | 1 | 1 | 0.14 | 0.26 |
| 5 | QZ ripple with DUT | [0.4] | [0.4] | Rectangular | √3 | 1 | [0.23] | [0.23] |
| 19 | Frequency flatness | 0.13 | 0.13 | Rectangular | √3 | 1 | 0.08 | 0.08 |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |  |
| 7 | Uncertainty of network analyzer | 0.13 | 0.20 | Normal | 1 | 1 | 0.13 | 0.20 |
| 8 | Mismatch (i.e. reference antenna, network analyser and reference cable) | 0.127 | 0.325 | U-shaped | √2 | 1 | 0.09 | 0.23 |
| 8 | Insertion loss variation | 0.18 | 0.18 | Rectangular | √3 | 1 | 0.10 | 0.10 |
| 17 | Switching uncertainty | 0.02 | 0.02 | Rectangular | √3 | 1 | 0.01 | 0.01 |
| Combined standard uncertainty (1σ) [dB] |  |  |  |  |  |  | 0.25 | 0.42 |
| Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  |  | [0.50] | [0.83] |

The agreed summation error (SE) of 0,75 dB is then root square sum combined with the per point values to give the following result (with 95% confidence level):

Total uncertainty f ≦ 3GHz (95% confidence level): [1.55] dB

Total uncertainty 3GHz < f ≦ 4.2 GHz (95% confidence level): [1.68] dB

10.4.2.5	Summary

Table 10.4.2.5-2: Test system specific measurement uncertainty values for the OTA absolute ACLR

|  | Expanded uncertainty ue [dB] |  |
| --- | --- | --- |
|  | f ≦ 3GHz | 3GHz < f ≦ 4.2 GHz |
| Indoor Anechoic Chamber | 1,1 | 1,3 |
| Compact Antenna Test Range | 2,2 | 2,7 |
| Near Field chamber |  |  |
| Reverberation chamber | 1.40 | 1.46 |
| Plane Wave Synthesizer | [1.76] | [1.88] |
| Common maximum accepted test system uncertainty | 2,2 | 2,7 |

Table 10.4.2.5-3: Test system specific measurement uncertainty values for the OTA relative ACLR

|  | Expanded uncertainty ue [dB] |  |
| --- | --- | --- |
|  | f ≦ 3GHz | 3GHz < f ≦ 4.2 GHz |
| Indoor Anechoic Chamber |  |  |
| Compact Antenna Test Range | 1,0 | 1,2 |
| Near Field chamber |  |  |
| Reverberation chamber | 1.40 | 1.46 |
| Plane Wave Synthesizer | [1.55] | [1.68] |
| Common maximum accepted test system uncertainty | 1,0 | 1,2 |

< End of Changes >

< Beginning of Changes >

10.4.3.4B	Plane Wave Synthesizer

##### 10.4.3.4B.1	General

The PWS is described in subclause 10.4.1.4B.1.

##### 10.4.3.4B.2	Calibration

[Note: This stage may be omitted provided calibration stage has been performed already during output power measurement]

Calibration should be carried out using the same procedure as in 10.4.1.4B.2.

Calibration shall be performed individually for each frequency at which unwanted emissions are measured. (This may involve calibration measurement or interpolation between calibration points).

10.4.3.4B.3	Procedure

Reference procedure in subclause 10.4.1.4B.3.

The appropriate parameters in step 5 is the mean power for OBUE test for each carrier arriving at the measurement equipment connector, denoted by PR_AAS_OBUE, C, and calculation of EIRP EIRPe using following formula:

EIRPe = PR_AAS_OBUE_C+ L

##### 10.4.3.4B.4	MU assessment

###### 10.4.3.4B.4.1	MU Budget

The MU budget is the same as in subclause 10.4.14B.4.1.

10.4.3.4B.4.2	 MU Value

The MU value is the same as in subclause 10.4.1.4B.4.2.

10.4.3.5	Summary

Table 10.4.3.5-1: Test system specific measurement uncertainty values for the OTA operating band unwanted emissions measurement

|  | Expanded uncertainty ue [dB] |  |
| --- | --- | --- |
|  | f ≦ 3GHz | 3GHz < f ≦ 4.2 GHz |
| Indoor Anechoic Chamber | 1,1 | 1,3 |
| Compact Antenna Test Range | 1,8 | 2,0 |
| Near Field chamber |  |  |
| Reverberation chamber | 1.40 | 1.46 |
| Plane Wave Synthesizer | [1.76] | [1.88] |
| Common maximum accepted test system uncertainty | 1,8 | 2,0 |

< End of Changes >

< Unchanged Text Deleted >

< Beginning of Changes >

10.4.4.4B	Plane Wave Synthesizer

##### 10.4.4.4B.1	General

The PWS is described in subclause 10.4.1.4B.1.

##### 10.4.4.4B.2	Calibration

[Note: This stage may be omitted provided calibration stage has been performed already during output power measurement]

Calibration should be carried out using the same procedure as in 10.4.1.4B.2.

Calibration shall be performed individually for each frequency at which unwanted emissions are measured. (This may involve calibration measurement or interpolation between calibration points).

10.4.4.4B.3	Procedure

Reference procedure in subclause 10.4.1.4B.3.

The appropriate parameters in step 5 is the mean power for OBUE test for each carrier arriving at the measurement equipment connector, denoted by PR_AAS_SEM, C, and calculation of Power Powere using following formula:

Powere = PR_AAS_SEM_C+ L

##### 10.4.4.4B.4	MU assessment

###### 10.4.4.4B.4.1	MU Budget

The MU budget is the same as in subclause 10.4.1.4B.4.1.

###### 10.4.4.4B.4.2	MU Value

The MU value is the same as in subclause 10.4.1.4B.4.2.

10.4.4.5	Summary

Table 10.4.4.5-1: Test system specific measurement uncertainty values for the OTA SEM

|  | Expanded uncertainty ue [dB] |  |
| --- | --- | --- |
|  | f ≦ 3GHz | 3GHz < f ≦ 4.2 GHz |
| Indoor Anechoic Chamber | 1,2 | 1,3 |
| Compact Antenna Test Range | 1,8 | 2,0 |
| Near Field chamber | [1.4] | [1.6] |
| Reverberation chamber | 1.40 | 1.46 |
| Plane Wave Synthesizer | [1.76] | [1.88] |
| Common maximum accepted test system uncertainty | 1,8 | 2,0 |

< End of Changes >