3GPP TSG-RAN WG4 Meeting #94-e 	R4-200xxxx

Online, 24 Feb - 06 Mar 2020

Source: 	Huawei

Title: 	TP to TR 37.9xx : Tx MU value derivation sub-clause updates

Agenda Item:	8.19

Document for:	Discussion

# Introduction

The external TR for OTA BS testing consolidates all the MU tables from the 3 internal TR’s. The generation of the TR is being used to review and correct all the MU calculations and tables.

The tables are being separated in a spreadsheet to enable consistency of the error values, once agreed they will be copied into the TR.

In preparation for the table to be copied into the TR a TP will be needed with all the MU derivation sections. This TP is extracts the relevant sub-clauses from the TR version 0.1.0 with the intention of being a template for the TP used to introduce all the tables.

Editors notes have been added in curly brackets to explain the process where necessary

Section 9.2. EIRP accuracy has been completed with the latest version of the tables to show how the tables will be imported once they are agreed.

The other sections have not to avoid duplicating the work effort if changes are made.

# TP to TR 39.0xx v0.1.0

--- Start of changes ---

{ EIRP accuracy, Normal test conditions }

#### 9.2.2.3	MU value derivation, FR1

Table 9.2.2.3-1 captures derivation of the expanded measurement uncertainty values for EIRP accuracy measurements in Indoor Anechoic Chamber (Normal test conditions, FR1).

Table 9.2.2.3-1: Indoor Anechoic Chamber measurement uncertainty value derivation for EIRP accuracy measurements, Normal test conditions, FR1

| Indoor anechoic |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB] |  |  |  |  |  |  |  |
|  |  | f<3 GHz | 3<f<4.2 GHz | 4.2<f<6 GHz |  |  |  | f<3 GHz |  | 3<f<4.2 GHz |  |  | 4.2<f<6 GHz |  |  |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| A1-1 | Positioning misalignment between the AAS BS and the reference antenna | 0.03 | 0.03 | 0.03 | Rectangular | 1.73 | 1 | 0.02 |  | 0.02 |  |  | 0.02 |  |  |
| A1-2 | Pointing misalignment between the AAS BS and the receiving antenna | 0.3 | 0.3 | 0.3 | Rectangular | 1.73 | 1 | 0.17 |  | 0.17 |  |  | 0.17 |  |  |
| A1-3 | Quality of quiet zone | 0.1 | 0.1 | 0.1 | Gaussian | 1.00 | 1 | 0.10 |  | 0.10 |  |  | 0.10 |  |  |
| A1-4a | Polarization mismatch between the AAS BS and the receiving antenna | 0.01 | 0.01 | 0.01 | Rectangular | 1.73 | 1 | 0.01 |  | 0.01 |  |  | 0.01 |  |  |
| A1-5 | Mutual coupling between the AAS BS and the receiving antenna | 0 | 0 | 0 | Rectangular | 1.73 | 1 | 0.00 |  | 0.00 |  |  | 0.00 |  |  |
| A1-6 | Phase curvature | 0.05 | 0.05 | 0.05 | Gaussian | 1.00 | 1 | 0.05 |  | 0.05 |  |  | 0.05 |  |  |
| C1-1 | RF power measurement equipment (e.g. spectrum analyzer, power meter) | 0.14 | 0.26 | 0.26 | Gaussian | 1.00 | 1 | 0.14 |  | 0.26 |  |  | 0.26 |  |  |
| A1-7 | Impedance mismatch in the receiving chain | 0.14 | 0.33 | 0.33 | U-shaped | 1.41 | 1 | 0.10 |  | 0.23 |  |  | 0.23 |  |  |
| A1-8 | Random uncertainty | 0.1 | 0.1 | 0.1 | Rectangular | 1.73 | 1 | 0.06 |  | 0.06 |  |  | 0.06 |  |  |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| A1-9 | Impedance mismatch between the receiving antenna and the network analyzer | 0.05 | 0.05 | 0.05 | U-shaped | 1.41 | 1 | 0.04 |  | 0.04 |  |  | 0.04 |  |  |
| A1-10 | Positioning and pointing misalignment between the reference antenna and the receiving antenna | 0.01 | 0.01 | 0.01 | Rectangular | 1.73 | 1 | 0.01 |  | 0.01 |  |  | 0.01 |  |  |
| A1-11 | Impedance mismatch between the reference antenna and the network analyzer. | 0.05 | 0.05 | 0.05 | U-shaped | 1.41 | 1 | 0.04 |  | 0.04 |  |  | 0.04 |  |  |
| A1-3 | Quality of quiet zone | 0.1 | 0.1 | 0.1 | Gaussian | 1.00 | 1 | 0.10 |  | 0.10 |  |  | 0.10 |  |  |
| A1-4b | Polarization mismatch between the reference antenna and the receiving antenna | 0.01 | 0.01 | 0.01 | Rectangular | 1.73 | 1 | 0.01 |  | 0.01 |  |  | 0.01 |  |  |
| A1-5 | Mutual coupling between the reference antenna and the receiving antenna | 0 | 0 | 0 | Rectangular | 1.73 | 1 | 0.00 |  | 0.00 |  |  | 0.00 |  |  |
| A1-6 | Phase curvature | 0.05 | 0.05 | 0.05 | Gaussian | 1.00 | 1 | 0.05 |  | 0.05 |  |  | 0.05 |  |  |
| C1-3 | Uncertainty of the network analyzer | 0.13 | 0.2 | 0.2 | Gaussian | 1.00 | 1 | 0.13 |  | 0.20 |  |  | 0.20 |  |  |
| A1-12 | Influence of the reference antenna feed cable | 0.05 | 0.05 | 0.05 | Rectangular | 1.73 | 1 | 0.03 |  | 0.03 |  |  | 0.03 |  |  |
| A1-13 | Reference antenna feed cable loss measurement uncertainty | 0.06 | 0.06 | 0.06 | Gaussian | 1.00 | 1 | 0.06 |  | 0.06 |  |  | 0.06 |  |  |
| A1-14 | Influence of the receiving antenna feed cable | 0.05 | 0.05 | 0.05 | Rectangular | 1.73 | 1 | 0.03 |  | 0.03 |  |  | 0.03 |  |  |
| C1-4 | Uncertainty of the absolute gain of the reference antenna | 0.50 | 0.43 | 0.43 | Rectangular | 1.73 | 1 | 0.29 |  | 0.25 |  |  | 0.25 |  |  |
| A1-15 | Uncertainty of the absolute gain of the receiving antenna | 0 | 0 | 0 | Rectangular | 1.73 | 1 | 0.00 |  | 0.00 |  |  | 0.00 |  |  |
| Combined standard uncertainty (1σ) [dB] |  |  |  |  |  |  |  |  | 0.44 |  | 0.54 |  |  | 0.54 |  |  |
| Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  |  |  |  | 0.87 |  | 1.06 |  |  | 1.06 |  |  |

--- Next change ---

#### 9.2.3.3	MU value derivation, FR1

Table 9.2.3.3-1 captures derivation of the expanded measurement uncertainty values for EIRP accuracy measurements in CATR (Normal test conditions, FR1).

Table 9.2.3.3-1: CATR MU value derivation for EIRP accuracy measurements, Normal test conditions, FR1

| CATR |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| UID | Uncertainty source | Uncertainty value |  |  |  | Distribution of the probability |  | Divisor based on distribution shape |  | ci |  |  | Standard uncertainty ui [dB] |  |  |  |  |  |  |  |  |  |  |
|  |  | f<3 GHz | 3<f<4.2 GHz | 4.2<f<6 GHz |  |  |  |  |  |  |  |  | f<3 GHz |  |  | 3<f<4.2 GHz |  |  |  | 4.2<f<6 GHz |  |  |  |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| A2-1a | Misalignment DUT & pointing error for EIRP | 0.00 | 0.00 |  | 0.00 |  | Exp. normal |  | 2.00 |  |  | 1 |  | 0.00 |  |  | 0.00 |  |  |  | 0.00 |  |  |  |
| C1-1 | RF power measurement equipment (e.g. spectrum analyzer, power meter) | 0.14 | 0.26 |  | 0.26 |  | Gaussian |  | 1.00 |  |  | 1 |  | 0.14 |  |  | 0.26 |  |  |  | 0.26 |  |  |  |
| A2-2a | Standing wave between DUT and test range antenna | 0.21 | 0.21 |  | 0.21 |  | U-shaped |  | 1.41 |  |  | 1 |  | 0.15 |  |  | 0.15 |  |  |  | 0.15 |  |  |  |
| A2-3 | RF leakage (SGH connector terminated & test range antenna connector cable terminated) | 0.00 | 0.00 |  | 0.00 |  | Normal |  | 1.00 |  |  | 1 |  | 0.00 |  |  | 0.00 |  |  |  | 0.00 |  |  |  |
| A2-4a | QZ ripple DUT | 0.09 | 0.09 |  | 0.09 |  | Normal |  | 1.00 |  |  | 1 |  | 0.09 |  |  | 0.09 |  |  |  | 0.09 |  |  |  |
| A2-12 | Frequency flatness | 0.25 | 0.25 |  | 0.25 |  | Normal |  | 1.00 |  |  | 1 |  | 0.25 |  |  | 0.25 |  |  |  | 0.25 |  |  |  |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| C1-3 | Uncertainty of the network analyzer | 0.13 | 0.20 |  | 0.20 |  | Gaussian |  | 1.00 |  |  | 1 |  | 0.13 |  |  | 0.20 |  |  |  | 0.00 |  |  |  |
| A2-6 | Mismatch of receiver chain | 0.13 | 0.33 |  | 0.33 |  | U-shaped |  | 1.41 |  |  | 1 |  | 0.09 |  |  | 0.23 |  |  |  | 0.23 |  |  |  |
| A2-3 | Insertion loss variation of receiver chain | 0.18 | 0.18 |  | 0.18 |  | Rectangular |  | 1.73 |  |  | 1 |  | 0.10 |  |  | 0.10 |  |  |  | 0.10 |  |  |  |
| A2-3 | RF leakage (SGH connector terminated & test range antenna connector cable terminated) | 0.00 | 0.00 |  | 0.00 |  | Normal |  | 1.00 |  |  | 1 |  | 0.00 |  |  | 0.00 |  |  |  | 0.00 |  |  |  |
| C1-4 | Influence of the calibration antenna feed cable: | 0.02 | 0.02 |  | 0.02 |  | U-shaped |  | 1.41 |  |  | 1 |  | 0.02 |  |  | 0.02 |  |  |  | 0.02 |  |  |  |
| C1-4 | Uncertainty of the absolute gain of the reference antenna | 0.50 | 0.43 |  | 0.43 |  | Rectangular |  | 1.73 |  |  | 1 |  | 0.29 |  |  | 0.25 |  |  |  | 0.25 |  |  |  |
| A2-1b | Misalignment positioning system | 0.00 | 0.00 |  | 0.00 |  | Exp. normal |  | 2.00 |  |  | 1 |  | 0.00 |  |  | 0.00 |  |  |  | 0.00 |  |  |  |
| A2-9 | Misalignment of calibration antenna and test range antenna | 0.50 | 0.50 |  | 0.50 |  | Exp. normal |  | 2.00 |  |  | 1 |  | 0.25 |  |  | 0.25 |  |  |  | 0.25 |  |  |  |
| A2-2b | Rotary Joints | 0.05 | 0.05 |  | 0.05 |  | U-shaped |  | 1.41 |  |  | 1 |  | 0.03 |  |  | 0.03 |  |  |  | 0.03 |  |  |  |
| A2-4b | Standing wave between calibration antenna and test range antenna | 0.09 | 0.09 |  | 0.09 |  | U-shaped |  | 1.41 |  |  | 1 |  | 0.06 |  |  | 0.06 |  |  |  | 0.06 |  |  |  |
| A2-11 | QZ ripple calibration antenna | 0.01 | 0.01 |  | 0.01 |  | Normal |  | 1.00 |  |  | 1 |  | 0.01 |  |  | 0.01 |  |  |  | 0.01 |  |  |  |
| A2-13 | Switching uncertainty | 0.26 | 0.26 |  | 0.26 |  | Rectangular |  | 1.73 |  |  | 1 |  | 0.15 |  |  | 0.15 |  |  |  | 0.15 |  |  |  |
| Combined standard uncertainty (1σ) [dB] |  |  |  |  |  |  |  |  |  |  | 0.57 |  |  |  | 0.65 |  |  |  | 0.61 |  |  |  |
| Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  |  |  |  |  |  | 1.11 |  |  |  | 1.27 |  |  |  | 1.20 |  |  |  |

--- Next change ---

#### 9.2.3.4	MU value derivation, FR2

The MU assessment was carried out using a CATR chamber only however other chamber types are not precluded if suitable MU assessment is done.

A CATR MU budget was assessed in order to determine acceptable MU for the EIRP accuracy measurement in FR2. The CATR test setup and calibration and measurement procedures for FR2 are expected to be similar to those of FR1, although the test chamber dimensions and associated MU values will scale due to the shorter wavelengths and larger relative array apertures.

Table 9.2.3.4-1 captures derivation of the expanded measurement uncertainty values for EIRP accuracy measurements in CATR (Normal test conditions, FR2).

Table 9.2.3.4-1: CATR MU value derivation for EIRP accuracy measurements, Normal test conditions, FR2

| CATR |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| UID | Uncertainty source | Uncertainty value |  | Distribution of the probability | Divisor based on distribution shape |  | ci | Standard uncertainty ui [dB] |  |  |  |  |
|  |  | 24.25<f <29.5GHz | 37<f <40GHz |  |  |  |  | 24.25<f <29.5GHz |  |  | 37<f <40GHz |  |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |  |  |  |
| A2-1a | Misalignment DUT & pointing error (EIRP) | 0.20 | 0.20 | Exp. normal | 2.00 |  | 1 | 0.10 |  |  | 0.10 |  |
| C1-7 | RF power measurement equipment (e.g. spectrum analyzer, power meter) -High power | 0.50 | 0.70 | Normal | 1.00 |  | 1 | 0.50 |  |  | 0.70 |  |
| A2-2a | Standing wave between DUT and test range antenna | 0.03 | 0.03 | U-shaped | 1.41 |  | 1 | 0.02 |  |  | 0.02 |  |
| A2-3 | RF leakage, test range antenna cable connector terminated. | 0.01 | 0.01 | Normal | 1.00 |  | 1 | 0.01 |  |  | 0.01 |  |
| A2-4a | QZ ripple with DUT | 0.40 | 0.40 | Normal | 1.00 |  | 1 | 0.40 |  |  | 0.40 |  |
| A2-12 | Frequency flatness | 0.25 | 0.25 | Normal | 1.00 |  | 1 | 0.25 |  |  | 0.25 |  |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |  |  |  |  |
| C1-3 | Network Analyzer | 0.30 | 0.30 | Normal | 1.00 |  | 1 | 0.30 |  |  | 0.30 |  |
| A2-5a | Mismatch of receiver chain | 0.43 | 0.57 | U-shaped | 1.41 |  | 1 | 0.30 |  |  | 0.40 |  |
| A2-6 | Insertion loss variation in receiver chain | 0.00 | 0.00 | Rectangular | 1.73 |  | 1 | 0.00 |  |  | 0.00 |  |
| A2-3 | RF leakage, (SGH connector terminated & test range antenna connector cable terminated) | 0.01 | 0.01 | Normal | 1.00 |  | 1 | 0.01 |  |  | 0.01 |  |
| A2-7 | Influence of the calibration antenna feed cable | 0.21 | 0.29 | U-shaped | 1.41 |  | 1 | 0.15 |  |  | 0.21 |  |
| C1-4 | SGH Calibration uncertainty | 0.52 | 0.52 | Rectangular | 1.73 |  | 1 | 0.30 |  |  | 0.30 |  |
| A2-8 | Misalignment positioning system | 0.00 | 0.00 | Exp. normal | 2.00 |  | 1 | 0.00 |  |  | 0.00 |  |
| A2-1b | Misalignment of calibration antenna and test range antenna | 0.00 | 0.00 | Exp. normal | 2.00 |  | 1 | 0.00 |  |  | 0.00 |  |
| A2-9 | Rotary joints | 0.00 | 0.00 | U-shaped | 1.41 |  | 1 | 0.00 |  |  | 0.00 |  |
| A2-2b | Standing wave between calibration antenna and test range antenna | 0.09 | 0.09 | U-shaped | 1.41 |  | 1 | 0.06 |  |  | 0.06 |  |
| A2-4b | QZ ripple calibration antenna | 0.01 | 0.01 | Normal | 1.00 |  | 1 | 0.01 |  |  | 0.01 |  |
| A2-11 | Switching uncertainty | 0.10 | 0.10 | Rectangular | 1.73 |  | 1 | 0.06 |  |  | 0.06 |  |
| Combined standard uncertainty (1σ) [dB] |  |  |  |  |  | 0.89 |  |  | 1.06 |
| Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  | 1.74 |  |  | 2.07 |

--- Next change ---

#### 9.2.4.3	MU value derivation, FR1

Table 9.2.4.3-1 captures derivation of the expanded measurement uncertainty values for EIRP accuracy measurements in One Dimensional Compact Range.

Table 9.2.4.3-1: One Dimensional Compact Range MU value derivation for EIRP accuracy measurements, FR1

| One Dimensional Compact Range Chamber |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  | Uncertainty value |  |  |  |  |  |  | Distribution of the probability |  |  | Divisor based on distribution shape |  |  |  | ci |  | Standard uncertainty ui [dB] |  |  |  |  |  |  |
|  |  |  |  | f<3 GHz | 3<f<4.2 GHz |  |  | 4.2<f<6 GHz |  |  |  |  |  |  |  |  |  |  |  | f<3 GHz |  |  | 3<f<4.2 GHz |  |  | 4.2<f<6 GHz |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| A4-1 | Misalignment DUT and pointing error | 0.00 | 0.00 | 0.00 | Exp. normal | 2.00 | 1 | 0.00 | 0.00 | 0.00 |
| A4-2a | Standing wave between DUT and test range antenna | 0.18 | 0.18 | 0.18 | U-shaped | 1.41 | 1 | 0.13 | 0.13 | 0.13 |
| A4-3a | Quiet zone ripple DUT | 0.03 | 0.03 | 0.03 | Gaussian | 1.00 | 1 | 0.03 | 0.03 | 0.03 |
| A4-4a | Phase curvature AAS | 0.01 | 0.01 | 0.01 | Gaussian | 1.00 | 1 | 0.01 | 0.01 | 0.01 |
| A4-5a | Polarization mismatch between DUT and receiving antenna | 0.05 | 0.05 | 0.05 | Rectangular | 1.73 | 1 | 0.03 | 0.03 | 0.03 |
| A4-6a | Mutual coupling between DUT and receiving antenna | 0.00 | 0.00 | 0.00 | Rectangular | 1.73 | 1 | 0.00 | 0.00 | 0.00 |
| C1-1 | RF power measurement equipment (e.g. spectrum analyzer, power meter) | 0.14 | 0.26 | 0.26 | Gaussian | 1.00 | 1 | 0.14 | 0.26 | 0.26 |
| A4-7 | Impedance mismatch in receiving chain | 0.01 | 0.01 | 0.01 | U-shaped | 1.41 | 1 | 0.00 | 0.01 | 0.01 |
| A4-8a | RF leakage (DUT connector terminated and test range antenna connector cable terminated) | 0.00 | 0.00 | 0.00 | Gaussian | 1.00 | 1 | 0.00 | 0.00 | 0.00 |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| A4-9 | Misalignment positioning system | 0.00 | 0.00 | 0.00 | Exp. normal | 2.00 | 1 | 0.00 | 0.00 | 0.00 |
| A4-10 | Pointing error between reference antenna and test range antenna | 0.00 | 0.00 | 0.00 | Rectangular | 1.73 | 1 | 0.00 | 0.00 | 0.00 |
| A4-11 | Impedance mismatch in path to reference antenna | 0.05 | 0.05 | 0.05 | U-shaped | 1.41 | 1 | 0.04 | 0.04 | 0.04 |
| A4-12 | Impedance mismatch in path to compact probe | 0.03 | 0.03 | 0.03 | U-shaped | 1.41 | 1 | 0.02 | 0.02 | 0.02 |
| A4-2b | Standing wave between reference antenna and receiving antenna | 0.09 | 0.09 | 0.09 | U-shaped | 1.41 | 1 | 0.06 | 0.06 | 0.06 |
| A4-3b | Quiet zone ripple reference antenna | 0.18 | 0.18 | 0.18 | Gaussian | 1.00 | 1 | 0.18 | 0.18 | 0.18 |
| A4-4b | Phase curvature refernce antenna | 0.00 | 0.00 | 0.00 | Gaussian | 1.00 | 1 | 0.00 | 0.00 | 0.00 |
| A4-5b | Polarization mismatch between reference antenna and receiving antenna | 0.05 | 0.05 | 0.05 | Rectangular | 1.73 | 1 | 0.03 | 0.03 | 0.03 |
| A4-6a | Mutual coupling between reference antenna and receiving antenna | 0.00 | 0.00 | 0.00 | Rectangular | 1.73 | 1 | 0.00 | 0.00 | 0.00 |
| C1-1 | RF power measurement equipment (e.g. spectrum analyzer, power meter) | 0.14 | 0.26 | 0.26 | Gaussian | 1.00 | 1 | 0.14 | 0.26 | 0.26 |
| A4-13 | Influence of the reference antenna feed cable (flexing cables, adapters, attenuators, connector repeatability) | 0.08 | 0.08 | 0.08 | Rectangular | 1.73 | 1 | 0.05 | 0.05 | 0.05 |
| A4-14 | Mismatch of receiver chain | 0.20 | 0.30 | 0.30 | U-shaped | 1.41 | 1 | 0.14 | 0.21 | 0.21 |
| A4-15 | Insertion loss of receiver chain | 0.18 | 0.18 | 0.18 | Rectangular | 1.73 | 1 | 0.10 | 0.10 | 0.10 |
| C1-4 | Uncertainty of the absolute gain of the reference antenna | 0.50 | 0.43 | 0.43 | Rectangular | 1.73 | 1 | 0.29 | 0.25 | 0.25 |
| A4-8b | RF leakage (SGH connector terminated and test range antenna connector cable terminated. | 0.00 | 0.00 | 0.00 | Gaussian | 1.00 | 1 | 0.00 | 0.00 | 0.00 |
| Combined standard uncertainty (1σ) [dB] |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 0.46 |  |  |  | 0.56 | 0.56 |  |  |
| Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 0.90 |  |  |  | 1.10 | 1.10 |  |  |

--- Next change ---

#### 9.2.5.3	MU value derivation, FR1

Table 9.2.5.3-1 captures derivation of the expanded measurement uncertainty values for EIRP accuracy measurements in Near Field Test Range.

Standard uncertainty values for the signal generator, network analyzer and reference antenna are according to the test equipment uncertainty values, as captured in annex C.

Table 9.2.5.3-1: NFTR measurement uncertainty value derivation for EIRP accuracy measurements, FR1

| Near field |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape |  | ci |  | Standard uncertainty ui [dB] |  |  |  |  |  |  |
|  |  | f<3 GHz | 3<f<4.2 GHz | 4.2<f<6 GHz |  |  |  |  |  | f<3 GHz |  |  |  | 3<f<4.2 GHz |  | 4.2<f<6 GHz |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |  |  |  |  |
| A3-1 | Axes Intersection | 0.00 | 0.00 | 0.00 | Gaussian | 1.00 |  | 1 |  | 0.00 |  |  |  | 0.00 |  | 0.00 |
| A3-2 | Axes Orthogonality | 0.00 | 0.00 | 0.00 | Gaussian | 1.00 |  | 1 |  | 0.00 |  |  |  | 0.00 |  | 0.00 |
| A3-3 | Horizontal Pointing | 0.00 | 0.00 | 0.00 | Gaussian | 1.00 |  | 1 |  | 0.00 |  |  |  | 0.00 |  | 0.00 |
| A3-4 | Probe Vertical Position | 0.00 | 0.00 | 0.00 | Gaussian | 1.00 |  | 1 |  | 0.00 |  |  |  | 0.00 |  | 0.00 |
| A3-5 | Probe H/V pointing | 0.00 | 0.00 | 0.00 | Gaussian | 1.00 |  | 1 |  | 0.00 |  |  |  | 0.00 |  | 0.00 |
| A3-6 | Measurement Distance | 0.00 | 0.00 | 0.00 | Gaussian | 1.00 |  | 1 |  | 0.00 |  |  |  | 0.00 |  | 0.00 |
| A3-7 | Amplitude and Phase Drift | 0.00 | 0.00 | 0.00 | Gaussian | 1.00 |  | 1 |  | 0.00 |  |  |  | 0.00 |  | 0.00 |
| A3-8 | Amplitude and Phase Noise | 0.02 | 0.02 | 0.02 | Gaussian | 1.00 |  | 1 |  | 0.02 |  |  |  | 0.02 |  | 0.02 |
| A3-9 | Leakage and Crosstalk | 0.00 | 0.00 | 0.00 | Gaussian | 1.00 |  | 1 |  | 0.00 |  |  |  | 0.00 |  | 0.00 |
| A3-10 | Amplitude Non-Linearity | 0.04 | 0.04 | 0.04 | Gaussian | 1.00 |  | 1 |  | 0.04 |  |  |  | 0.04 |  | 0.04 |
| A3-11 | Amplitude and Phase Shift in rotary joints | 0.00 | 0.00 | 0.00 | Gaussian | 1.00 |  | 1 |  | 0.00 |  |  |  | 0.00 |  | 0.00 |
| A3-12 | Channel Balance Amplitude and Phase | 0.00 | 0.00 | 0.00 | Gaussian | 1.00 |  | 1 |  | 0.00 |  |  |  | 0.00 |  | 0.00 |
| A3-13 | Probe Polarization Amplitude and Phase | 0.00 | 0.00 | 0.00 | Gaussian | 1.00 |  | 1 |  | 0.00 |  |  |  | 0.00 |  | 0.00 |
| A3-14 | Probe Pattern Knowledge | 0.00 | 0.00 | 0.00 | Gaussian | 1.00 |  | 1 |  | 0.00 |  |  |  | 0.00 |  | 0.00 |
| A3-15 | Multiple Reflections | 0.00 | 0.00 | 0.00 | Gaussian | 1.00 |  | 1 |  | 0.00 |  |  |  | 0.00 |  | 0.00 |
| A3-16 | Room Scattering | 0.09 | 0.09 | 0.09 | Gaussian | 1.00 |  | 1 |  | 0.09 |  |  |  | 0.09 |  | 0.09 |
| A3-17 | DUT support Scattering | 0.00 | 0.00 | 0.00 | Gaussian | 1.00 |  | 1 |  | 0.00 |  |  |  | 0.00 |  | 0.00 |
| A3-18 | Scan Area Truncation | 0.00 | 0.00 | 0.00 | Gaussian | 1.00 |  | 1 |  | 0.00 |  |  |  | 0.00 |  | 0.00 |
| A3-19 | Sampling Point Offset | 0.01 | 0.01 | 0.01 | Gaussian | 1.00 |  | 1 |  | 0.01 |  |  |  | 0.01 |  | 0.01 |
| A3-20 | Spherical Mode Truncation | 0.02 | 0.02 | 0.02 | Gaussian | 1.00 |  | 1 |  | 0.02 |  |  |  | 0.02 |  | 0.02 |
| A3-21 | Positioning | 0.03 | 0.03 | 0.03 | Rectangular | 1.73 |  | 1 |  | 0.02 |  |  |  | 0.02 |  | 0.02 |
| A3-22 | Probe Array Uniformity | 0.06 | 0.06 | 0.06 | Gaussian | 1.00 |  | 1 |  | 0.06 |  |  |  | 0.06 |  | 0.06 |
| A3-23 | Mismatch of receiver chain | 0.28 | 0.28 | 0.28 | U-Shaped | 1.41 |  | 1 |  | 0.20 |  |  |  | 0.20 |  | 0.20 |
| A3-24 | Insertion loss of receiver chain | 0.00 | 0.00 | 0.00 | Gaussian | 1.00 |  | 1 |  | 0.00 |  |  |  | 0.00 |  | 0.00 |
| A3-25 | Uncertainty of the absolute gain of the probe antenna | 0.00 | 0.00 | 0.00 | Gaussian | 1.00 |  | 1 |  | 0.00 |  |  |  | 0.00 |  | 0.00 |
| C1-1 | RF power measurement equipment (e.g. spectrum analyzer, power meter) | 0.14 | 0.26 | 0.26 | Gaussian | 1.00 |  | 1 |  | 0.14 |  |  |  | 0.26 |  | 0.26 |
| A3-26 | Measurement repeatability - positioning repeatability | 0.15 | 0.15 | 0.15 | Gaussian | 1.00 |  | 1 |  | 0.15 |  |  |  | 0.15 |  | 0.15 |
| A3-33 | Test system frequency flatness | 0.25 | 0.25 | 0.25 | Normal | 1.00 |  | 1 |  | 0.25 |  |  |  | 0.25 |  | 0.25 |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |  |  |  |  |  |
| C1-3 | Uncertainty of the network analyzer | 0.13 | 0.20 | 0.20 | Gaussian | 1.00 |  | 1 |  | 0.13 |  |  |  | 0.20 |  | 0.20 |
| A3-27 | Mismatch of receiver chain | 0.00 | 0.00 | 0.00 | Gaussian | 1.00 |  | 1 |  | 0.00 |  |  |  | 0.00 |  | 0.00 |
| A3-28 | Insertion loss of receiver chain | 0.00 | 0.00 | 0.00 | Gaussian | 1.00 |  | 1 |  | 0.00 |  |  |  | 0.00 |  | 0.00 |
| A3-29 | Mismatch in the connection of the calibration antenna | 0.02 | 0.02 | 0.02 | U-Shaped | 1.41 |  | 1 |  | 0.01 |  |  |  | 0.01 |  | 0.01 |
| A3-30 | Influence of the calibration antenna feed cable | 0.00 | 0.00 | 0.00 | Gaussian | 1.00 |  | 1 |  | 0.00 |  |  |  | 0.00 |  | 0.00 |
| A3-31 | Influence of the probe antenna cable | 0.00 | 0.00 | 0.00 | Gaussian | 1.00 |  | 1 |  | 0.00 |  |  |  | 0.00 |  | 0.00 |
| C1-4 | Uncertainty of the absolute gain of the reference antenna | 0.50 | 0.43 | 0.43 | Rectangular | 1.73 |  | 1 |  | 0.29 |  |  |  | 0.25 |  | 0.25 |
| A3-32 | Short term repeatability | 0.09 | 0.09 | 0.09 | Gaussian | 1.00 |  | 1 |  | 0.09 |  |  |  | 0.09 |  | 0.09 |
| Combined standard uncertainty (1σ) [dB] |  |  |  |  |  |  | 0.52 |  | 0.56 |  |  | 0.56 |  |  |
| Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  |  | 1.01 |  | 1.10 |  |  | 1.10 |  |  |

--- Next change ---

#### 9.2.6.3	MU value derivation, FR1

Table 9.2.6.3-1 captures derivation of the expanded measurement uncertainty values for EIRP accuracy measurements in PWS.

Table 9.2.6.3-1: PWS measurement uncertainty value derivation for EIRP accuracy measurements, FR1

| Plane wave synthesizer |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape |  | ci |  | Standard uncertainty ui [dB] |  |  |  |  |  |  |
|  |  | f<3 GHz | 3<f<4.2 GHz | 4.2<f<6 GHz |  |  |  |  |  | f<3 GHz |  |  |  | 3<f<4.2 GHz |  | 4.2<f<6 GHz |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |  |  |
| A7-1a | Misalignment DUT & pointing error | 0.10 | 0.10 | 0.10 | Rectangular | 1.73 |  | 1 |  | 0.06 |  |  |  | 0.06 |  | 0.06 |
| C1-1 | RF power measurement equipment (e.g. spectrum analyzer, power meter) | 0.14 | 0.26 | 0.26 | Gaussian | 1.00 |  | 1 |  | 0.14 |  |  |  | 0.26 |  | 0.26 |
| A7-2a | Longitudinal position uncertainty (i.e. standing wave and imperfect field synthesis) for DUT antenna | 0.05 | 0.14 | 0.14 | Rectangular | 1.73 |  | 1 |  | 0.03 |  |  |  | 0.08 |  | 0.08 |
| A7-3 | RF leakage (calibration antenna connector terminated) | 0.09 | 0.09 | 0.09 | Normal | 1.00 |  | 1 |  | 0.09 |  |  |  | 0.09 |  | 0.09 |
| A7-4a | QZ ripple with DUT | 0.42 | 0.43 | 0.43 | Rectangular | 1.73 |  | 1 |  | 0.24 |  |  |  | 0.25 |  | 0.25 |
| A7-5 | Miscellaneous Uncertainty | 0.00 | 0.00 | 0.00 | Normal | 1.00 |  | 1 |  | 0.00 |  |  |  | 0.00 |  | 0.00 |
| A7-14 | System non-linearity | 0.06 | 0.06 | 0.06 | Rectangular | 1.73 |  | 1 |  | 0.04 |  |  |  | 0.04 |  | 0.04 |
| A7-13 | Frequency Flatness | 0.13 | 0.13 | 0.13 | Rectangular | 1.73 |  | 1 |  | 0.08 |  |  |  | 0.08 |  | 0.08 |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |  |  |  |
| C1-3 | Uncertainty of the network analyzer | 0.13 | 0.20 | 0.20 | Gaussian | 1.00 |  | 1 |  | 0.13 |  |  |  | 0.20 |  | 0.20 |
| A7-6 | Mismatch (i.e. reference antenna, network analyzer and reference cable) | 0.13 | 0.33 | 0.33 | U-shaped | 1.41 |  | 1 |  | 0.09 |  |  |  | 0.23 |  | 0.23 |
| A7-7 | Insertion loss variation | 0.18 | 0.18 | 0.18 | Rectangular | 1.73 |  | 1 |  | 0.10 |  |  |  | 0.10 |  | 0.10 |
| A7-3 | RF leakage (calibration antenna connector terminated) | 0.09 | 0.09 | 0.09 | Normal | 1.00 |  | 1 |  | 0.09 |  |  |  | 0.09 |  | 0.09 |
| A7-8 | Influence of the calibration antenna feed cable | 0.10 | 0.10 | 0.10 | Rectangular | 1.73 |  | 1 |  | 0.06 |  |  |  | 0.06 |  | 0.06 |
| C1-4 | Uncertainty of the absolute gain of the reference antenna | 0.50 | 0.43 | 0.43 | Rectangular | 1.73 |  | 1 |  | 0.29 |  |  |  | 0.25 |  | 0.25 |
| A7-9 | Misalignment of positioning system | 0.00 | 0.00 | 0.00 | Exp. normal | 2.00 |  | 1 |  | 0.00 |  |  |  | 0.00 |  | 0.00 |
| A7-1b | Misalignment of calibration antenna & pointing error | 0.05 | 0.05 | 0.05 | Rectangular | 1.73 |  | 1 |  | 0.03 |  |  |  | 0.03 |  | 0.03 |
| A7-10 | Rotary joints | 0.00 | 0.00 | 0.00 | U-shaped | 1.73 |  | 1 |  | 0.00 |  |  |  | 0.00 |  | 0.00 |
| A7-2b | Longitudinal position uncertainty (i.e. standing wave and imperfect field synthesis) for calibration antenna | 0.12 | 0.12 | 0.12 | Rectangular | 1.73 |  | 1 |  | 0.07 |  |  |  | 0.07 |  | 0.07 |
| A7-4a | QZ ripple with calibration antenna | 0.20 | 0.20 | 0.20 | Rectangular | 1.73 |  | 1 |  | 0.12 |  |  |  | 0.12 |  | 0.12 |
| A7-11 | Switching uncertainty | 0.02 | 0.02 | 0.02 | Rectangular | 1.73 |  | 1 |  | 0.01 |  |  |  | 0.01 |  | 0.01 |
| A7-12 | Field repeatability | 0.06 | 0.12 | 0.12 | Normal | 1.00 |  | 1 |  | 0.06 |  |  |  | 0.12 |  | 0.12 |
| Combined standard uncertainty (1σ) [dB] |  |  |  |  |  |  | 0.50 |  | 0.60 |  |  | 0.60 |  |  |
| Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  |  | 0.98 |  | 1.18 |  |  | 1.18 |  |  |

--- Next change ---

### 9.2.7	Maximum accepted test system uncertainty

Maximum test system uncertainties derivation methodology was described in subclause 5.1. The maximum accepted test system uncertainty values was derived based on test system specific values.

According to the methodology referred above, the common maximum accepted test system uncertainty values for the EIRP accuracy test can be derived from values captured in table 9.2.7-1, separately for each of the defined frequency ranges. The common maximum accepted test system uncertainty values are applicable for all test methods addressing EIRP accuracy test requirement in Normal test conditions. Based on input values in table 9.2.7-1, the expanded uncertainty ue (1.96σ - confidence interval of 95 %) values were derived for frequency ranges as below:

- For the frequency range up to 4.2 GHz (with the breakdown point at 3 GHz), the same MU values as for E-UTRA in TS 37.145-2 [4] were adopted also for NR operation below 4.2 GHz. It is expected that the test chamber setup, calibration and measurement procedures for E-UTRA and NR will be highly similar. All uncertainty factors were judged to be the same. The MU value was thus agreed to be 1.1 dB for up to 3 GHz bands.

- For the frequency range 4.2 - 6 GHz, all MU factors including instrumentation related MU were judged to be the same as for the 3 - 4.2 GHz range, and thus the total MU for 4.2 – 6 GHz is the same as for 3 - 4.2 GHz. This assessment was made under the assumption of testing BS designed for licensed spectrum; for unlicensed spectrum the MU may differ. The MU value was thus agreed to be 1.3 dB for 3 – 6 GHz bands. The MU in 4.2 - 6 GHz is valid for BS designed to operate in licensed spectrum.

- Based on CATR inputs in subcluase 9.2.3.4, for the frequency range 24.25 < f < 29.5 GHz the MU was decided to be 1.7 dB.

- Based on CATR inputs in subcluase 9.2.3.4, for the frequency range 37 < f < 40 GHz the MU was decided to be 2.0 dB.

Table 9.2.7-1: OTA test system specific measurement uncertainty values for the EIRP accuracy, FR1, Normal test conditions

|  | Expanded uncertainty [dB] |  |  |
| --- | --- | --- | --- |
|  | f<3 GHz | 3<f<4.2 GHz | 4.2<f<6 GHz |
| Indoor Anechoic Chamber | 0.87 | 1.06 | 1.06 |
| Compact Antenna Test Range | 1.11 | 1.27 | 1.20 |
| One Dimensional Compact Range Chamber | 0.90 | 1.10 | 1.10 |
| Near Field Test Range | 1.01 | 1.10 | 1.10 |
| PWS | 0.98 | 1.18 | 1.18 |
| Common maximum accepted test system uncertainty | 1.10 | 1.30 | 1.30 |

Table 9.2.7-2: OTA test system specific measurement uncertainty values for the EIRP accuracy, FR2, Normal test conditions

|  | Expanded uncertainty [dB] |  |
| --- | --- | --- |
|  | 24.25<f <29.5GHz | 37<f <40GHz |
| Indoor Anechoic Chamber |  |  |
| Compact Antenna Test Range | 1.74 | 2.07 |
| One Dimensional Compact Range Chamber |  |  |
| Near Field Test Range |  |  |
| PWS |  |  |
| Agreed value | 1.70 | 2.00 |

An overview of the MU values for all the requirements is captured in clause 16.

### 9.2.8	Test Tolerance for EIRP accuracy, Normal test conditions

Considering the methodology described in subclause 5.1, Test Tolerance values for EIRP were derived based on values captured in subclause 9.2.7.

The TT was decided to be the same as the MU for EIRP accuracy in FR1.

The TT was decided to be the same as the MU for EIRP accuracy in FR2.

Frequency range specific Test Tolerance values for the EIRP accuracy test are defined in table 9.2.8-1.

Table 9.2.8-1: Test Tolerance values for the EIRP accuracy in Normal test conditions FR1

|  | f ≦ 3GHz | 3GHz < f ≦ 4.2 GHz | 4.2<f<6 GHz |
| --- | --- | --- | --- |
| Test Tolerance (dB) | 1.1 | 1.3 | 1.3 |

Table 9.2.8-2: Test Tolerance values for the EIRP accuracy in Normal test conditions FR2

|  | 24.25<f <29.5GHz | 37<f <40GHz |
| --- | --- | --- |
| Test Tolerance (dB) | 1.7 | 2.0 |

An overview of the TT values for all the requirements is captured in clause 17.

--- Next change ---

#### 9.3.2.3	MU value derivation

Table 9.3.2.3-1: Indoor Anechoic Chamber measurement uncertainty value derivation for EIRP accuracy measurements in Extreme test conditions

--- Next change ---

#### 9.3.3.4	MU value derivation, FR1

Table 9.3.3.4-1: CATR MU value derivation for EIRP accuracy measurements in Extreme test conditions

--- Next change ---

#### 9.3.3.5	MU value derivation, FR2

Table 9.3.3.5-1: CATR MU value derivation for EIRP accuracy measurements in Extreme test conditions, FR2

--- Next change ---

### 9.3.4	Maximum accepted test system uncertainty

For the frequency range up to 4.2 GHz, the same MU values as for E-UTRA in TS 37.145-2 [4] were adopted. It is expected that the test chamber setup, calibration and measurement procedures for E-UTRA and NR will be highly similar. All uncertainty factors were judged to be the same.

For the frequency range 4.2 - 6 GHz, all MU factors, including instrumentation related MU were judged to be the same as for the 3 - 4.2 GHz range, and thus the total MU for 4.2 – 6 GHz is the same as for 3 - 4.2 GHz. This assessment was made under the assumption of testing BS designed for licensed spectrum; for unlicensed spectrum the MU may differ.

The MU value was thus agreed to be 2.5 dB for up to 3 GHz bands and 2.6 dB for 3 – 6 GHz bands. The MU in 4.2 - 6 GHz is valid for BS designed to operate in licensed spectrum.

For FR2, for the direct far field method the MU budget is very similar to the existing MU budget for the EIRP accuracy requirement. However there are a number of additional sources of uncertainty due to the environmental enclosure that need to be added to the budget. For FR2 only the CATR MU budget has been assessed however other suitable camber types are not precluded.

Based on the above evaluation, the MU was decided to be 3.1 dB for the frequency range 24.25 < f < 29.5 GHz and 3.3 dB for the frequency range 37 < f < 40 GHz.

Test system specific measurement uncertainty values for the EIRP accuracy in Extreme test conditions are summarised in table 9.3.4-1.

Table 9.3.4-1: Test system specific measurement uncertainty values for the EIRP accuracy in Extreme test conditions, FR1

|  | Expanded uncertainty ue (dB) |  |  |
| --- | --- | --- | --- |
|  | f ≦ 3GHz | 3GHz < f ≦ 4.2 GHz | 4.2GHz < f ≦ 6GHz |
| Indoor Anechoic Chamber |  |  |  |
| Compact Antenna Test Range |  |  |  |
| Common maximum accepted test system uncertainty | 2.5 | 2.6 | 2.6 |

Table 9.3.4-2: Test system specific measurement uncertainty values for the EIRP accuracy in Extreme test conditions, FR2

|  | Expanded uncertainty ue (dB) |  |  |
| --- | --- | --- | --- |
|  | 24.25<f <29.5GHz | 37GHz < f ≦ 40 GHz |
| Compact Antenna Test Range |  |  |
| Common maximum accepted test system uncertainty | 3.1 | 3.3 |

An overview of the MU values for all the requirements is captured in clause 16.

### 9.3.5	Test Tolerance for EIRP accuracy, Extreme test conditions

Considering the methodology described in subclause 5.1, Test Tolerance values for EIRP were derived based on values captured in subclause 9.3.4.

The TT was decided to be the same as the MU for EIRP accuracy in FR1.

The TT was decided to be the same as the MU for EIRP accuracy in FR2.

Frequency range specific Test Tolerance values for the EIRP accuracy test are defined in table 9.3.5-1.

Table 9.3.5-1: Test Tolerance values for the EIRP accuracy in Extreme test conditions, FR1

|  | f ≦ 3GHz | 3GHz < f ≦ 4.2 GHz | 4.2GHz < f ≦ 6GHz |
| --- | --- | --- | --- |
| Test Tolerance (dB) | 2.5 | 2.6 | 2.6 |

Table 9.3.5-2: Test Tolerance values for the EIRP accuracy in Extreme test conditions, FR2

|  | 24.25<f <29.5GHz | 37GHz < f ≦ 40 GHz |
| --- | --- | --- |
| Test Tolerance (dB) | 3.1 | 3.3 |

An overview of the TT values for all the requirements is captured in clause 17.

--- Next change ---

#### 9.4.2.3	MU value derivation

As the DL RS power is an absolute measurement most of the uncertainties form the EIRP accuracy remain the same. Also it can be noted that the measured signal is a wanted signal and hence will be beam formed in the same way as the wanted signal, hence any errors which may be dependent on the beam shape will be the same as for the EIRP accuracy measurement.

Table 9.4.2.4-1: Indoor Anechoic Chamber measurement uncertainty value derivation for OTA E-UTRA DL RS power measurement

--- Next change ---

#### 9.4.3.3	MU value derivation, FR1

As the DL RS power is an absolute measurement most of the uncertainties form the EIRP accuracy remain the same. Also it can be noted that the measured signal is a wanted signal and hence will be beam formed in the same way as the wanted signal, hence any errors which may be dependent on the beam shape will be the same as for the EIRP accuracy measurement.

Table 6.4.3.4-1: CATR MU value derivation for OTA E-UTRA DL RS power measurement

--- Next change ---

#### 9.4.4.3	MU value derivation

Table 9.4.4.3-1: NFTR measurement uncertainty value derivation for OTA E-UTRA DL RS power measurement

--- Next change ---

### 9.4.5	Maximum accepted test system uncertainty

Maximum test system uncertainties derivation methodology was described in subclause 5.1. The maximum accepted test system uncertainty values was derived based on test system specific values.

According to the methodology referred above, the common maximum accepted test system uncertainty values for the OTA E-UTRA DL RS power test can be derived from values captured in table 6.4.5-1, separately for each of the defined frequency ranges. The common maximum values are applicable for all test methods addressing OTA E-UTRA DL RS power test requirement. Based on the input values, the expanded uncertainty ue (1.96σ - confidence interval of 95 %) values were derived.

Table 6.4.5-1: Test system specific measurement uncertainty values for the OTA E-UTRA DL RS power test

|  | Expanded uncertainty ue (dB) |  |  |
| --- | --- | --- | --- |
|  | f ≦ 3GHz | 3GHz < f ≦ 4.2 GHz | 4.2GHz < f ≦ 6GHz |
| Indoor Anechoic Chamber |  |  |  |
| Compact Antenna Test Range |  |  |  |
| Near Field Test Range |  |  |  |
| Common maximum accepted test system uncertainty | 1.3 | 1.5 | 1.5 |

### 9.4.5	Test Tolerance for OTA E-UTRA DL RS power

The TT was decided to be the same as the MU in FR1

Table 9.4.5-1: Test Tolerance values for the OTA E-UTRA DL RS power

|  | f ≦ 3GHz | 3GHz < f ≦ 4.2 GHz | 4.2GHz < f ≦ 6GHz |
| --- | --- | --- | --- |
| Test Tolerance (dB) | 2.5 | 2.6 | 2.6 |

--- Next change ---

#### 9.5.2.3	MU value derivation

As the output power dynamics are relative measurements most of the uncertainties form the EIRP accuracy cancel out as the same error will be applied to both of the measured OTA signals.

This includes all calibration errors, misalignment errors, impedance mismatch and mutual coupling.

As the both the measured OTA signal will have the same beam pattern quiet zone errors, phase curvature errors also can be expected to be the same for both signals.

The uncertainty budget descriptions are the same as those in table 6.4.2.3-1 with the addition descriptions in tble 6.5.2.3-1.

The MU uncertainty assessment is shown in table 6.5.2.4-1, zero values have been omitted in the table for the sake of space, but still be considered as part of the budget.

Table 9.5.2.3-1: Indoor Anechoic Chamber measurement uncertainty value derivation for OTA total power dynamic range measurement

The same uncertainty assessments have been carried out for the UTRA dynamic range requirements, i.e.

- Power control steps,

- Power control dynamic range,

- Total power dynamic range,

- IPDL Time mask.

In each case the uncertainty for the conducted measurement is the same as that for the conducted MU in TS 25.141 [2] as follows:

- Power control steps, Uncertainty of conducted measurement = 0.1 dB, Expanded OTA uncertainty = 0.15 dB.

- Power control dynamic range, Uncertainty of conducted measurement = 1.1 dB, Expanded OTA uncertainty = 1.11 dB.

- Total power dynamic range, Uncertainty of conducted measurement = 0.3 dB, Expanded OTA uncertainty = 0.32 dB.

- IPDL Time mask, Uncertainty of conducted measurement = 0.7dB, Expanded OTA uncertainty = 0.71 dB.

--- Next change ---

#### 9.5.3.3	MU value derivation

As the output power dynamics are relative measurements most of the uncertainties form the EIRP accuracy cancel out as the same error will be applied to both of the measured OTA signals.

This includes all calibration errors, misalignment errors, impedance mismatch and mutual coupling.

As the both the measured OTA signal will have the same beam pattern quiet zone errors, phase curvature errors also can be expected to be the same for both signals.

The uncertainty budget descriptions are the same as those in table 6.4.3.3-1 with the addition descriptions in table 6.5.3.3-1.

The MU uncertainty assessment is shown in table 6.5.3.4-1, zero values have been omitted in the table for the sake of space, but still be considered as part of the budget.

Table 9.5.3.3-1: CATR MU value derivation for OTA total power dynamic range measurement

The same uncertainty assessments have been carried out for the UTRA dynamic range requirements, i.e.

- Power control steps,

- Power control dynamic range,

- Total power dynamic range,

- IPDL Time mask.

In each case the uncertainty for the conducted measurement is the same as that for the conducted MU in TS 25.141 [2] as follows:

- Power control steps: Uncertainty of conducted measurement = 0.1 dB, Expanded OTA uncertainty = 0.1 dB.

- Power control dynamic range: Uncertainty of conducted measurement = 1.1 dB, Expanded OTA uncertainty = 1.1 dB.

- Total power dynamic range: Uncertainty of conducted measurement = 0.3 dB, Expanded OTA uncertainty = 0.3 dB.

- IPDL Time mask: Uncertainty of conducted measurement = 0.7 dB, Expanded OTA uncertainty = 0.7 dB.

- --- Next change ---

#### 9.5.4.3	MU value derivation

As the output power dynamics are relative measurements most of the uncertainties form the EIRP accuracy cancel out as the same error will be applied to both of the measured OTA signals.

The uncertainty budget descriptions are the same as those in table 6.4.4.4-1 (excluding uncertainties from the NF to FF transformation, since the transformation is not needed) with the addition descriptions in table 6.5.4.3-1:

The MU uncertainty assessment is shown in table 6.5.4.4-1, zero values have been omitted in the table for the sake of space, but still be considered as part of the budget.

Table 9.5.4.3-1: NFTR MU value derivation for OTA E-UTRA total power dynamic range measurement

--- Next change ---

### 9.5.5	Maximum accepted test system uncertainty

Maximum test system uncertainties derivation methodology was described in subclause 5.1. The maximum accepted test system uncertainty values was derived based on test system specific values.

According to the methodology referred above, the common maximum accepted test system uncertainty values for the OTA output power dynamics tests can be derived from values captured in tables 6.5.5-1 to 6.5.5-5, separately for each of the defined frequency ranges. The common maximum values are applicable for all test methods addressing certain OTA output power dynamics test requirement. Based on the input values, the expanded uncertainty ue (1.96σ - confidence interval of 95 %) values were derived.

The output power dynamic range MU for FR1 for up to 4.2 GHz was agreed to be the same as for eAAS. It is expected that the measurement chamber setup, calibration and measurement procedures and the MU budget will be identical for E-UTRA and NR.

Furthermore, for the frequency range 4.2 – 6 GHz, the chamber and instrument uncertainties are the same as those for the frequency range 3 – 4.2 GHz, assuming testing of a BS designed for operation in licensed spectrum. The MU value was agreed to be +/- 0.4 dB for all bands from 0 to 6 GHz.

For FR2: Similarly to FR1, since the OTA output power dynamic range requirement is a relative measurement, only the test equipment uncertainty is of importance. Based on expected test equipment uncertainty, the output power dynamic range MU for FR2 was decided to be +/- 0.4 dB (same as for FR1).

Table 9.5.5-1: Test system specific measurement uncertainty values for the OTA E-UTRA and NR dynamic range test

|  | Expanded uncertainty ue (dB) |  |  |
| --- | --- | --- | --- |
|  | f ≦ 3GHz | 3GHz < f ≦ 4.2 GHz | 4.2GHz < f ≦ 6GHz |
| Indoor Anechoic Chamber |  |  |  |
| Compact Antenna Test Range |  |  |  |
| Common maximum accepted test system uncertainty | 0.4 | 0.4 | 0.4 |

Table 9.5.5-2: Test system specific measurement uncertainty values for the UTRA power control steps test

|  | Expanded uncertainty ue (dB) |  |  |
| --- | --- | --- | --- |
|  | f ≦ 3GHz | 3GHz < f ≦ 4.2 GHz | 4.2GHz < f ≦ 6GHz |
| Indoor Anechoic Chamber |  |  |  |
| Compact Antenna Test Range |  |  |  |
| Common maximum accepted test system uncertainty | 0.1 | 0.1 | 0.1 |

Table 9.5.5-3: Test system specific measurement uncertainty values for the UTRA power control dynamic range test

|  | Expanded uncertainty ue (dB) |  |  |
| --- | --- | --- | --- |
|  | f ≦ 3GHz | 3GHz < f ≦ 4.2 GHz | 4.2GHz < f ≦ 6GHz |
| Indoor Anechoic Chamber |  |  |  |
| Compact Antenna Test Range |  |  |  |
| Common maximum accepted test system uncertainty | 1.1 | 1.1 | 1.1 |

Table 9.5.5-4: Test system specific measurement uncertainty values for the UTRA total power dynamic range test

|  | Expanded uncertainty ue (dB) |  |  |
| --- | --- | --- | --- |
|  | f ≦ 3GHz | 3GHz < f ≦ 4.2 GHz | 4.2GHz < f ≦ 6GHz |
| Indoor Anechoic Chamber |  |  |  |
| Compact Antenna Test Range |  |  |  |
| Common maximum accepted test system uncertainty | 0.3 | 0.3 | 0.3 |

Table 9.5.5-5: Test system specific measurement uncertainty values for the UTRA IPDL time mask test

|  | Expanded uncertainty ue (dB) |  |  |
| --- | --- | --- | --- |
|  | f ≦ 3GHz | 3GHz < f ≦ 4.2 GHz | 4.2GHz < f ≦ 6GHz |
| Indoor Anechoic Chamber |  |  |  |
| Compact Antenna Test Range |  |  |  |
| Common maximum accepted test system uncertainty | 0.7 | 0.7 | 0.7 |

An overview of the MU values for all the requirements is captured in clause 16.

### 9.5.6	Test Tolerance for OTA output power dynamics

Considering the methodology described in subclause 5.1, Test Tolerance values for OTA output power dynamics were derived based on values captured in subclause 9.5.5.

The TT was decided to be the same as the MU in FR1

Table 9.5.6-1: Test Tolerance values for the OTA output power dynamic range, FR1

|  | f ≦ 3GHz | 3GHz < f ≦ 4.2 GHz | 4.2GHz < f ≦ 6GHz |
| --- | --- | --- | --- |
| Test Tolerance (dB) | 0.4 | 0.4 | 0.4 |

Table 9.5.6-2: Test Tolerance values for the OTA output power dynamic range, FR1

|  | 24.25<f <29.5GHz | 37<f <40GHz |
| --- | --- | --- |
| Test Tolerance (dB) | 0.4 | 0.4 |

Table 9.5.6-3: Test Tolerance values for the UTRA power control steps

|  | f ≦ 3GHz | 3GHz < f ≦ 4.2 GHz | 4.2GHz < f ≦ 6GHz |
| --- | --- | --- | --- |
| Test Tolerance (dB) | 0.1 | 0.1 | 0.1 |

Table 9.5.6-4: Test Tolerance values for the UTRA power control dynamic range

|  | f ≦ 3GHz | 3GHz < f ≦ 4.2 GHz | 4.2GHz < f ≦ 6GHz |
| --- | --- | --- | --- |
| Test Tolerance (dB) | 1.1 | 1.1 | 1.1 |

Table 9.5.6-5: Test Tolerance values for the UTRA total power dynamic range

|  | f ≦ 3GHz | 3GHz < f ≦ 4.2 GHz | 4.2GHz < f ≦ 6GHz |
| --- | --- | --- | --- |
| Test Tolerance (dB) | 0.3 | 0.3 | 0.3 |

Table 9.5.6-6: Test Tolerance values for the UTRA IPDL time mask

|  | f ≦ 3GHz | 3GHz < f ≦ 4.2 GHz | 4.2GHz < f ≦ 6GHz |
| --- | --- | --- | --- |
| Test Tolerance (dB) | 0.7 | 0.7 | 0.7 |

An overview of the TT values for all the requirements is captured in clause 17.

--- Next change ---

#### 9.6.2.3	MU value derivation

The uncertainty in the power accuracy of the OTA test chamber will not affect the frequency error uncertainty.

Possible phase variation in the chamber due to variation in the signal BW is not significant to affect the frequency error measurement uncertainty.

The conducted MU value is adopted for the OTA MU.

--- Next change ---

#### 9.6.3.3	MU value derivation

The uncertainty in the power accuracy of the OTA test chamber will not affect the frequency error uncertainty.

Possible phase variation in the chamber due to variation in the signal BW is not significant to affect the frequency error measurement uncertainty.

The conducted MU value is adopted for the OTA MU.

--- Next change ---

#### 9.6.4.3	MU value derivation

The uncertainty in the power accuracy of the OTA test chamber will not affect the frequency error uncertainty.

Possible phase variation in the chamber due to variation in the signal BW is not significant to affect the frequency error measurement uncertainty.

The conducted MU value is adopted for the OTA MU.

--- Next change ---

### 9.6.5	Maximum accepted test system uncertainty

The uncertainty in the power accuracy of the OTA test chamber will not affect the frequency error uncertainty.

Possible phase variation in the chamber due to variation in the signal BW is not significant to affect the frequency error measurement uncertainty. The frequency error MU is ±12 Hz.

The MU for FR1 for frequency error for up to 4.2 GHz was agreed to be the same as for eAAS. It is expected that the measurement chamber setup, calibration and measurement procedures and the MU budget will be identical for E-UTRA and NR.

Furthermore, for the frequency range 4.2 – 6 GHz, the chamber and instrument uncertainties are the same as those for the frequency range 3 – 4.2 GHz, assuming testing of a BS designed for operation in licensed spectrum.

For FR2, similarly to FR1, as long as the link budget is sufficient to provide a signal at the test equipment that is within its operating range, then the signal quality requirements MU depends only on the test equipment MU. In the beam center, for a CATR, there is sufficient link budget and thus the MU was decided based on expected test equipment performance.

The MU values are ± 12 Hz for frequency error.

An overview of the MU values for all the requirements is captured in clause 16.

### 9.6.6	Test Tolerance for frequency error

Considering the methodology described in subclause 5.1, Test Tolerance values for frequency error were derived based on values captured in subclause 9.6.5.

The TT values were agreed to be the same as the MU values.

Frequency range specific Test Tolerance values for the frequency error test are defined in table 9.7.6-1.

Table 9.6.6-1: Test Tolerance values for frequency error, FR1

|  | f ≦ 3GHz | 3GHz < f ≦ 4.2 GHz | 4.2GHz < f ≦ 6GHz |
| --- | --- | --- | --- |
| Test Tolerance (Hz) | 12 | 12 | 12 |

Table 9.6.6-2: Test Tolerance values for frequency error, FR2

|  | 24.25<f <29.5GHz | 37<f <40GHz |
| --- | --- | --- |
| Test Tolerance (Hz) | 12 | 12 |

An overview of the TT values for all the requirements is captured in clause 17.

#### 9.7.2.3	MU value derivation

As both the wanted signal and the noise signal are at the same frequency they will be measured at the same time the requirement is effectively differential and most of the OTA chamber errors will cancel out.

The wanted signal will be beam formed and hence the errors used for the EIRP accuracy will be valid, however the co-channel noise may not be beam formed and hence could suffer different errors due to the chamber quite zone, and phase profile. These items are included in both the calibration error and the measurement error, as the requirement is differential if there is a difference between the wanted and the unwanted it will only be due to the measurement phase. The calibration errors will cancel as calibration is only done one so they will be the same for both wanted and unwanted signals.

Potentially, the EVM may vary in space due to different patterns of wanted signal and distortion. Thus for narrow beams, it may be possible that beam pointing and alignment errors could impact EVM results.

As EVM is also dependent on the phase of the calibrated path it is possible that phase ripple in the quite zone or elsewhere, which arises due to multipath reflections, may lead to frequency ripple and cause additional EVM errors which do not appear in a power accuracy analysis as done for EIRP accuracy.

The potential impacts of both beam pointing misalignment and scattering within the chamber on the received waveform and measurement accuracy were investigated. The potential deviation in the measured EVM arising from beam pointing errors was examined considering a worst case scenario, in which variation in space of EVM is maximal due to the ideal signal being correlated and the distortion uncorrelated; hence the impact of misalignment error would be the difference between array gain and element gain. Even in this circumstance, alignment errors of several degrees would not lead to a significant error in the measured EVM. Considering all likely chamber sizes, for E-UTRA any scattering would fall within the cyclic prefix of the OFDM symbol and hence not cause ISI. Furthermore, the likely delay spread of any scattering would relate to coherence bandwidths much larger than any UTRA/E-UTRA channel bandwidth. Even if the scattered energy would cause interference, the interference level would anyhow not lead to a significant EVM increase. Thus it was concluded that the impact scattering within the measurement chamber would be negligible.

The uncertainty causing by power variations when measuring OTA EVM is indicated in table 6.7.2.4-1:

Table 9.7.2.3-1: IAC MU value derivation for power uncertainty aspects of OTA EVM, FR1

| Indoor anechoic |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB] |  |  |
|  |  | f<3 GHz | 3<f<4.2 GHz | 4.2<f<6 GHz |  |  |  | f<3 GHz | 3<f<4.2 GHz | 4.2<f<6 GHz |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |  |  |
| A1-3 | Quality of quiet zone | 0.1 | 0.1 | 0.1 | Gaussian | 1 | 1 | 0.10 | 0.10 | 0.10 |
| A1-6 | Phase curvature | 0.05 | 0.05 | 0.05 | Gaussian | 1 | 1 | 0.05 | 0.05 | 0.05 |
| A1-8 | Random uncertainty | 0.1 | 0.1 | 0.1 | Rectangular | 1.73 | 1 | 0.06 | 0.06 | 0.06 |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |  |  |  |
| Combined standard uncertainty (1σ) [dB] |  |  |  |  |  |  |  | 0.13 | 0.13 | 0.13 |
| Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  |  |  | 0.25 | 0.25 | 0.25 |

The indoor anechoic chamber budget is carried out without consideration of the measurement equipment as this MU is given in %, converting to dB gives, for example:

2% is equivalent to 20*log10(2/100) = -33.98 dB

If the unwanted signal is 0.25 dB higher than the wanted due to the test system then this will be degraded to -33.73 dB, and

-33.73dB is equivalent to:	10(-33.73/20) *100 = 2.06%

Additional error due to potential phase error has not been considered however the potential increase due to then OTA test equipment is well within the contribution allowable with a 1% linear MU.

NOTE:	Analysis of the phase uncertainties indicates that the contributions are not significant to affect the final MU value, however if future work indicates that phase or any other errors not related to amplitude calibration may affect the EVM measurement uncertainty the MU analysis may be re-examined.

--- Next change ---

#### 9.7.3.3 	MU value derivation

As both the wanted signal and the noise signal are at the same frequency they will be measured at the same time the requirement is effectively differential and most of the OTA chamber errors will cancel out.

The wanted signal will be beam formed and hence the errors used for the EIRP accuracy will be valid, however the co-channel noise may not be beam formed and hence could suffer different errors due to the chamber quite zone, and phase profile. These items are included in both the calibration error and the measurement error, as the requirement is differential if there is a difference between the wanted and the unwanted it will only be due to the measurement phase. The calibration errors will cancel as calibration is only done one so they will be the same for both wanted and unwanted signals.

Potentially, the EVM may vary in space due to different patterns of wanted signal and distortion. Thus for narrow beams, it may be possible that beam pointing and alignment errors could impact EVM results.

As EVM is also dependent on the phase of the calibrated path it is possible that phase ripple in the quite zone or elsewhere, which arises due to multipath reflections, may lead to frequency ripple and cause additional EVM errors which do not appear in a power accuracy analysis as done for EIRP accuracy.

The potential impacts of both beam pointing misalignment and scattering within the chamber on the received waveform and measurement accuracy were investigated. The potential deviation in the measured EVM arising from beam pointing errors was examined considering a worst case scenario, in which variation in space of EVM is maximal due to the ideal signal being correlated and the distortion uncorrelated; hence the impact of misalignment error would be the difference between array gain and element gain. Even in this circumstance, alignment errors of several degrees would not lead to a significant error in the measured EVM. Considering all likely chamber sizes, for E-UTRA any scattering would fall within the cyclic prefix of the OFDM symbol and hence not cause ISI. Furthermore, the likely delay spread of any scattering would relate to coherence bandwidths much larger than any UTRA/E-UTRA channel bandwidth. Even if the scattered energy would cause interference, the interference level would anyhow not lead to a significant EVM increase. Thus it was concluded that the impact scattering within the measurement chamber would be negligible.

The uncertainty causing by power variations when measuring OTA EVM is indicated in table 6.7.3.4-1:

Table 9.7.3.3-1: CATR MU value derivation for power uncertainty aspects of OTA EVM, FR1

| CATR |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB] |  |  |
|  |  | f<3 GHz | 3<f<4.2 GHz | 4.2<f<6 GHz |  |  |  | f<3 GHz | 3<f<4.2 GHz | 4.2<f<6 GHz |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |  |  |
| A2-2a | Standing wave between DUT and test range antenna | 0.21 | 0.21 | 0.21 | U-shaped | 1.41 | 1 | 0.15 | 0.15 | 0.15 |
| A2-4a | QZ ripple DUT | 0.09 | 0.09 | 0.09 | Normal | 1 | 1 | 0.09 | 0.09 | 0.09 |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |  |  |  |
| Combined standard uncertainty (1σ) [dB] |  |  |  |  |  |  |  | 0.18 | 0.18 | 0.18 |
| Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  |  |  | 0.34 | 0.34 | 0.34 |

The CATR budget is carried out without consideration of the measurement equipment as this MU is given in %, converting to dB gives, for example:

2% is equivalent to 20*log10(2/100) = -33.98 dB

If the unwanted signal is 0.35 dB higher than the wanted due to the test system then this will be degraded to -33.63 dB, and

-33.63 dB is equivalent to:	10(-33.63/20) *100 = 2.08%

Additional error due to potential phase error has not been considered however the potential increase due to then OTA test equipment is well within the contribution allowable with a 1% linear MU.

NOTE:	Analysis of the phase uncertainties indicates that the contributions are not significant to affect the final MU value, however if future work indicates that phase or any other errors not related to amplitude calibration may affect the EVM measurement uncertainty the MU analysis may be re-examined.

--- Next change ---

#### 9.7.4.3	MU value derivation

EVM is a relative measurement given that the wanted signal and noise signal are at the same frequency and measured at the same time therefore most of the OTA anechoic chamber errors will cancel out.

Nearly all of uncertainty terms for the EVM measurement and EVM calibration are the same and hence EVM is a differential or relative measurement.

Potentially, the EVM may vary in space due to different patterns of wanted signal and distortion. Thus for narrow beams, it may be possible that beam pointing and alignment errors could impact EVM results. The importance and impact of such effects is likely to be even smaller than for far field based measurements.

As EVM is also dependent on the phase of the calibrated path it is possible that phase ripple in the quiet zone or elsewhere, which arises due to multipath reflections, may lead to frequency ripple and cause additional EVM errors which do not appear in a power accuracy analysis as done for EIRP accuracy. Analysis of the effect of such effects is that the effects will be even smaller than for far field based measurements.

The Near Field budget is carried out without consideration of the measurement equipment therefore the MU is given in % and can be converted to dB, for example:

2% is equivalent to 20*log10(2/100) = -33.98 dB

If the unwanted signal is 0.35 dB higher than the wanted due to the test system then this will be degraded to -33.63 dB, and

-33.63 dB is equivalent to:	10(-33.63/20) *100 = 2.08%

Additional error due to potential phase error has not been considered however the potential increase due to then OTA test equipment is well within the contribution allowable with a 1% linear MU.

For Near Field Test Range a new measurement uncertainty term shall be added to the MU. This term will take into account the fact that in Near Field the phase pattern will sum up so that the signal level is increasing while the noise level is the same. This MU term will only cause an error in the direction of increasing the reported EVM value and not decreasing it, and will depend on the implementation of the BS.

NOTE:	Analysis of the phase uncertainties indicates that the contributions are not significant enough to affect the final MU value, however if future work indicates that phase or any other errors not related to amplitude calibration may affect the EVM measurement uncertainty, the MU analysis may need to be re-examined.

--- Next change ---

### 9.7.5	Maximum accepted test system uncertainty

Without consideration of any phase uncertainty, the amplitude error analysis shows the conducted MU of 1% can be maintained for the OTA MU (subject to the clarification noted in the limitations clause that the reported EVM may be greater than the real EVM due to the difference between near field and far field EVM values. The extent of such a difference is dependent on the architecture of the BS).

NOTE :	Analysis of the phase uncertainties indicates that the contributions are not significant to affect the final MU value, however if future work indicates that phase or any other errors not related to amplitude calibration may affect the EVM measurement uncertainty the MU analysis may be re-examined.

The MU for FR1 for EVM for 0 - 4.2 GHz was agreed to be the same as for eAAS. It is expected that the measurement chamber setup, calibration and measurement procedures and the MU budget will be identical for E-UTRA and NR.

Furthermore, for the frequency range 4.2 – 6 GHz, the chamber and instrument uncertainties are the same as those for the frequency range 3 – 4.2 GHz, assuming testing of a BS designed for operation in licensed spectrum.

For FR2, similarly to FR1, as long as the link budget is sufficient to provide a signal at the test equipment that is within its operating range, then the signal quality requirements MU depends only on the test equipment MU. In the beam center, for a CATR, there is sufficient link budget and thus the MU was decided based on expected test equipment performance.

The MU values are 1% for EVM.

An overview of the MU values for all the requirements is captured in clause 16.

### 9.7.6	Test Tolerance for EVM

Considering the methodology described in subclause 5.1, Test Tolerance values for EVM were derived based on values captured in subclause 9.7.5.

The TT values were agreed to be the same as the MU values.

Frequency range specific Test Tolerance values for the EVM test are defined in table 9.7.6-1.

Table 9.7.6-1: Test Tolerance values for EVM, FR1

|  | f ≦ 3GHz | 3GHz < f ≦ 4.2 GHz | 4.2GHz < f ≦ 6GHz |
| --- | --- | --- | --- |
| Test Tolerance (%) | 1 | 1 | 1 |

Table 9.7.6-2: Test Tolerance values for EVM, FR2

|  | 24.25<f <29.5GHz | 37<f <40GHz |
| --- | --- | --- |
| Test Tolerance (%) | 1 | 1 |

An overview of the TT values for all the requirements is captured in clause 17.

--- Next change ---

#### 9.8.2.3	MU value derivation

The uncertainty in the power accuracy of the OTA test chamber will not affect the TAE measurement uncertainty.

Possible phase variation in the chamber due to variation in the signal BW is not significant to affect the TAE measurement uncertainty.

The conducted MU value is adopted for the OTA MU.

--- Next change ---

#### 9.8.3.3	MU value derivation

The uncertainty in the power accuracy of the OTA test chamber will not affect the TAE measurement uncertainty.

Possible phase variation in the chamber due to variation in the signal BW is not significant to affect the TAE measurement uncertainty.

The conducted MU value is adopted for the OTA MU.

--- Next change ---

#### 9.8.4.3	MU value derivation

Refer to subclause 6.8.3.3 for the OTA TAE measurement in CATR.

--- Next change ---

### 9.8.5	Maximum accepted test system uncertainty

The uncertainty in the power accuracy of the OTA test chamber will not affect the TAE measurement uncertainty.

Possible phase variation in the chamber due to variation in the signal BW is not significant to affect the TAE measurement uncertainty.

The conducted MU value is adopted for the OTA MU, and is ±25 ns.

The MU for FR1 for TAE for 0 - 4.2 GHz was agreed to be the same as for eAAS. It is expected that the measurement chamber setup, calibration and measurement procedures and the MU budget will be identical for E-UTRA and NR.

Furthermore, for the frequency range 4.2 – 6 GHz, the chamber and instrument uncertainties are the same as those for the frequency range 3 – 4.2 GHz, assuming testing of a BS designed for operation in licensed spectrum.

For FR2, similarly to FR1, as long as the link budget is sufficient to provide a signal at the test equipment that is within its operating range, then the signal quality requirements MU depends only on the test equipment MU. In the beam center, for a CATR, there is sufficient link budget and thus the MU was decided based on expected test equipment performance.

The MU values are 25 ns for time alignment error.

An overview of the MU values for all the requirements is captured in clause 16.

### 9.7.6	Test Tolerance for TAE

Considering the methodology described in subclause 5.1, Test Tolerance values for TAE were derived based on values captured in subclause 9.7.5.

The TT values were agreed to be the same as the MU values.

Frequency range specific Test Tolerance values for the TAE test are defined in table 9.7.6-1 and 9.7.6-2..

Table 9.7.6-1: Test Tolerance values for TAE, FR1

|  | f ≦ 3GHz | 3GHz < f ≦ 4.2 GHz | 4.2GHz < f ≦ 6GHz |
| --- | --- | --- | --- |
| Test Tolerance (ns) | 25 | 25 | 25 |

Table 9.7.6-2: Test Tolerance values for TAE, FR2

|  | 24.25<f <29.5GHz | 37<f <40GHz |
| --- | --- | --- |
| Test Tolerance (ns) | 25 | 25 |

An overview of the TT values for all the requirements is captured in clause 17.

--- Next change ---

#### 9.9.2.3	MU value derivation

#### The MU value is the same as the conducted value given in table 9.1-1.

--- Next change ---

#### 9.9.3.3	MU value derivation

#### The MU value is the same as the conducted value given in table 9.1-1.

--- Next change ---

#### 9.9.4.3	MU value derivation

#### The MU value is the same as the conducted value given in table 9.1-1.

--- Next change ---

### 9.9.5	Maximum accepted test system uncertainty

For NR in FR1, the MU for the OTA occupied bandwidth depends on the roll-off of the transmitted signal and the instrument MU, not on the OTA chamber related MU. In principle, the occupied bandwidth MU was agreed to be the same as for E-UTRA/eAAS [9]. However NR covers larger BS channel bandwidths than E-UTRA, and thus additional MU relating to larger channel bandwidths were estimated. The MU was decided based on the density of power measurements within the channel bandwidth and the expected test instrument performance.

For NR in FR2, similarly to FR1, as long as the link budget is sufficient to provide a signal at the test equipment that is within its operating range, then the occupied bandwidth MU depends only on the test equipment MU. In the beam center, for a CATR, there is sufficient link budget and thus the MU was decided based on expected test equipment performance, considering the agreed density of power measurements in the frequency domain.

For E-UTRA, the OTA occupied BW MU is the same as the conducted MU and is as follows:

1.4 MHz, 3 MHz Channel BW: 30 kHz

5 MHz, 10 MHz Channel BW: 100 kHz

15 MHz, ≥20 MHz: Channel BW: 300 kHz

For NR in FR1, the MU was agreed to be:

+/- 100 kHz for BS channel bandwidths up to 10 MHz,

+/- 300 kHz for 10 MHz < BS channel bandwidth ≤ 50 MHz, and

+/- 300 kHz for 50 MHz < BS channel bandwidths ≤ 100 MHz.

For NR in FR2, based on expected test equipment performance, the MU for occupied bandwidth was decided to be:

+/- 600 kHz (same as for FR1 for channel bandwidths greater than 50 MHz).

An overview of the MU values for all the requirements is captured in clause 16.

### 9.9.6	Test Tolerance for OTA occupied bandwidth

Considering the methodology described in subclause 5.1, Test Tolerance values for OTA occupied bandwidth were derived based on values captured in subclause 8.9.5.

For NR in FR1 as well as in FR2, the TT value was agreed to be 0 Hz.

Frequency range specific Test Tolerance values for the OTA occupied bandwidth test are defined in table 9.6.6-1 and table 9.6.6-2.

Table 9.9.6-1: Test Tolerance values for the EIRP accuracy in Normal test conditions, FR1

|  | f ≦ 3GHz | 3GHz < f ≦ 4.2 GHz | 4.2GHz < f ≦ 6GHz |
| --- | --- | --- | --- |
| Test Tolerance (Hz) | 0 | 0 | 0 |

Table 9.9.6-2: Test Tolerance values for the EIRP accuracy in Normal test conditions, FR2

|  | 24.25<f <29.5GHz | 37<f <40GHz |
| --- | --- | --- |
| Test Tolerance (Hz) | 0 | 0 |

An overview of the TT values for all the requirements is captured in clause 17.

--- Next change ---

#### 9.10.2.3	MU value derivation, FR2

The MU assessment was carried out using a CATR chamber only. However other chamber types are not precluded if suitable MU assessment is done.

The CATR test setup and calibration and measurement procedures for FR2 are expected to be similar to those of FR1, although the test chamber dimensions and associated MU values will scale due to the shorter wavelengths and larger relative array apertures. However, it is noted that in order to achieve the test instrument uncertainties that were assumed, calibration of the spectrum analyzer may be needed.

Table 9.10.2.3-1: Compact antenna test range uncertainty assessment for EIRP measurements for transmitter OFF power and transmitter transient period

--- Next change ---

### 9.10.3	Maximum accepted test system uncertainty

The MU assessment was carried out using a CATR chamber only. However other chamber types are not precluded if suitable MU assessment is done.

Based on the above evaluation and root sum square combining of the dB values for the MU, the MU was decided to be 2.9 dB for the frequency range 24.25 < f < 29.5 GHz and 3.3 dB for the frequency range 37 < f < 40 GHz.

--- Next change ---

### 9.10.4	Test Tolerance for OTA TX OFF power and transmitter transient period

Considering the methodology described in subclause 5.1, Test Tolerance values for OTA TX OFF power and transmitter transient period were derived based on values captured in subclause 9.10.3.

The TT value was agreed to be the same as the MU.

Frequency range specific Test Tolerance values for the OTA TX OFF power and transmitter transient period test are defined in table 9.10.4-1.

Table 9.10.4-1: Test Tolerance values for the EIRP accuracy, FR2

|  | 24.25 < f < 29.5 GHz | 37 < f < 40 GHz |
| --- | --- | --- |
| Test Tolerance (dB) | 2.9 | 3.3 |

An overview of the TT values for all the requirements is captured in clause 17.

# 11	In-band TRP requirements

--- Next change ---

#### 11.2.2.3	MU value derivation, FR1

Table 11.2.2.3-1 captures derivation of the expanded measurement uncertainty values for OTA BS output power measurements in Indoor Anechoic Chamber (Normal test conditions, FR1).

Table 11.2.2.3-1: IAC MU value derivation for EIRP measurement of OTA BS output power

| Indoor anechoic |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB] |  |  |  |  |  |  |  |
|  |  | f<3 GHz | 3<f<4.2 GHz | 4.2<f<6 GHz |  |  |  | f<3 GHz |  | 3<f<4.2 GHz |  |  | 4.2<f<6 GHz |  |  |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| A1-1 | Positioning misalignment between the AAS BS and the reference antenna | 0.03 | 0.03 | 0.03 | Rectangular | 1.73 | 1 | 0.02 |  | 0.02 |  |  | 0.02 |  |  |
| A1-2 | Pointing misalignment between the AAS BS and the receiving antenna | 0.3 | 0.3 | 0.3 | Rectangular | 1.73 | 1 | 0.17 |  | 0.17 |  |  | 0.17 |  |  |
| A1-3 | Quality of quiet zone | 0.1 | 0.1 | 0.1 | Gaussian | 1.00 | 1 | 0.10 |  | 0.10 |  |  | 0.10 |  |  |
| A1-4a | Polarization mismatch between the AAS BS and the receiving antenna | 0.01 | 0.01 | 0.01 | Rectangular | 1.73 | 1 | 0.01 |  | 0.01 |  |  | 0.01 |  |  |
| A1-5 | Mutual coupling between the AAS BS and the receiving antenna | 0 | 0 | 0 | Rectangular | 1.73 | 1 | 0.00 |  | 0.00 |  |  | 0.00 |  |  |
| A1-6 | Phase curvature | 0.05 | 0.05 | 0.05 | Gaussian | 1.00 | 1 | 0.05 |  | 0.05 |  |  | 0.05 |  |  |
| C1-1 | Uncertainty of the RF Power Measurement Equipment | 0.14 | 0.26 | 0.26 | Gaussian | 1.00 | 1 | 0.14 |  | 0.26 |  |  | 0.26 |  |  |
| A1-7 | Impedance mismatch in the receiving chain | 0.14 | 0.33 | 0.33 | U-shaped | 1.41 | 1 | 0.10 |  | 0.23 |  |  | 0.23 |  |  |
| A1-8 | Random uncertainty | 0.1 | 0.1 | 0.1 | Rectangular | 1.73 | 1 | 0.06 |  | 0.06 |  |  | 0.06 |  |  |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| A1-9 | Impedance mismatch between the receiving antenna and the network analyzer | 0.05 | 0.05 | 0.05 | U-shaped | 1.41 | 1 | 0.04 |  | 0.04 |  |  | 0.04 |  |  |
| A1-10 | Positioning and pointing misalignment between the reference antenna and the receiving antenna | 0.01 | 0.01 | 0.01 | Rectangular | 1.73 | 1 | 0.01 |  | 0.01 |  |  | 0.01 |  |  |
| A1-11 | Impedance mismatch between the reference antenna and the network analyzer. | 0.05 | 0.05 | 0.05 | U-shaped | 1.41 | 1 | 0.04 |  | 0.04 |  |  | 0.04 |  |  |
| A1-3 | Quality of quiet zone | 0.10 | 0.10 | 0.10 | Gaussian | 1.00 | 1 | 0.10 |  | 0.10 |  |  | 0.10 |  |  |
| A1-4b | Polarization mismatch for reference antenna | 0.01 | 0.01 | 0.01 | Rectangular | 1.73 | 1 | 0.01 |  | 0.01 |  |  | 0.01 |  |  |
| A1-5 | Mutual coupling between the reference antenna and the receiving antenna | 0.00 | 0.00 | 0.00 | Rectangular | 1.73 | 1 | 0.00 |  | 0.00 |  |  | 0.00 |  |  |
| A1-6 | Phase curvature | 0.05 | 0.05 | 0.05 | Gaussian | 1.00 | 1 | 0.05 |  | 0.05 |  |  | 0.05 |  |  |
| C1-3 | Uncertainty of the network analyzer | 0.13 | 0.20 | 0.20 | Gaussian | 1.00 | 1 | 0.13 |  | 0.20 |  |  | 0.20 |  |  |
| A1-12 | Influence of the reference antenna feed cable | 0.05 | 0.05 | 0.05 | Rectangular | 1.73 | 1 | 0.03 |  | 0.03 |  |  | 0.03 |  |  |
| A1-13 | Reference antenna feed cable loss measurement uncertainty | 0.06 | 0.06 | 0.06 | Gaussian | 1.00 | 1 | 0.06 |  | 0.06 |  |  | 0.06 |  |  |
| A1-14 | Influence of the receiving antenna feed cable | 0.05 | 0.05 | 0.05 | Rectangular | 1.73 | 1 | 0.03 |  | 0.03 |  |  | 0.03 |  |  |
| C1-4 | Uncertainty of the absolute gain of the reference antenna | 0.50 | 0.43 | 0.43 | Rectangular | 1.73 | 1 | 0.29 |  | 0.25 |  |  | 0.25 |  |  |
| A1-15 | Uncertainty of the absolute gain of the receiving antenna | 0.00 | 0.00 | 0.00 | Rectangular | 1.73 | 1 | 0.00 |  | 0.00 |  |  | 0.00 |  |  |
| Combined standard uncertainty (1σ) [dB] |  |  |  |  |  |  |  |  | 0.44 |  |  | 0.54 |  |  | 0.54 |  |  |
| Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  |  |  |  | 0.87 |  |  | 1.06 |  |  | 1.06 |  |  |
| TRP Summation error |  |  |  |  |  |  |  |  | 0.75 |  |  | 0.75 |  |  | 0.75 |  |  |
| Total MU |  |  |  |  |  |  |  |  | 1.15 |  |  | 1.30 |  |  | 1.30 |  |  |

--- Next change ---

#### 11.2.3.3	MU value derivation, FR1

Table 11.2.3.3-1 captures derivation of the expanded measurement uncertainty values for OTA BS output power measurements in CATR (Normal test conditions, FR1).

Table 11.2.3.3-1: CATR MU value derivation for OTA BS output power measurement, FR1

--- Next change ---

#### 11.2.3.4	MU value derivation, FR2

A CATR MU budget was assessed in order to determine acceptable MU for the EIRP accuracy measurement in FR2. The CATR test setup and calibration and measurement procedures for FR2 are expected to be similar to those of FR1, although the test chamber dimensions and associated MU values will scale due to the shorter wavelengths and larger relative array apertures.

Table 11.2.3.4-1 captures derivation of the expanded measurement uncertainty values for OTA BS output power measurements in CATR (Normal test conditions, FR2).

Table 11.2.3.4-1: CATR MU value derivation for OTA BS output power measurement, FR2

--- Next change ---

#### 11.2.4.3	MU value derivation, FR1

Table 11.2.4.3-1 captures derivation of the expanded measurement uncertainty values for OTA BS output power measurements in NFTR (Normal test conditions, FR1).

Table 11.2.4.3-1: NFTR MU value derivation for power density pattern measurement

MUper-point2+SE2

f≤3GHzMUTRP=0.592+0.752

3GHz≤f≤4.2GHzMUTRP=0.702+0.752

--- Next change ---

#### 11.2.5.3	MU value derivation, FR1

Table 11.2.5.3-1 captures derivation of the expanded measurement uncertainty values for OTA BS output power measurements in Reverberation Chamber (Normal test conditions, FR1).

Table 11.2.5.3-1: Reverberation chamber MU value derivation for OTA BS output power

--- Next change ---

#### 11.2.5.4	MU value derivation, FR2

Table 11.2.5.4-1 captures derivation of the expanded measurement uncertainty values for OTA BS output power measurements in Reverberation Chamber (Normal test conditions, FR2).

Table 11.2.5.4-1: Reverberation chamber MU value derivation for OTA BS output power, FR2

--- Next change ---

#### 11.2.6.3	MU value derivation, FR1

Table 11.2.6.3-1 captures derivation of the expanded measurement uncertainty values for OTA BS output power measurements in PWS.

Table 11.2.6.3-1: PWS MU value derivation for OTA BS output power

--- Next change ---

### 11.2.7	Maximum accepted test system uncertainty

For the frequency range up to 4.2 GHz, the same MU values as for E-UTRA were adopted [9]. It is expected that the test chamber setup, calibration and measurement procedures for E-UTRA and NR will be highly similar. All uncertainty factors were judged to be the same.

For the frequency range 4.2 - 6 GHz, all MU factors, including instrumentation related MU were judged to be the same as for the 3 - 4.2 GHz range, and thus the total MU for 4.2 – 6 GHz is the same as for 3 - 4.2 GHz. This assessment was made under the assumption of testing BS designed for licensed spectrum; for unlicensed spectrum the MU may differ.

For CATR the expanded MU is established as a root sum square combining of the dB values for the MU and the SE (see clause 12.10), the MU was decided to be 2.1 dB for the frequency range 24.25<f<29.5GHz and 2.4 dB for the frequency range 37<f<40GHz.

Maximum test system uncertainties derivation methodology was described in subclause 5.1. The maximum accepted test system uncertainty values was derived based on test system specific values.

According to the methodology referred above, the common maximum accepted test system uncertainty values for OTA BS output power test can be derived from values captured in table 11.2.7-1, derived based on the expanded uncertainty ue (1.96σ - confidence interval of 95 %) values. The common maximum accepted test system uncertainty values are applicable for all test methods addressing OTA BS output power test requirement.

Table 11.2.7-1: Test system specific MU values for the OTA BS output power test, Normal test conditions, FR1

|  | Expanded uncertainty ue (dB) |  |  |
| --- | --- | --- | --- |
|  | f ≦ 3GHz | 3GHz < f ≦ 4.2 GHz | 4.2GHz < f ≦ 6GHz |
| Indoor Anechoic Chamber |  |  |  |
| Compact Antenna Test Range |  |  |  |
| Near Field chamber |  |  |  |
| Reverberation chamber |  |  |  |
| Plane Wave Synthesizer |  |  |  |
| Common maximum accepted test system uncertainty | 1.4 | 1.5 | 1.5 |

Table 11.2.7-2: Test system specific MU values for the OTA BS output power test, Normal test conditions, FR2

|  | Expanded uncertainty ue (dB) |  |
| --- | --- | --- |
|  | 24.25 < f < 29.5 GHz | 37 < f < 40 GHz |
| Indoor Anechoic Chamber |  |  |
| Compact Antenna Test Range |  |  |
| Near Field chamber |  |  |
| Reverberation chamber |  |  |
| Plane Wave Synthesizer |  |  |
| Common maximum accepted test system uncertainty | 2.1 | 2.4 |

An overview of the MU values for all the requirements is captured in clause 16.

### 11.2.8	Test Tolerance for OTA BS output power

Considering the methodology described in subclause 5.1, Test Tolerance values for OTA BS output power were derived based on values captured in subclause 11.2.7.

The TT value was agreed to be the same as the MU value, both for FR1 and FR2.

Frequency range specific Test Tolerance values for the OTA BS output power test are defined in table 11.2.8-1.

Table 11.2.8-1: Test Tolerance values for the OTA BS output power, Normal test conditions, FR1

|  | f ≦ 3GHz | 3GHz < f ≦ 4.2 GHz | 4.2GHz < f ≦ 6GHz |
| --- | --- | --- | --- |
| Test Tolerance (dB) | 1.4 | 1.5 | 1.5 |

Table 11.2.8-2: Test Tolerance values for the OTA BS output power, Normal test conditions, FR2

|  | 24.25 < f < 29.5 GHz | 37 < f < 40 GHz |
| --- | --- | --- |
| Test Tolerance (dB) | 1.4 | 1.5 |

An overview of the TT values for all the requirements is captured in clause 17.

--- Next change ---

#### 11.3.2.3	MU value derivation, FR1

Table 11.3.2.3-1: IAC MU value derivation for the absolute ACLR measurement

--- Next change ---

#### 11.3.3.3	MU value derivation, FR1

Table 11.3.3.3-1: CATR MU value derivation for the EIRP measurement of the absolute OTA ACLR, FR1

Table 11.3.3.3-2: CATR MU value derivation for the EIRP measurement of the relative OTA ACLR, FR1

--- Next change ---

#### 11.3.3.4	MU value derivation, FR2

A CATR MU budget was assessed in order to determine acceptable MU for the EIRP accuracy measurement in FR2. The CATR test setup and calibration and measurement procedures for FR2 are expected to be similar to those of FR1, although the test chamber dimensions and associated MU values will scale due to the shorter wavelengths and larger relative array apertures. However, it is noted that in order to achieve the test instrument uncertainties that were assumed, calibration of the spectrum analyzer may be needed.

For relative ACLR, the MU budget for CATR was assessed as follows:

Table 11.3.3.4-1: CATR MU value derivation for the EIRP measurement of the absolute OTA ACLR, FR2

Table 11.3.3.4-2: CATR MU value derivation for the EIRP measurement of the relative OTA ACLR, FR2

--- Next change ---

#### 11.3.4.3	MU value derivation, FR1

Table 11.3.4.3-1 captures derivation of the expanded measurement uncertainty values for OTA ACLR measurements in NFTR (Normal test conditions, FR1).

Table 11.3.4.3-1: NFTR MU value derivation for absolute ACLR measurement

--- Next change ---

#### 11.3.5.3	MU value derivation, FR1

Table 11.3.5.3-1 captures derivation of the expanded measurement uncertainty values for relative ACLR measurements in Reverberation Chamber (Normal test conditions, FR1).

Table 11.3.5.3-2 captures derivation of the expanded measurement uncertainty values for absolute ACLR measurements in Reverberation Chamber (Normal test conditions, FR1).

Table 11.3.5.3-1: Reverberation chamber MU value derivation for relative ACLR measurement

Table 11.3.5.3-2: Reverberation chamber MU value derivation for absolute ACLR measurement

--- Next change ---

#### 11.3.5.4	MU value derivation, FR2

Table 11.3.5.4-1 captures derivation of the expanded measurement uncertainty values for relative ACLR measurements in Reverberation Chamber (Normal test conditions, FR2).

Table 11.3.5.4-2 captures derivation of the expanded measurement uncertainty values for absolute ACLR measurements in Reverberation Chamber (Normal test conditions, FR2).

Table 11.3.5.4-1: Reverberation chamber MU value derivation for relative ACLR measurement, FR2

Table 11.3.5.4-2: Reverberation chamber MU value derivation for absolute ACLR measurement, FR2

--- Next change ---

#### 11.3.6.3	MU value derivation, FR1

The MU value derivation for absolute ACLR measurements is the same as in subclause 11.2.6.3 (i.e. OTA BS output power).

Table 11.3.6.3-1 captures derivation of the expanded measurement uncertainty values for relative ACLR measurements in PWS.

Table 11.3.6.3-1: PWS MU value derivation for relative ACLR measurement

Table 11.3.6.3-2: PWS MU value derivation for absolute ACLR measurement

--- Next change ---

### 11.3.7	Maximum accepted test system uncertainty

For the frequency range up to 4.2 GHz, the same MU values as for E-UTRA were adopted [9]. It is expected that the test chamber setup, calibration and measurement procedures for E-UTRA and NR will be highly similar. All uncertainty factors were judged to be the same.

For the frequency range 4.2 - 6 GHz, all MU factors, including instrumentation related MU were judged to be the same as for the 3 - 4.2 GHz range, and thus the total MU for 4.2 – 6 GHz is the same as for 3 - 4.2 GHz. This assessment was made under the assumption of testing BS designed for licensed spectrum; for unlicensed spectrum the MU may differ.

For CATR the expanded MU is established as a root sum square combining of the dB values for the MU and the SE (see subclause 6.3.6), the MU for absolute ACLR was decided to be 2.7 dB for the frequency range 24.25<f<29.5GHz and 2.7 dB for the frequency range 37<f<40GHz. The MU for relative ACLR was decided to be 2.3 dB for the frequency range 24.25<f<29.5GHz and 2.6 dB for the frequency range 37<f<40GHz.

Maximum test system uncertainties derivation methodology was described in subclause 5.1. The maximum accepted test system uncertainty values was derived based on test system specific values.

According to the methodology referred above, the common maximum accepted test system uncertainty values for absolute and relative OTA ACLR tests can be derived from values captured in table 11.3.7-1 and 11.3.7-2, derived based on the expanded uncertainty ue (1.96σ - confidence interval of 95 %) values. The common maximum accepted test system uncertainty values are applicable for all test methods addressing OTA ACLR test requirement.

Table 11.3.7-1: Test system specific MU values for the absolute OTA ACLR, Normal test conditions, FR1

|  | Expanded uncertainty ue (dB) |  |  |
| --- | --- | --- | --- |
|  |  | f ≦ 3GHz | 3GHz < f ≦ 4.2 GHz | 4.2GHz < f ≦ 6GHz |
| Indoor Anechoic Chamber |  |  |  |  |
| Compact Antenna Test Range |  |  |  |  |
| Near Field chamber |  |  |  |  |
| Reverberation chamber |  |  |  |  |
| Plane Wave Synthesizer |  |  |  |  |
| Common maximum accepted test system uncertainty |  | 2,2 | 2,7 | 2,7 |

Table 11.3.7-2: Test system specific MU values for the relative OTA ACLR, Normal test conditions, FR1

|  | Expanded uncertainty ue (dB) |  |  |
| --- | --- | --- | --- |
|  |  | f ≦ 3GHz | 3GHz < f ≦ 4.2 GHz | 4.2GHz < f ≦ 6GHz |
| Indoor Anechoic Chamber |  |  |  |  |
| Compact Antenna Test Range |  |  |  |  |
| Near Field chamber |  |  |  |  |
| Reverberation chamber |  |  |  |  |
| Plane Wave Synthesizer |  |  |  |  |
| Common maximum accepted test system uncertainty |  | 1,0 | 1,2 | 1,2 |

Table 11.3.7-3: Test system specific MU values for the absolute OTA ACLR, Normal test conditions, FR2

|  | Expanded uncertainty ue (dB) |  |
| --- | --- | --- |
|  | 24.25 < f < 29.5 GHz | 37 < f < 40 GHz |
| Indoor Anechoic Chamber |  |  |
| Compact Antenna Test Range |  |  |
| Near Field chamber |  |  |
| Reverberation chamber |  |  |
| Plane Wave Synthesizer |  |  |
| Common maximum accepted test system uncertainty | 2,7 | 2,7 |

Table 11.3.7-4: Test system specific MU values for the relative OTA ACLR, Normal test conditions, FR2

|  | Expanded uncertainty ue (dB) |  |
| --- | --- | --- |
|  | 24.25 < f < 29.5 GHz | 37 < f < 40 GHz |
| Indoor Anechoic Chamber |  |  |
| Compact Antenna Test Range |  |  |
| Near Field chamber |  |  |
| Reverberation chamber |  |  |
| Plane Wave Synthesizer |  |  |
| Common maximum accepted test system uncertainty | 2.3 | 2.6 |

For relative ACLR, the MU value was agreed to be 1.0 dB for 0 – 3 GHz bands and 1.2 dB for 3 – 6 GHz bands. The MU in 4.2-6 GHz is valid for BS designed to operate in licensed spectrum.

For absolute ACLR, the MU value was agreed to be 2.2 dB for 0 – 3 GHz bands and 2.7 dB for 3 – 6 GHz bands. The MU in 4.2-6 GHz is valid for BS designed to operate in licensed spectrum.

An overview of the MU values for all the requirements is captured in clause 16.

### 11.3.8	Test Tolerance for OTA ACLR

Considering the methodology described in subclause 5.1, Test Tolerance values for OTA ACLR were derived based on values captured in subclause 11.3.7.

For FR1:

- For relative ACLR, the TT was agreed to be the same as the MU.

- For absolute ACLR, the TT was agreed to be 0 dB.

For FR2: The TT was agreed to be the same as the MU.

Frequency range specific Test Tolerance values for the OTA ACLR test are defined in table 11.3.8-1 and 11.3.8-2.

Table 11.3.8-1: Test Tolerance values for the absolute OTA ACLR, Normal test conditions, FR1

|  | f ≦ 3GHz | 3GHz < f ≦ 4.2 GHz | 4.2GHz < f ≦ 6GHz |
| --- | --- | --- | --- |
| Test Tolerance (dB) | 0 | 0 | 0 |

Table 11.3.8-2: Test Tolerance values for the relative OTA ACLR, Normal test conditions, FR1

|  | f ≦ 3GHz | 3GHz < f ≦ 4.2 GHz | 4.2GHz < f ≦ 6GHz |
| --- | --- | --- | --- |
| Test Tolerance (dB) | 1,0 | 1,2 | 1,2 |

Table 11.3.8-3: Test Tolerance values for the absolute OTA ACLR, Normal test conditions, FR2

|  | 24.25 < f < 29.5 GHz | 37 < f < 40 GHz |
| --- | --- | --- |
| Test Tolerance (dB) | 2.7 | 2.7 |

Table 11.3.8-4: Test Tolerance values for the relative OTA ACLR, Normal test conditions, FR2

|  | 24.25 < f < 29.5 GHz | 37 < f < 40 GHz |
| --- | --- | --- |
| Test Tolerance (dB) | 2.3 | 2.6 |

An overview of the TT values for all the requirements is captured in clause 17.

--- Next change ---

#### 11.4.2.3	MU value derivation, FR1

Table 11.4.2.3-1 captures derivation of the expanded measurement uncertainty values for OTA OBUE measurements in Indoor Anechoic Chamber (Normal test conditions, FR1).

Table 11.4.2.3-1: IAC MU value derivation for OTA OBUE measurement, FR1

--- Next change ---

#### 11.4.3.3	MU value derivation, FR1

Table 10.4.3.3-1 captures derivation of the expanded measurement uncertainty values for OTA OBUE measurements in CATR (Normal test conditions, FR1).

Table 11.4.3.3-1: CATR MU value derivation for OTA OBUE measurement, FR1

--- Next change ---

#### 11.4.3.4	MU value derivation, FR2

A CATR MU budget was assessed in order to determine acceptable MU for the EIRP accuracy measurement in FR2. The CATR test setup and calibration and measurement procedures for FR2 are expected to be similar to those of FR1, although the test chamber dimensions and associated MU values will scale due to the shorter wavelengths and larger relative array apertures. However, it is noted that in order to achieve the test instrument uncertainties that were assumed, calibration of the spectrum analyzer may be needed.

Table 10.4.3.4-1 captures derivation of the expanded measurement uncertainty values for OTA OBUE measurements in CATR (Normal test conditions, FR2).

Table 11.4.3.4-1: CATR MU value derivation for OTA OBUE measurement, FR2

--- Next change ---

#### 11.4.4.3	MU value derivation, FR1

Refer to subclause 11.2.4.3 (i.e. the MU value derivation for the OTA BS output power measurement in NFTR) for MU value per point measurement.

MUTRP=MUper-point2+SE2

f≤3GHzMUTRP=1.152+0.752

3GHz≤f≤4.2GHzMUTRP=1.372+0.752

--- Next change ---

#### 11.4.5.3	MU value derivation, FR1

Table 11.4.5.3-1 captures derivation of the expanded measurement uncertainty values for OTA OBUE measurement in Reverberation chamber (Normal test conditions, FR1).

Table 11.4.5.3-1: Reverberation chamber MU value derivation for OTA OBUE

#### 11.4.5.4	MU value derivation, FR2

Table 11.4.5.4-1 captures derivation of the expanded measurement uncertainty values for OTA OBUE measurements in Reverberation chamber (Normal test conditions, FR2).

Table 11.4.5.4-1: Reverberation chamber MU value derivation for OTA OBUE measurement, FR2

--- Next change ---

#### 11.4.6.3	MU value derivation, FR1

The MU value for OTA OBUE in PWS is the same as in subclause 11.2.6.3 (i.e. OTA BS output power MU in PWS).

--- Next change ---

### 11.4.7	Maximum accepted test system uncertainty

For the frequency range up to 4.2 GHz, the same MU values as for E-UTRA were adopted [9]. It is expected that the test chamber setup, calibration and measurement procedures for E-UTRA and NR will be highly similar. All uncertainty factors were judged to be the same.

For the frequency range 4.2 - 6 GHz, all MU factors, including instrumentation related MU were judged to be the same as for the 3 - 4.2 GHz range, and thus the total MU for 4.2 – 6 GHz is the same as for 3 - 4.2 GHz. This assessment was made under the assumption of testing BS designed for licensed spectrum; for unlicensed spectrum the MU may differ.

Maximum test system uncertainties derivation methodology was described in subclause 5.1. The maximum accepted test system uncertainty values was derived based on test system specific values.

According to the methodology referred above, the common maximum accepted test system uncertainty values for OTA OBUE test can be derived from values captured in table 11.4.7-1, derived based on the expanded uncertainty ue (1.96σ - confidence interval of 95 %) values. The common maximum accepted test system uncertainty values are applicable for all test methods addressing OTA OBUE test requirement.

Table 11.4.7-1: Test system specific MU values for the OTA OBUE measurement, FR1

|  | Expanded uncertainty ue (dB) |  |  |
| --- | --- | --- | --- |
|  | f ≦ 3GHz | 3GHz < f ≦ 4.2 GHz | 4.2GHz < f ≦ 6GHz |
| Indoor Anechoic Chamber |  |  |  |
| Compact Antenna Test Range |  |  |  |
| Near Field chamber |  |  |  |
| Reverberation chamber |  |  |  |
| Plane Wave Synthesizer |  |  |  |
| Common maximum accepted test system uncertainty | 1,8 | 2,0 | 2.0 |

Table 11.4.7-2: Test system specific MU values for the OTA OBUE measurement, FR2

|  | Expanded uncertainty ue (dB) |  |
| --- | --- | --- |
|  | 24.25<f<29.5GHz | 37<f<40GHz |
| Indoor Anechoic Chamber |  |  |
| Compact Antenna Test Range |  |  |
| Near Field chamber |  |  |
| Reverberation chamber |  |  |
| Plane Wave Synthesizer |  |  |
| Common maximum accepted test system uncertainty | 2.7 | 2.7 |

The MU value was agreed to be 1.4 dB for 0 – 3 GHz bands and 1.5 dB for 3 – 6 GHz bands. The MU in 4.2-6 GHz is valid for BS designed to operate in licensed spectrum.

For CATR the expanded MU is established as a root sum square combining of the dB values for the MU and the SE (see subclause 6.3.6), the MU was decided to be 2.7 dB for the frequency range 24.25<f<29.5GHz and 2.7 dB for the frequency range 37<f<40GHz.

An overview of the MU values for all the requirements is captured in clause 16.

### 11.4.8	Test Tolerance for OTA OBUE

Considering the methodology described in subclause 5.1, Test Tolerance values for OTA OBUE were derived based on values captured in subclause 11.4.7.

FR1: The TT value was agreed to be the same as the MU value for 0-10MHz from the carrier and 0dB for >10MHz from the carrier.

FR2: The TT value was agreed to be the same as the MU (2.7dB) for 0-10% of the BS channel bandwidth away from the carrier, and 0dB for >10% of BS channel bandwidth from the carrier.

Frequency range specific Test Tolerance values for the OTA OBUE test are defined in table 11.4.8-1.

Table 11.4.8-1: Test Tolerance values for the OTA BS output power, Normal test conditions, FR1

|  | f ≦ 3GHz | 3 GHz < f ≦ 4.2 GHz | 4.2GHz < f ≦ 6GHz |
| --- | --- | --- | --- |
| Test Tolerance (dB) | 1.8 | 2.0 | 2.0 |

Table 11.4.8-2: Test Tolerance values for the OTA BS output power, Normal test conditions, FR2

|  | 24.25 < f < 29.5 GHz | 37 < f < 40 GHz |
| --- | --- | --- |
| Test Tolerance (dB) | 2.7 | 2.7 |

An overview of the TT values for all the requirements is captured in clause 17.

--- Next change ---

#### 11.5.2.3	MU value derivation, FR1

Table 11.5.2.3-1 captures derivation of the expanded measurement uncertainty values for the OTA SEM in Indoor Anechoic Chamber (Normal test conditions, FR1).

Table 11.5.2.3-1: IAC MU value derivation for  OTA SEM

--- Next change ---

#### 11.5.3.3	MU value derivation, FR1

Table 11.5.3.3-1 captures derivation of the expanded measurement uncertainty values for the OTA SEM in CATR (Normal test conditions, FR1).

Table 11.5.3.3-1: CATR MU value derivation for OTA SEM

--- Next change ---

#### 11.5.4.3	MU value derivation, FR1

Refer to subclause 11.2.4.2.3 (i.e. the MU value derivation for the OTA BS output power measurement in NFTR) for MU value per point measurement.

MUTRP=MUper-point2+SE2

f≤3GHzMUTRP=1.152+0.752

3GHz≤f≤4.2GHzMUTRP=1.372+0.752

--- Next change ---

#### 11.5.5.3	MU value derivation, FR1

Table 11.5.5.3-1 captures derivation of the expanded measurement uncertainty values for OTA SEM measurement in Reverberation chamber (Normal test conditions, FR1).

Table 11.5.5.3-1: Reverberation chamber MU value derivation for OTA SEM

--- Next change ---

#### 11.5.6.3	MU value derivation, FR1

The MU value for OTA OBUE in PWS is the same as in subclause 11.2.6.3 (i.e. OTA BS output power MU in PWS).

--- Next change ---

### 11.5.7	Maximum accepted test system uncertainty

Maximum test system uncertainties derivation methodology was described in subclause 5.1. The maximum accepted test system uncertainty values was derived based on test system specific values.

According to the methodology referred above, the common maximum accepted test system uncertainty values for OTA SEM test can be derived from values captured in table 11.5.7-1, derived based on the expanded uncertainty ue (1.96σ - confidence interval of 95 %) values. The common maximum accepted test system uncertainty values are applicable for all test methods addressing OTA SEM test requirement.

Table 11.5.7-1: Test system specific MU values for the OTA SEM measurement

|  | Expanded uncertainty ue (dB) |  |
| --- | --- | --- |
|  | f ≦ 3GHz | 3GHz < f ≦ 4.2 GHz |
| Indoor Anechoic Chamber |  |  |
| Compact Antenna Test Range |  |  |
| Near Field chamber |  |  |
| Reverberation chamber |  |  |
| Plane Wave Synthesizer |  |  |
| Common maximum accepted test system uncertainty | 1,8 | 2,0 |

An overview of the MU values for all the requirements is captured in clause 16.

### 11.5.8	Test Tolerance for OTA SEM

Considering the methodology described in subclause 5.1, Test Tolerance values for OTA SEM were derived based on values captured in subclause 11.5.7.

Frequency range specific Test Tolerance values for the OTA SEM test are defined in table 11.5.8-1.

Table 11.5.8-1: Test Tolerance values for the OTA BS output power, Normal test conditions

|  | f ≦ 3GHz | 3 GHz < f ≦ 4.2 GHz |
| --- | --- | --- |
| Test Tolerance (dB) | 1.8 | 2.0 |

An overview of the TT values for all the requirements is captured in clause 17.

--- Next change ---

#### 12.2.2.3	MU value derivation, FR1

Table 12.2.2.3-1 captures derivation of the expanded measurement uncertainty values for OTA TX spurious emissions measurements in general directional Chamber (Normal test conditions, FR1).

Table 12.2.2.3-1: General directional chamber MU value derivation for the TX spurious emissions, FR1

#### 12.2.3.x	MU value derivation, FR2

Table 12.2.3.3-1 captures derivation of the expanded measurement uncertainty values for OTA TX spurious emissions measurements in CATR (Normal test conditions, FR2).

Table 12.2.3.x-1: IAC value derivation for TX spurious emissions, FR2

#### 12.2.3.3	MU value derivation, FR2

Table 12.2.3.3-1 captures derivation of the expanded measurement uncertainty values for OTA TX spurious emissions measurements in CATR (Normal test conditions, FR2).

Table 12.2.3.4-1: CATR value derivation for TX spurious emissions, FR2

--- Next change ---

#### 12.2.3.4	MU value derivation

Table 12.2.3.4-1 captures derivation of the expanded measurement uncertainty values for OTA TX spurious emissions measurements in Reverberation Chamber (Normal test conditions).

Table 12.2.3.4-1: Reverberation Chamber value derivation for TX spurious emissions, 380 MHz – 26 GHz

Table 12.2.3.4-1-2: Reverberation Chamber value derivation for TX spurious emissions, 18 GHz – 60 GHz

--- Next change ---

### 12.2.4	Maximum accepted test system uncertainty

Considering the methodology described in sub-clause 5.1, Test Tolerance values for OTA transmitter spurious emissions were derived based on values captured in sub-clause 12.2.3.

Frequency range specific Test Tolerance values for the OTA SEM test are defined in table 12.7.1-1.

Hence, we have the following MU values for the whole spurious emissions range (for FR1 and FR2 cases):

Table 12.2.4-1: Spurious emissions MU values

| Frequency range | MU (dB) |
| --- | --- |
| 30 MHz < f ≦ 6 GHz | 2.3 |
| 6 GHz < f ≦ 18 GHz | 2.7 |
| 18 GHz < f ≦ 40 GHz | 2.7 |
| 40 GHz < f ≦ 60 GHz | 5 |

### 12.2.5	Test Tolerance for OTA TX spurious emissions

The conduced test tolerance for the mandatory spurious emissions requirements is zero. As the requirements are set by regulatory limits the same test tolerance is used for OTA.

TT = 0.

--- Next change ---

#### 12.3.2.3	MU value derivation

Table 12.3.2.3-1 captures derivation of the expanded measurement uncertainty values for OTA RX spurious emissions measurements in Indoor Anechoic Chamber (Normal test conditions, FR1).

Table 12.3.2.3-1: General chamber MU value derivation for RX spurious emissions

--- Next change ---

### 12.3.3	Maximum accepted test system uncertainty

The TRP MU is very similar to that for the transmitter mandatory spurious emissions. However, the receiver requirements are at a much lower power level so TRP calculation may be affected by the noise floor of the measurement system.

This range can be split into a number of regions:

30 MHz < f ≤ 6 GHz

The same value is assumed for receiver spurious emissions as for transmitter spurious emissions. This is the same as the in band FR2 MU value.

6 GHz < f ≤ 18 GHz

The same value is assumed for receiver spurious emissions as for transmitter spurious emissions. This is the same as the in band FR2 MU value.

18 GHz < f ≤ 40 GHz

The same value is assumed for receiver spurious emissions as for transmitter spurious emissions. This is calculated in clause 12.7.1.2.

40 GHz < f ≤ 60 GHz

The same value is assumed for receiver spurious emissions as for transmitter spurious emissions. This is calculated in clause 12.7.1.2.

Table 12.7.2.2-1: Receiver Spurious emissions MUvalues

| Frequency range | MU (dB) |
| --- | --- |
| 30 MHz < f ≦ 6 GHz | 2.5 |
| 6 GHz < f ≦ 18 GHz | 2.7 |
| 18 GHz < f ≦ 40 GHz | 2.7 |
| 40 GHz < f ≦ 60 GHz | 5 |

### 12.3.4	Test Tolerance for OTA RX spurious emissions

The conduced test tolerance for the receiver spurious emissions requirements is zero. However for OTA BS the receiver spurious emissions requirements only apply to TTD in OFF mode. As such the limit is set by RAN4 to be considerably lower than the equivalent regulatory requirement.

In addition due to the difficulty in measuring low levels of TRP close to the measurement system noise floor the risk of false failures is high. As the risk is due to the noise floor of the measurement system it cannot be mitigated by BS design.

Hence it has been agreed that for receiver spurious emissions the TT = MU.

--- Next change ---

#### 12.4.2.3	MU value derivation

Table 12.4.2.3-1 captures derivation of the expanded measurement uncertainty values for additional (co-existence) OTA TX spurious emissions measurements in CATR (Normal test conditions, FR1).

Table 12.4.2.3-1: CATR MU value derivation for additional (co-existence) OTA TX spurious emissions

--- Next change ---

### 12.4.3	Maximum accepted test system uncertainty

The additional spurious emission requirements including the co-existence with other BS in the same geographical area are based on existing co-existence with other 3GPP bands so the frequency ranges for the uncertainty assessment are different from the mandatory spurious emissions. In this case the uncertainty budgets for the in-band chambers are considered in the MU analysis. As the TRP level is at a low power level the measurement system dynamic range uncertainty is also considered as with the receiver spurious emissions.

Table 12.3.3-1: Test system specific MU values for the Additional (coexistence) spurious emissions measurement

| chamber | Expanded uncertainty [dB] |  |  |
| --- | --- | --- | --- |
|  | 30MHz<f≤3 GHz | 3GHz<f ≤4.2GHz | 4.2<f<6 GHz |
| CATR |  |  |  |
| Common maximum accepted test system uncertainty | 2.60 | 3.00 | 3.50 |

NOTE: There are currently no additional spurious emissions requirements or co-existence requirements for FR2.

### 12.4.4	Test Tolerance for additional spurious emissions requirements

The conduced test tolerance for the additional spurious emissions requirements is zero.

However for OTA BS, the difficulty in measuring TRP of the additional spurious emissions requirements at low levels close to the measurement system noise floor means the high risk of false failures. As the risk is due to the noise floor of the measurement system it cannot be mitigated by BS design.

As the 3GPP to 3GPP co-existence requirements are not regulatory but set by RAN4 to assist with co-existence of 3GPP systems in the same geographical area it is acceptable for RAN4 to set the TT value to be non-zero.

Hence it has been agreed that for 3GPP to 3GPP co-existence spurious emissions the TT = MU.

Table 12.4.4-1: TT values for the Additional (coexistence) spurious emissions measurement

|  | 30MHz<f≤3 GHz | 3GHz<f ≤4.2GHz | 4.2<f<6 GHz |
| --- | --- | --- | --- |
| Test Tolerance (dB) | 1.8 | 2.0 |  |

For PHS, and public safety additional requirements the TT = 0 dB.

Some additional requirements such as the protection of PHS and the 700 and 800 MHz public safety bands, are regulatory so it is not possible to have a non zero TT, hence for these requirements the TT is zero.

--- End of changes ---