## Summary
| Expanded uncertainty [dB] | [23 empty] | A method exceeds agreed value |  |  | Review comments (RAN4#94)
| IAC |  |  | CATR |  |  | NF |  |  | 1D |  |  | PWS |  |  | Reverb |  |  | General chamber |  |  | Common maximum accepted test system uncertainty
| f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz
EIRP | 0.86972 | 1.0574 | 1.0574 | 1.1143 | 1.2663 | 1.2663 | 1.0143 | 1.1041 | 1.1041 | 0.90418 | 1.0952 | 1.0952 | 0.98073 | 1.1826 | 1.1826 | x | x | x | x | x | x | 1.1 | 1.3 | 1.3
EIRP - Extreme temp | 2.46 | 2.5324 | 2.5324 | 2.5103 | 2.5814 | 2.5515 | [15 empty] | 2.5 | 2.6 | 2.6
DL RS | 1.1519 | 1.4364 | 1.4364 | 1.3461 | 1.5965 | 1.5965 | 1.1658 | 1.387 | 1.387 | x | x | x | x | x | x | x | x | x | x | x | x | 1.3 | 1.5 | 1.5 |  | x | x
power dynamics | 0.40801 | 0.40801 | 0.40801 | 0.392 | 0.392 | 0.392 | 0.392 | 0.392 | 0 | x | x | x | x | x | x | x | x | x | x | x | x | 0.4 | 0.4 | 0.4
EVM (%) | 1.0288 | 1.0288 | 1.0288 | 1.0403 | 1.0403 | 1.0403 | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | 1 | 1 | 1
In-band TRP | 1.1484 | 1.2964 | 1.2964 | 1.3854 | 1.5104 | 1.5104 | 1.2611 | 1.3343 | 1.3343 | x | x | x | 1.2346 | 1.4003 | 1.4003 | 1.3653 | 1.4619 | 1.4619 | x | x | x | 1.4 | 1.5 | 1.5
ACLR- abs | 1.1484 | 1.2964 | 1.2964 | 1.3854 | 1.5104 | 1.5104 | 1.2611 | 1.3343 | 1.3343 | x | x | x | 1.2346 | 1.4003 | 1.4003 | 1.3653 | 1.4619 | 1.4619 | x | x | x | 2.2 | 2.7 | 2.7
ACLR-rel | x | x | x | 1.1181 | 1.302 | 1.302 | x | x | x | x | x | x | 1.0179 | 1.2215 | 1.2215 | 1.3653 | 1.4619 | 1.4619 | x | x | x | 1 | 1.2 | 1.2 | x | x | x
OBUE | 1.1484 | 1.2964 | 1.2964 | 1.3854 | 1.5104 | 1.5104 | 1.2611 | 1.3343 | 1.3343 | x | x | x | 1.2346 | 1.4003 | 1.4003 | 1.3974 | 1.4619 | 1.4619 | x | x | x | 1.8 | 2 | 2
SEM | 1.1484 | 1.2964 | 1.2964 | 1.3854 | 1.5104 | 1.5104 | 1.2611 | 1.3343 | 1.3343 | x | x | x | 1.2346 | 1.4003 | 1.4003 | 1.3974 | 1.4619 | 1.4619 | x | x | x | 1.8 | 2 | 2
COEX EM | 2.6153 | 3.052 | 3.4661 | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | 2.6 | 3 | 3.5 |  | x
| General Chamber |  |  | Reverb |  |  | Agreed value |  |  | A method exceeds agreed value |  |  |  | Review comments (RAN4#94)
| 30MHz<f≤6 GHz | 6GHz<f ≤19GHz | 19GHz<f ≤26GHz | 30MHz<f≤6 GHz | 6GHz<f ≤19GHz | 19GHz<f ≤26GHz | 30MHz<f≤6 GHz | 6GHz<f ≤19GHz | 19GHz<f ≤26GHz | 30MHz<f≤6 GHz | 6GHz<f ≤19GHz | 19GHz<f ≤26GHz
OOB EM | 2.2939 | 4.1358 | 4.1358 | 3.2895 | 3.3763 | 3.3763 | 2.3 | 4.2 | 4.2 | x
RX EM | 2.5023 | 4.2549 | 4.2549 | 0 | 0 | 0 | 2.5 | 4.2 | 4.2 |  | x | x

## TE
Review comments (RAN4#94)
Test Equipment Errors
UID | Instrument | Uncertainty value (for the prob dist type) |  |  | Probability distribution | divisor | Standard uncertainty σ (dB)
|  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz |  |  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz
C1-1 | RF power measurement equipment (e.g. spectrum analyzer, power meter) | 0.14 | 0.26 | 0.26 | Gaussian | 1 | 0.14 | 0.26 | 0.26
C1-2 | Uncertainty of the RF signal generator | 0.46 | 0.46 | 0.46 | Gaussian | 1 | 0.46 | 0.46 | 0.46
C1-3 | Uncertainty of the network analyzer | 0.13 | 0.2 | 0.2 | Gaussian | 1 | 0.13 | 0.2 | 0.2
C1-4 | Uncertainty of the absolute gain of the reference antenna | 0.50229 | 0.43301 | 0.43301 | Rectangular | 1.7321 | 0.29 | 0.25 | 0.25
C1-5 | Measurement Reciever (Co-location) | 0.41 | 0.74 | 0.8 | Gaussian | 1 | 0.41 | 0.74 | 0.8
C1-6 | Noise figure measurememt accuracy | 0.2 | 0.2 | 0.2 | Gaussian | 1 | 0.2 | 0.2 | 0.2
Test Equipment Errors (same errors and values are consistant with above, buit greater freq range)
UID | Instrument | Uncertainty value (for the prob dist type) |  |  |  |  | Probability distribution | divisor | Standard uncertainty σ (dB)
|  | f≤3 GHz | 3<f<≤.8 GHz | 3.8<f≤12.75 GHz | 12.75<f≤19 GHz | 19<f≤26 GHz |  |  | f≤3 GHz | 3<f<≤.8 GHz | 3.8<f≤12.75 GHz | 12.75<f≤19 GHz | 19<f≤26 GHz
C1-1 | RF power measurement equipment (e.g. spectrum analyzer, power meter) | 0.14 | 0.26 | 0.26 | 0.37 | 0.37 | Gaussian | 1 | 0.14 | 0.26 | 0.26 | 0.37 | 0.37
C1-3 | Uncertainty of the network analyzer | 0.13 | 0.2 | 0.2 | 0.2 | 0.2 | Gaussian | 1 | 0.13 | 0.2 | 0.2 | 0.2 | 0.2
Conducted MU (minus missmatch) - up to 6GHz
| Instrument | Uncertainty value (for the prob dist type) |  |  | Probability distribution | divisor | Standard uncertainty σ (dB)
|  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz |  |  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz
C3-1 | DL-RS MU derived from conducted spec | 0.41 | 0.56 | 0.56 | Gaussian | 1 | 0.41 | 0.56 | 0.56
C3-2 | Total power dynamic range conducted unbcertainty | 0.2 | 0.2 | 0.2 | Gaussian | 1 | 0.2 | 0.2 | 0.2
C3-5 | Additional (COEX) emssions - Condcuted Uncertainty (minus missmatch) | 1.02 | 1.28 | 1.53 | Gaussian | 1 | 1.02 | 1.28 | 1.53
Conducted MU (minus missmatch) - Whole Spectrum
| Instrument | Uncertainty value (for the prob dist type) |  |  | Probability distribution | divisor | Standard uncertainty σ (dB)
|  | 30MHz<f≤6 GHz | 6GHz<f ≤19GHz | 19GHz<f ≤26GHz |  |  | 30MHz<f≤6 GHz | 6GHz<f ≤19GHz | 19GHz<f ≤26GHz
C3-3 | Transmitter mandatory spurious emissions - Conducted Uncertainty (minus miss match) | 1 | 2 | 2 | Gaussian | 1 | 1 | 2 | 2
C3-4 | Receiver spurious emissions - Conducted Uncertainty (minus missmatch) | 1 | 2 | 2 | Gaussian | 1 | 1 | 2 | 2

## IA-Er
Indoor anechoic | [10 empty] | Review comments (RAN4#94)
|  | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz |  |  |  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz
Stage 2: DUT measurement
A1-1 | Positioning misalignment between the AAS BS and the reference antenna | 0.03 | 0.03 | 0.03 | Rectangular | 1.7321 | 1 | 0.017321 | 0.017321 | 0.017321
A1-2 | Pointing misalignment between the AAS BS and the receiving antenna | 0.3 | 0.3 | 0.3 | Rectangular | 1.7321 | 1 | 0.17321 | 0.17321 | 0.17321
A1-3 | Quality of quiet zone | 0.1 | 0.1 | 0.1 | Gaussian | 1 | 1 | 0.1 | 0.1 | 0.1
A1-4a | Polarization mismatch between the AAS BS and the receiving antenna | 0.01 | 0.01 | 0.01 | Rectangular | 1.7321 | 1 | 0.0057735 | 0.0057735 | 0.0057735
A1-5 | Mutual coupling between the AAS BS and the receiving antenna | 0 | 0 | 0 | Rectangular | 1.7321 | 1 | 0 | 0 | 0
A1-6 | Phase curvature | 0.05 | 0.05 | 0.05 | Gaussian | 1 | 1 | 0.05 | 0.05 | 0.05
C1-1 | RF power measurement equipment (e.g. spectrum analyzer, power meter) | 0.14 | 0.26 | 0.26 | Gaussian | 1 | 1 | 0.14 | 0.26 | 0.26
A1-7 | Impedance mismatch in the receiving chain | 0.14 | 0.33 | 0.33 | U-shaped | 1.4142 | 1 | 0.098995 | 0.23335 | 0.23335
A1-8 | Random uncertainty | 0.1 | 0.1 | 0.1 | Rectangular | 1.7321 | 1 | 0.057735 | 0.057735 | 0.057735
A1-17 | Quality of quiet zone (extreme) | 0.6 | 0.6 | 0.6 | Gaussian | 1 | 1 | 0.6 | 0.6 | 0.6
A1-19 | Radome loss variation | 0.4 | 0.4 | 0.4 | Rectangular | 1.7321 | 1 | 0.23094 | 0.23094 | 0.23094
A1-18 | Wet radome loss variation | 0.95 | 0.95 | 0.95 | Gaussian | 1 | 1 | 0.95 | 0.95 | 0.95
A1-20 | Change in absorber behaviour | 0.1 | 0.1 | 0.1 | Gaussian | 1 | 1 | 0.1 | 0.1 | 0.1
A1-16 | Frequency flatness | 0.25 | 0.25 | 0.25 | Gaussian | 1 | 1 | 0.25 | 0.25 | 0.25
Stage 1: Calibration measurement
A1-9 | Impedance mismatch between the receiving antenna and the network analyzer | 0.05 | 0.05 | 0.05 | U-shaped | 1.4142 | 1 | 0.035355 | 0.035355 | 0.035355
A1-10 | Positioning and pointing misalignment between the reference antenna and the receiving antenna | 0.01 | 0.01 | 0.01 | Rectangular | 1.7321 | 1 | 0.0057735 | 0.0057735 | 0.0057735
A1-11 | Impedance mismatch between the reference antenna and the network analyzer. | 0.05 | 0.05 | 0.05 | U-shaped | 1.4142 | 1 | 0.035355 | 0.035355 | 0.035355
A1-3 | Quality of quiet zone | 0.1 | 0.1 | 0.1 | Gaussian | 1 | 1 | 0.1 | 0.1 | 0.1
A1-4b | Polarization mismatch between the reference antenna and the receiving antenna | 0.01 | 0.01 | 0.01 | Rectangular | 1.7321 | 1 | 0.0057735 | 0.0057735 | 0.0057735
A1-5 | Mutual coupling between the reference antenna and the receiving antenna | 0 | 0 | 0 | Rectangular | 1.7321 | 1 | 0 | 0 | 0
A1-6 | Phase curvature | 0.05 | 0.05 | 0.05 | Gaussian | 1 | 1 | 0.05 | 0.05 | 0.05
C1-3 | Uncertainty of the network analyzer | 0.13 | 0.2 | 0.2 | Gaussian | 1 | 1 | 0.13 | 0.2 | 0.2
A1-12 | Influence of the reference antenna feed cable | 0.05 | 0.05 | 0.05 | Rectangular | 1.7321 | 1 | 0.028868 | 0.028868 | 0.028868
A1-13 | Reference antenna feed cable loss measurement uncertainty | 0.06 | 0.06 | 0.06 | Gaussian | 1 | 1 | 0.06 | 0.06 | 0.06
A1-14 | Influence of the receiving antenna feed cable | 0.05 | 0.05 | 0.05 | Rectangular | 1.7321 | 1 | 0.028868 | 0.028868 | 0.028868
C1-4 | Uncertainty of the absolute gain of the reference antenna | 0.50229 | 0.43301 | 0.43301 | Rectangular | 1.7321 | 1 | 0.29 | 0.25 | 0.25
A1-15 | Uncertainty of the absolute gain of the receiving antenna | 0 | 0 | 0 | Rectangular | 1.7321 | 1 | 0 | 0 | 0
A1-17 | Impedance mismatch in the transmitting chain | 0.14 | 0.23 | 0.23 | U-shaped | 1.4142 | 1 | 0.098995 | 0.16263 | 0.16263
A1-18 | Gain variation of power amplifier | 0.02 | 0.02 | 0.02 | Rectangular | 1.7321 | 1 | 0.011547 | 0.011547 | 0.011547
A1-19 | Influence of the transmitting antenna feed cable a) Flexing cables, adapters, attenuators, and connector repeatability | 0.05 | 0.05 | 0.05 | Rectangular | 1.7321 | 1 | 0.028868 | 0.028868 | 0.028868

## CATR-Er
CATR | [10 empty] | Review comments (RAN4#94)
|  | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz |  |  |  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz
Stage 2: DUT measurement
A2-1a | Misalignment DUT & pointing error for EIRP | 0 | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 | 0 | RK: this is 0 in the EIPR measurement in TR 37.842 and 0.3 in the TRP meaurements in TR37.843 - is this correct? And why is the error different?
A2-18 | Misalignment DUT & pointing error for TRP | 0.3 | 0.3 | 0.3 | Rectangular | 1.7321 | 1 | 0.17321 | 0.17321 | 0.17321 | RK:added separate error for TRP after discussion with Torbjorn(Ericsson)
C1-1 | RF power measurement equipment (e.g. spectrum analyzer, power meter) | 0.14 | 0.26 | 0.26 | Gaussian | 1 | 1 | 0.14 | 0.26 | 0.26
A2-2a | Standing wave between DUT and test range antenna | 0.21 | 0.21 | 0.21 | U-shaped | 1.4142 | 1 | 0.14849 | 0.14849 | 0.14849
A2-3 | RF leakage (SGH connector terminated & test range antenna connector cable terminated) | 0.0012 | 0.0012 | 0.0012 | Gaussian | 1 | 1 | 0.0012 | 0.0012 | 0.0012
A2-4a | QZ ripple DUT | 0.0928 | 0.0928 | 0.0928 | Gaussian | 1 | 1 | 0.0928 | 0.0928 | 0.0928
A2-10 | Miscellaneous uncertainty
A2-12 | Frequency flatness | 0.25 | 0.25 | 0.25 | Gaussian | 1 | 1 | 0.25 | 0.25 | 0.25
A2-17 | Measurement system dynamic range uncertainty | 0.51 | 0.51 | 0.51 | Gaussian | 1 | 1 | 0.51 | 0.51 | 0.51
A2-13 | Quality of quiet zone (extreme) | 0.6 | 0.6 | 0.6 | Gaussian | 1 | 1 | 0.6 | 0.6 | 0.6
A2-14 | Wet radome loss variation | 0.4 | 0.4 | 0.4 | Rectangular | 1.7321 | 1 | 0.23094 | 0.23094 | 0.23094
A2-15 | Radome loss variation | 0.95 | 0.95 | 0.95 | Gaussian | 1 | 1 | 0.95 | 0.95 | 0.95
A2-16 | Change in absorber behaviour | 0.1 | 0.1 | 0.1 | Gaussian | 1 | 1 | 0.1 | 0.1 | 0.1
Stage 1: Calibration measurement
C1-3 | Uncertainty of the network analyzer | 0.13 | 0.2 | 0.2 | Gaussian | 1 | 1 | 0.13 | 0.2 | 0.2
A2-5 | Mismatch of receiver chain | 0.127 | 0.325 | 0.325 | U-shaped | 1.4142 | 1 | 0.089803 | 0.22981 | 0.22981
A2-6 | Insertion loss variation of receiver chain | 0.18 | 0.18 | 0.18 | Rectangular | 1.7321 | 1 | 0.10392 | 0.10392 | 0.10392
A2-3 | RF leakage, (SGH connector terminated & test range antenna connector cable terminated) | 0.0012 | 0.0012 | 0.0012 | Gaussian | 1 | 1 | 0.0012 | 0.0012 | 0.0012 | RK:repeat
A2-7 | Influence of the calibration antenna feed cable: | 0.022 | 0.022 | 0.022 | U-shaped | 1.4142 | 1 | 0.015556 | 0.015556 | 0.015556
C1-4 | Uncertainty of the absolute gain of the reference antenna | 0.50229 | 0.43301 | 0.43301 | Rectangular | 1.7321 | 1 | 0.29 | 0.25 | 0.25
A2-8 | Misalignment positioning system | 0 | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 | 0
A2-1b | Misalignment of calibration antenna and test range antenna | 0.5 | 0.5 | 0.5 | Exp. normal | 2 | 1 | 0.25 | 0.25 | 0.25
A2-9 | Rotary Joints | 0.048 | 0.048 | 0.048 | U-shaped | 1.4142 | 1 | 0.033941 | 0.033941 | 0.033941
A2-2b | Standing wave between calibration antenna and test range antenna | 0.09 | 0.09 | 0.09 | U-shaped | 1.4142 | 1 | 0.06364 | 0.06364 | 0.06364
A2-4b | QZ ripple calibration antenna | 0.009 | 0.009 | 0.009 | Gaussian | 1 | 1 | 0.009 | 0.009 | 0.009
A2-11 | Switching uncertainty | 0.26 | 0.26 | 0.26 | Rectangular | 1.7321 | 1 | 0.15011 | 0.15011 | 0.15011
A2-13 | Quality of quiet zone (extreme)
A2-14 | Wet radome loss variation
A2-15 | Radome loss variation
A2-16 | Change in absorber behaviour

## NF-Er
Near Field | [10 empty] | Review comments (RAN4#94)
|  | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz |  |  |  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz
Stage 2: DUT measurement
A3-1 | Axes Intersection | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0
A3-2 | Axes Orthogonality | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0
A3-3 | Horizontal Pointing | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0
A3-4 | Probe Vertical Position | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0
A3-5 | Probe H/V pointing | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0
A3-6 | Measurement Distance | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0
A3-7 | Amplitude and Phase Drift | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0
A3-8 | Amplitude and Phase Noise | 0.02 | 0.02 | 0.02 | Gaussian | 1 | 1 | 0.02 | 0.02 | 0.02
A3-9 | Leakage and Crosstalk | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0
A3-10 | Amplitude Non-Linearity | 0.04 | 0.04 | 0.04 | Gaussian | 1 | 1 | 0.04 | 0.04 | 0.04
A3-11 | Amplitude and Phase Shift in rotary joints | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0
A3-12 | Channel Balance Amplitude and Phase | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0
A3-13 | Probe Polarization Amplitude and Phase | 0.0001 | 0.0001 | 0.0001 | Gaussian | 1 | 1 | 0.0001 | 0.0001 | 0.0001
A3-14 | Probe Pattern Knowledge | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0
A3-15 | Multiple Reflections | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0
A3-16 | Room Scattering | 0.09 | 0.09 | 0.09 | Gaussian | 1 | 1 | 0.09 | 0.09 | 0.09
A3-17 | DUT support Scattering | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0
A3-18 | Scan Area Truncation | 0.003 | 0.003 | 0.003 | Gaussian | 1 | 1 | 0.003 | 0.003 | 0.003
A3-19 | Sampling Point Offset | 0.0058 | 0.0058 | 0.0058 | Gaussian | 1 | 1 | 0.0058 | 0.0058 | 0.0058
A3-20 | Spherical Mode Truncation | 0.015 | 0.015 | 0.015 | Gaussian | 1 | 1 | 0.015 | 0.015 | 0.015
A3-21 | Positioning | 0.03 | 0.03 | 0.03 | Rectangular | 1.73 | 1 | 0.017341 | 0.017341 | 0.017341
A3-22 | Probe Array Uniformity | 0.055 | 0.055 | 0.055 | Gaussian | 1 | 1 | 0.055 | 0.055 | 0.055
A3-23 | Mismatch of receiver chain | 0.284 | 0.284 | 0.284 | U-Shaped | 1.41 | 1 | 0.20142 | 0.20142 | 0.20142
A3-24 | Insertion loss of receiver chain | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0
A3-25 | Uncertainty of the absolute gain of the probe antenna | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0
C1-1 | RF power measurement equipment (e.g. spectrum analyzer, power meter) | 0.14 | 0.26 | 0.26 | Gaussian | 1 | 1 | 0.14 | 0.26 | 0.26
A3-26 | Measurement repeatability - positioning repeatability | 0.15 | 0.15 | 0.15 | Gaussian | 1 | 1 | 0.15 | 0.15 | 0.15
A3-33 | Test system frequency flatness | 0.25 | 0.25 | 0.25 | Gaussian | 1 | 1 | 0.25 | 0.25 | 0.25
Stage 1: Calibration measurement
C1-3 | Uncertainty of the network analyzer | 0.13 | 0.2 | 0.2 | Gaussian | 1 | 1 | 0.13 | 0.2 | 0.2
A3-27 | Mismatch of receiver chain | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0
A3-28 | Insertion loss of receiver chain | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0
A3-29 | Mismatch in the connection of the calibration antenna | 0.02 | 0.02 | 0.02 | U-Shaped | 1.41 | 1 | 0.014184 | 0.014184 | 0.014184
A3-30 | Influence of the calibration antenna feed cable | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0
A3-31 | Influence of the probe antenna cable | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0
C1-4 | Uncertainty of the absolute gain of the reference antenna | 0.50229 | 0.43301 | 0.43301 | Rectangular | 1.7321 | 1 | 0.29 | 0.25 | 0.25
A3-32 | Short term repeatability | 0.088 | 0.088 | 0.088 | Gaussian | 1 | 1 | 0.088 | 0.088 | 0.088

## 1D-Er
One Dimensional Compact Range Chamber | [10 empty] | Review comments (RAN4#94)
|  | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz |  |  |  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz
Stage 2: DUT measurement
A4-1 | Misalignment DUT and pointing error | 0 | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 | 0
A4-2a | Standing wave between DUT and test range antenna | 0.18 | 0.18 | 0.18 | U-shaped | 1.4142 | 1 | 0.12728 | 0.12728 | 0.12728
A4-3a | Quiet zone ripple DUT | 0.0325 | 0.0325 | 0.0325 | Gaussian | 1 | 1 | 0.0325 | 0.0325 | 0.0325
A4-4a | Phase curvature AAS | 0.01 | 0.01 | 0.01 | Gaussian | 1 | 1 | 0.01 | 0.01 | 0.01
A4-5a | Polarization mismatch between DUT and receiving antenna | 0.05 | 0.05 | 0.05 | Rectangular | 1.7321 | 1 | 0.028868 | 0.028868 | 0.028868
A4-6a | Mutual coupling between DUT and receiving antenna | 0 | 0 | 0 | Rectangular | 1.7321 | 1 | 0 | 0 | 0
C1-1 | RF power measurement equipment (e.g. spectrum analyzer, power meter) | 0.14 | 0.26 | 0.26 | Gaussian | 1 | 1 | 0.14 | 0.26 | 0.26
A4-7 | Impedance mismatch in receiving chain | 0.0056 | 0.01 | 0.01 | U-shaped | 1.4142 | 1 | 0.0039598 | 0.0070711 | 0.0070711
A4-8a | RF leakage (DUT connector terminated and test range antenna connector cable terminated) | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0
Stage 1: Calibration measurement
A4-9 | Misalignment positioning system | 0 | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 | 0
A4-10 | Pointing error between reference antenna and test range antenna | 0 | 0 | 0 | Rectangular | 1.7321 | 1 | 0 | 0 | 0
A4-11 | Impedance mismatch in path to reference antenna | 0.05 | 0.05 | 0.05 | U-shaped | 1.4142 | 1 | 0.035355 | 0.035355 | 0.035355
A4-12 | Impedance mismatch in path to compact probe | 0.03 | 0.03 | 0.03 | U-shaped | 1.4142 | 1 | 0.021213 | 0.021213 | 0.021213
A4-2b | Standing wave between reference antenna and receiving antenna | 0.09 | 0.09 | 0.09 | U-shaped | 1.4142 | 1 | 0.06364 | 0.06364 | 0.06364
A4-3b | Quiet zone ripple reference antenna | 0.178 | 0.178 | 0.178 | Gaussian | 1 | 1 | 0.178 | 0.178 | 0.178
A4-4b | Phase curvature refernce antenna | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0
A4-5b | Polarization mismatch between reference antenna and receiving antenna | 0.05 | 0.05 | 0.05 | Rectangular | 1.7321 | 1 | 0.028868 | 0.028868 | 0.028868
A4-6a | Mutual coupling between reference antenna and receiving antenna | 0 | 0 | 0 | Rectangular | 1.7321 | 1 | 0 | 0 | 0
C1-1 | RF power measurement equipment (e.g. spectrum analyzer, power meter) | 0.14 | 0.26 | 0.26 | Gaussian | 1 | 1 | 0.14 | 0.26 | 0.26
A4-13 | Influence of the reference antenna feed cable (flexing cables, adapters, attenuators, connector repeatability) | 0.082 | 0.082 | 0.082 | Rectangular | 1.7321 | 1 | 0.047343 | 0.047343 | 0.047343
A4-14 | Mismatch of receiver chain | 0.2 | 0.3 | 0.3 | U-shaped | 1.4142 | 1 | 0.14142 | 0.21213 | 0.21213
A4-15 | Insertion loss of receiver chain | 0.18 | 0.18 | 0.18 | Rectangular | 1.7321 | 1 | 0.10392 | 0.10392 | 0.10392
C1-4 | Uncertainty of the absolute gain of the reference antenna | 0.50229 | 0.43301 | 0.43301 | Rectangular | 1.7321 | 1 | 0.29 | 0.25 | 0.25
A4-8b | RF leakage (SGH connector terminated and test range antenna connector cable terminated. | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0

## PWS-Er
Plane wave synthesiser | [10 empty] | Review comments (RAN4#94)
|  | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz |  |  |  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz
Stage 2: DUT measurement
A7-1a | Misalignment DUT & pointing error | 0.1 | 0.1 | 0.1 | Rectangular | 1.7321 | 1 | 0.057735 | 0.057735 | 0.057735
C1-1 | RF power measurement equipment | 0.14 | 0.26 | 0.26 | Gaussian | 1 | 1 | 0.14 | 0.26 | 0.26
A7-2a | Longitudinal position uncertainty (i.e. standing wave and imperfect field synthesis) for DUT antenna | 0.05 | 0.14 | 0.14 | Rectangular | 1.7321 | 1 | 0.028868 | 0.080829 | 0.080829
A7-3 | RF leakage (calibration antenna connector terminated) | 0.086 | 0.086 | 0.086 | Gaussian | 1 | 1 | 0.086 | 0.086 | 0.086
A7-4a | QZ ripple with DUT | 0.42 | 0.43 | 0.43 | Rectangular | 1.7321 | 1 | 0.24249 | 0.24826 | 0.24826
A7-5 | Miscellaneous Uncertainty | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0
A7-14 | System non-linearity | 0.062 | 0.062 | 0.062 | Rectangular | 1.7321 | 1 | 0.035796 | 0.035796 | 0.035796
A7-13 | Frequency Flatness | 0.13 | 0.13 | 0.13 | Rectangular | 1.7321 | 1 | 0.075056 | 0.075056 | 0.075056
Stage 1: Calibration measurement
C1-3 | Uncertainty of network analyzer | 0.13 | 0.2 | 0.2 | Gaussian | 1 | 1 | 0.13 | 0.2 | 0.2
A7-6 | Mismatch (i.e. reference antenna, network analyser and reference cable) | 0.127 | 0.325 | 0.325 | U-shaped | 1.4142 | 1 | 0.089803 | 0.22981 | 0.22981
A7-7 | Insertion loss variation | 0.18 | 0.18 | 0.18 | Rectangular | 1.7321 | 1 | 0.10392 | 0.10392 | 0.10392
A7-3 | RF leakage (calibration antenna connector terminated) | 0.086 | 0.086 | 0.086 | Gaussian | 1 | 1 | 0.086 | 0.086 | 0.086 | RK: same as above
A7-8 | Influence of the calibration antenna feed cable | 0.103 | 0.104 | 0.104 | Rectangular | 1.7321 | 1 | 0.059467 | 0.060044 | 0.060044
C1-4 | Uncertainty of the absolute gain of the reference antenna | 0.50229 | 0.43301 | 0.43301 | Rectangular | 1.7321 | 1 | 0.29 | 0.25 | 0.25
A7-9 | Misalignment of positioning system | 0 | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 | 0
A7-1b | Misalignment of calibration antenna & pointing error | 0.05 | 0.05 | 0.05 | Rectangular | 1.7321 | 1 | 0.028868 | 0.028868 | 0.028868
A7-10 | Rotary joints | 0 | 0 | 0 | U-shaped | 1.7321 | 1 | 0 | 0 | 0
A7-2b | Longitudinal position uncertainty (i.e. standing wave and imperfect field synthesis) for calibration antenna | 0.12 | 0.12 | 0.12 | Rectangular | 1.7321 | 1 | 0.069282 | 0.069282 | 0.069282
A7-4a | QZ ripple with calibration antenna | 0.2 | 0.2 | 0.2 | Rectangular | 1.7321 | 1 | 0.11547 | 0.11547 | 0.11547
A7-11 | Switching uncertainty | 0.02 | 0.02 | 0.02 | Rectangular | 1.7321 | 1 | 0.011547 | 0.011547 | 0.011547
A7-12 | Field repeatability | 0.06 | 0.12 | 0.12 | Gaussian | 1 | 1 | 0.06 | 0.12 | 0.12

## Reverb-Er
Reverberation Chamber - In-ban | [14 empty] | Review comments (RAN4#94)
|  | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz |  |  |  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz
Stage 2: DUT measurement
UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f<3 GHz | 3<f<4.2 GHz | 4.2<f<6 GHz |  |  |  | f<3 GHz | 3<f<4.2 GHz | 4.2<f<6 GHz
Stage 2: DUT measurement
C1-1 | Uncertainty of the measurement equipment | 0.14 | 0.26 | 0.26 | Gaussian | 1 | 1 | 0.14 | 0.26 | 0.26
A6-1 | Impedance mismatch in the receiving chain | 0.2 | 0.2 | 0.2 | U-shaped | 1.4142 | 1 | 0.14142 | 0.14142 | 0.14142 |  |  |  |  | RK: use 0.2 (Torbjorn by email)
A6-2 | Random uncertainty | 0.1 | 0.1 | 0.1 | Rectangular | 1.7321 | 1 | 0.057735 | 0.057735 | 0.057735
Stage 1: Calibration measurement
A6-3 | Reference antenna radiation efficiency | 0.2 | 0.2 | 0.2 | Gaussian | 1 | 1 | 0.2 | 0.2 | 0.2
A6-4 | Mean value estimation of reference antenna radiation efficiency | 0.15 | 0.15 | 0.15 | Gaussian | 1 | 1 | 0.15 | 0.15 | 0.15
C1-3 | Uncertainty of the Network Analyzer | 0.2 | 0.2 | 0.2 | Gaussian | 1 | 1 | 0.2 | 0.2 | 0.2
A6-5 | Influence of the reference antenna feed cable | 0.2 | 0.2 | 0.2 | Gaussian | 1 | 1 | 0.2 | 0.2 | 0.2
A6-6 | Mean value estimation of transfer function | 0.27 | 0.27 | 0.27 | Gaussian | 1 | 1 | 0.27 | 0.27 | 0.27
A6-7 | Uniformity of transfer function | 0.5 | 0.5 | 0.5 | Gaussian | 1 | 1 | 0.5 | 0.5 | 0.5
Reverberation Chamber -OOB
UID | Uncertainty source | Uncertainty value |  |  |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | 380 MHz < f ≦ 3 GHz | 3 GHz < f ≦ 3.8 GHz | 3.8 GHz < f ≦ 12.75 GHz | 12.75 GHz < f ≦ 19 GHz | 19GHz<f ≤26GHz |  |  |  | 380 MHz < f ≦ 3 GHz | 3 GHz < f ≦ 3.8 GHz | 3.8 GHz < f ≦ 12.75 GHz | 12.75 GHz < f ≦ 19 GHz | 19GHz<f ≤26GHz
Stage 2: DUT measurement
C1-1 | RF power measurement equipment (e.g. spectrum analyzer, power meter) | 0.14 | 0.26 | 0.26 | 0.37 | 0.37 | Gaussian | 1 | 1 | 0.14 | 0.26 | 0.26 | 0.37 | 0.37
A6-1 | Impedance mismatch in the receiving chain | 0.2 | 0.2 | 0.45 | 0.45 | 0.45 | U-shaped | 1.4142 | 1 | 0.28284 | 0.28284 | 0.6364 | 0.6364 | 0.6364
A6-2 | Random uncertainty | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | Rectangular | 1.7321 | 1 | 0.17321 | 0.17321 | 0.17321 | 0.17321 | 0.17321
Stage 1: Calibration measurement
A6-3 | Reference antenna radiation efficiency | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | Gaussian | 1 | 1 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5
A6-4 | Mean value estimation of reference antenna radiation efficiency | 0.27 | 0.27 | 0.27 | 0.27 | 0.27 | Gaussian | 1 | 1 | 0.27 | 0.27 | 0.27 | 0.27 | 0.27
C1-3 | Uncertainty of the network analyzer | 0.13 | 0.2 | 0.2 | Gaussian | 1 | 0.13 | 0.2 | 1 | 0.026 | 0.04 | 0.04 | #VALUE! | 0.2
A6-5 | Influence of the reference antenna feed cable | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | Gaussian | 1 | 1 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2
A6-6 | Mean value estimation of transfer function | 0.27 | 0.27 | 0.27 | 0.27 | 0.27 | Gaussian | 1 | 1 | 0.27 | 0.27 | 0.27 | 0.27 | 0.27
A6-7 | Uniformity of transfer function | 1.5 | 1.5 | 1.5 | 1.5 | 1.5 | Gaussian | 1 | 1 | 1.5 | 1.5 | 1.5 | 1.5 | 1.5

## Gen-Er
General Directional Chamber | [10 empty] | Review comments (RAN4#94)
UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | 30MHz<f≤6 GHz | 6GHz<f ≤19GHz | 19GHz<f ≤26GHz |  |  |  | 30MHz<f≤6 GHz | 6GHz<f ≤19GHz | 19GHz<f ≤26GHz
Stage 2: DUT measurement
A5-1 | Positioning misalignment between the AAS BS and the reference antenna | 0.03 | 0.03 | 0.03 | Rectangular | 1.7321 | 1 | 0.017321 | 0.017321 | 0.017321
A5-2 | Pointing misalignment between the AAS BS and the receiving antenna | 0 | 0 | 0 | Rectangular | 1.7321 | 1 | 0 | 0 | 0
A5-3 | Quality of quiet zone | 0.1 | 0.1 | 0.1 | Gaussian | 1 | 1 | 0.1 | 0.1 | 0.1
A5-4a | Polarization mismatch between the AAS BS and the receiving antenna | 0.01 | 0.01 | 0.01 | Rectangular | 1.7321 | 1 | 0.0057735 | 0.0057735 | 0.0057735
A5-5a | Mutual coupling between the AAS BS and the receiving antenna | 0 | 0 | 0 | Rectangular | 1.7321 | 1 | 0 | 0 | 0
A5-6a | Phase curvature | 0.05 | 0.05 | 0.05 | Gaussian | 1 | 1 | 0.05 | 0.05 | 0.05
C3-3 | Transmitter mandatory spurious emissions - Conducted Uncertainty (minus miss match) | 1 | 2 | 2 | Gaussian | 1 | 1 | 1 | 2 | 2
A5-7 | Impedance mismatch in the receiving chain | 0.2 | 0.45 | 0.45 | U-shaped | 1.4142 | 1 | 0.14142 | 0.3182 | 0.3182
A5-8 | Random uncertainty | 0.1 | 0.1 | 0.1 | Rectangular | 1.7321 | 1 | 0.057735 | 0.057735 | 0.057735
A5-17 | Measurement antenna frequency variation | 0.1 | 0.1 | 0.1 | Rectangular | 1.7321 | 1 | 0.057735 | 0.057735 | 0.057735
A5-18 | FSPL estimation error | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0
A5-16 | Test system frequency flatness | 0.25 | 0.25 | 0.25 | Gaussian | 1 | 1 | 0.25 | 0.25 | 0.25
A5-19 | Measurement system dynamic range uncertainty | 0.51 | 0.51 | 0.51 | Gaussian | 1 | 1 | 0.51 | 0.51 | 0.51
A5-20 | Reflections in anechoic chamber | 0.01 | 0.01 | 0.01 | Gaussian | 1 | 1 | 0.01 | 0.01 | 0.01
Stage 1: Calibration measurement
A5-9 | Impedance mismatch between the receiving antenna and the network analyzer | 0.05 | 0.05 | 0.05 | U-shaped | 1.4142 | 1 | 0.035355 | 0.035355 | 0.035355
A5-10 | Positioning and pointing misalignment between the reference antenna and the receiving antenna | 0.01 | 0.01 | 0.01 | Rectangular | 1.7321 | 1 | 0.0057735 | 0.0057735 | 0.0057735
A5-11 | Impedance mismatch between the reference antenna and the network analyzer. | 0.05 | 0.05 | 0.05 | U-shaped | 1.4142 | 1 | 0.035355 | 0.035355 | 0.035355
A5-3 | Quality of quiet zone | 0.1 | 0.1 | 0.1 | Gaussian | 1 | 1 | 0.1 | 0.1 | 0.1
A5-4b | Polarization mismatch for reference antenna | 0.01 | 0.01 | 0.01 | Rectangular | 1.7321 | 1 | 0.0057735 | 0.0057735 | 0.0057735
A5-5b | Mutual coupling between the reference antenna and the receiving antenna | 0 | 0 | 0 | Rectangular | 1.7321 | 1 | 0 | 0 | 0
A5-6b | Phase curvature | 0.05 | 0.05 | 0.05 | Gaussian | 1 | 1 | 0.05 | 0.05 | 0.05
C1-3 | Uncertainty of the network analyzer | 0.13 | 0.2 | 0.2 | Gaussian | 1 | 1 | 0.13 | 0.2 | 0.2
A5-12 | Influence of the reference antenna feed cable | 0.05 | 0.05 | 0.05 | Rectangular | 1.7321 | 1 | 0.028868 | 0.028868 | 0.028868
A5-13 | Reference antenna feed cable loss measurement uncertainty | 0.06 | 0.06 | 0.06 | Gaussian | 1 | 1 | 0.06 | 0.06 | 0.06
A5-14 | Influence of the receiving antenna feed cable | 0.05 | 0.05 | 0.05 | Rectangular | 1.7321 | 1 | 0.028868 | 0.028868 | 0.028868
C1-4 | Uncertainty of the absolute gain of the reference antenna | 0.50229 | 0.43301 | 0.43301 | Rectangular | 1.7321 | 1 | 0.29 | 0.25 | 0.25
A5-15 | Uncertainty of the absolute gain of the receiving antenna | 0 | 0 | 0 | Rectangular | 1.7321 | 1 | 0 | 0 | 0

## EIRP
Table 9.2.7-1: OTA test system specific measurement uncertainty values for the EIRP accuracy, FR1, Gaussian test conditions
|  | Expanded uncertainty [dB] | [11 empty] | Review comments (RAN4#94)
|  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz
| Indoor Anechoic Chamber | 0.86972 | 1.0574 | 1.0574
| Compact Antenna Test Range | 1.1143 | 1.2663 | 1.2663
| One Dimensional Compact Range Chamber | 0.90418 | 1.0952 | 1.0952
| Near Field Test Range | 1.0143 | 1.1041 | 1.1041
| PWS | 0.98073 | 1.1826 | 1.1826
| Common maximum accepted test system uncertainty | 1.1 | 1.3 | 1.3
Table 9.2.2.3-1: Indoor Anechoic Chamber measurement uncertainty value derivation for EIRP accuracy measurements, Gaussian test conditions, FR1
Indoor anechoic | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz |  |  |  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz
Stage 2: DUT measurement
A1-1 | Positioning misalignment between the AAS BS and the reference antenna | 0.03 | 0.03 | 0.03 | Rectangular | 1.7321 | 1 | 0.017321 | 0.017321 | 0.017321 | 0.0003 | 0.0003 | 0.0003
A1-2 | Pointing misalignment between the AAS BS and the receiving antenna | 0.3 | 0.3 | 0.3 | Rectangular | 1.7321 | 1 | 0.17321 | 0.17321 | 0.17321 | 0.03 | 0.03 | 0.03
A1-3 | Quality of quiet zone | 0.1 | 0.1 | 0.1 | Gaussian | 1 | 1 | 0.1 | 0.1 | 0.1 | 0.01 | 0.01 | 0.01
A1-4a | Polarization mismatch between the AAS BS and the receiving antenna | 0.01 | 0.01 | 0.01 | Rectangular | 1.7321 | 1 | 0.0057735 | 0.0057735 | 0.0057735 | 3.3333e-05 | 3.3333e-05 | 3.3333e-05
A1-5 | Mutual coupling between the AAS BS and the receiving antenna | 0 | 0 | 0 | Rectangular | 1.7321 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A1-6 | Phase curvature | 0.05 | 0.05 | 0.05 | Gaussian | 1 | 1 | 0.05 | 0.05 | 0.05 | 0.0025 | 0.0025 | 0.0025
C1-1 | RF power measurement equipment (e.g. spectrum analyzer, power meter) | 0.14 | 0.26 | 0.26 | Gaussian | 1 | 1 | 0.14 | 0.26 | 0.26 | 0.0196 | 0.0676 | 0.0676
A1-7 | Impedance mismatch in the receiving chain | 0.14 | 0.33 | 0.33 | U-shaped | 1.4142 | 1 | 0.098995 | 0.23335 | 0.23335 | 0.0098 | 0.05445 | 0.05445
A1-8 | Random uncertainty | 0.1 | 0.1 | 0.1 | Rectangular | 1.7321 | 1 | 0.057735 | 0.057735 | 0.057735 | 0.0033333 | 0.0033333 | 0.0033333
Stage 1: Calibration measurement | [10 empty] | 0 | 0 | 0
A1-9 | Impedance mismatch between the receiving antenna and the network analyzer | 0.05 | 0.05 | 0.05 | U-shaped | 1.4142 | 1 | 0.035355 | 0.035355 | 0.035355 | 0.00125 | 0.00125 | 0.00125
A1-10 | Positioning and pointing misalignment between the reference antenna and the receiving antenna | 0.01 | 0.01 | 0.01 | Rectangular | 1.7321 | 1 | 0.0057735 | 0.0057735 | 0.0057735 | 3.3333e-05 | 3.3333e-05 | 3.3333e-05
A1-11 | Impedance mismatch between the reference antenna and the network analyzer. | 0.05 | 0.05 | 0.05 | U-shaped | 1.4142 | 1 | 0.035355 | 0.035355 | 0.035355 | 0.00125 | 0.00125 | 0.00125
A1-3 | Quality of quiet zone | 0.1 | 0.1 | 0.1 | Gaussian | 1 | 1 | 0.1 | 0.1 | 0.1 | 0.01 | 0.01 | 0.01
A1-4b | Polarization mismatch between the reference antenna and the receiving antenna | 0.01 | 0.01 | 0.01 | Rectangular | 1.7321 | 1 | 0.0057735 | 0.0057735 | 0.0057735 | 3.3333e-05 | 3.3333e-05 | 3.3333e-05
A1-5 | Mutual coupling between the reference antenna and the receiving antenna | 0 | 0 | 0 | Rectangular | 1.7321 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A1-6 | Phase curvature | 0.05 | 0.05 | 0.05 | Gaussian | 1 | 1 | 0.05 | 0.05 | 0.05 | 0.0025 | 0.0025 | 0.0025
C1-3 | Uncertainty of the network analyzer | 0.13 | 0.2 | 0.2 | Gaussian | 1 | 1 | 0.13 | 0.2 | 0.2 | 0.0169 | 0.04 | 0.04
A1-12 | Influence of the reference antenna feed cable | 0.05 | 0.05 | 0.05 | Rectangular | 1.7321 | 1 | 0.028868 | 0.028868 | 0.028868 | 0.00083333 | 0.00083333 | 0.00083333
A1-13 | Reference antenna feed cable loss measurement uncertainty | 0.06 | 0.06 | 0.06 | Gaussian | 1 | 1 | 0.06 | 0.06 | 0.06 | 0.0036 | 0.0036 | 0.0036
A1-14 | Influence of the receiving antenna feed cable | 0.05 | 0.05 | 0.05 | Rectangular | 1.7321 | 1 | 0.028868 | 0.028868 | 0.028868 | 0.00083333 | 0.00083333 | 0.00083333
C1-4 | Uncertainty of the absolute gain of the reference antenna | 0.50229 | 0.43301 | 0.43301 | Rectangular | 1.7321 | 1 | 0.29 | 0.25 | 0.25 | 0.0841 | 0.0625 | 0.0625
A1-15 | Uncertainty of the absolute gain of the receiving antenna | 0 | 0 | 0 | Rectangular | 1.7321 | 1 | 0 | 0 | 0 | 0 | 0 | 0
Combined standard uncertainty (1σ) [dB] | [7 empty] | 0.44373 | 0.53949 | 0.53949 | 0.44373 | 0.53949 | 0.53949
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [7 empty] | 0.86972 | 1.0574 | 1.0574 | 0.86972 | 1.0574 | 1.0574
same as 37.842
Table 9.2.3.3-1: CATR MU value derivation for EIRP accuracy measurements, Gaussian test conditions, FR1
CATR | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz |  |  |  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz
Stage 2: DUT measurement
A2-1a | Misalignment DUT & pointing error for EIRP | 0 | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | RK:0 is used in 37.842 and 0.3 in 37.843
C1-1 | RF power measurement equipment (e.g. spectrum analyzer, power meter) | 0.14 | 0.26 | 0.26 | Gaussian | 1 | 1 | 0.14 | 0.26 | 0.26 | 0.0196 | 0.0676 | 0.0676
A2-2a | Standing wave between DUT and test range antenna | 0.21 | 0.21 | 0.21 | U-shaped | 1.4142 | 1 | 0.14849 | 0.14849 | 0.14849 | 0.02205 | 0.02205 | 0.02205
A2-3 | RF leakage (SGH connector terminated & test range antenna connector cable terminated) | 0.0012 | 0.0012 | 0.0012 | Gaussian | 1 | 1 | 0.0012 | 0.0012 | 0.0012 | 1.44e-06 | 1.44e-06 | 1.44e-06
A2-4a | QZ ripple DUT | 0.0928 | 0.0928 | 0.0928 | Gaussian | 1 | 1 | 0.0928 | 0.0928 | 0.0928 | 0.0086118 | 0.0086118 | 0.0086118
A2-12 | Frequency flatness | 0.25 | 0.25 | 0.25 | Gaussian | 1 | 1 | 0.25 | 0.25 | 0.25 | 0.0625 | 0.0625 | 0.0625 | RK:Not in 37.842, added to be consistant with 37.843 and other methods.
Stage 1: Calibration measurement | [10 empty] | 0 | 0 | 0
C1-3 | Uncertainty of the network analyzer | 0.13 | 0.2 | 0.2 | Gaussian | 1 | 1 | 0.13 | 0.2 | 0.2 | 0.0169 | 0.04 | 0.04
A2-6 | Mismatch of receiver chain | 0.127 | 0.325 | 0.325 | U-shaped | 1.4142 | 1 | 0.089803 | 0.22981 | 0.22981 | 0.0080645 | 0.052812 | 0.052812
A2-3 | Insertion loss variation of receiver chain | 0.18 | 0.18 | 0.18 | Rectangular | 1.7321 | 1 | 0.10392 | 0.10392 | 0.10392 | 0.0108 | 0.0108 | 0.0108
A2-3 | RF leakage (SGH connector terminated & test range antenna connector cable terminated) | 0.0012 | 0.0012 | 0.0012 | Gaussian | 1 | 1 | 0.0012 | 0.0012 | 0.0012 | 1.44e-06 | 1.44e-06 | 1.44e-06
C1-4 | Influence of the calibration antenna feed cable: | 0.022 | 0.022 | 0.022 | U-shaped | 1.4142 | 1 | 0.015556 | 0.015556 | 0.015556 | 0.000242 | 0.000242 | 0.000242
C1-4 | Uncertainty of the absolute gain of the reference antenna | 0.50229 | 0.43301 | 0.43301 | Rectangular | 1.7321 | 1 | 0.29 | 0.25 | 0.25 | 0.0841 | 0.0625 | 0.0625
A2-1b | Misalignment positioning system | 0 | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A2-9 | Misalignment of calibration antenna and test range antenna | 0.5 | 0.5 | 0.5 | Exp. normal | 2 | 1 | 0.25 | 0.25 | 0.25 | 0.0625 | 0.0625 | 0.0625
A2-2b | Rotary Joints | 0.048 | 0.048 | 0.048 | U-shaped | 1.4142 | 1 | 0.033941 | 0.033941 | 0.033941 | 0.001152 | 0.001152 | 0.001152
A2-4b | Standing wave between calibration antenna and test range antenna | 0.09 | 0.09 | 0.09 | U-shaped | 1.4142 | 1 | 0.06364 | 0.06364 | 0.06364 | 0.00405 | 0.00405 | 0.00405
A2-11 | QZ ripple calibration antenna | 0.009 | 0.009 | 0.009 | Gaussian | 1 | 1 | 0.009 | 0.009 | 0.009 | 8.1e-05 | 8.1e-05 | 8.1e-05
A2-13 | Switching uncertainty | 0.26 | 0.26 | 0.26 | Rectangular | 1.7321 | 1 | 0.15011 | 0.15011 | 0.15011 | 0.022533 | 0.022533 | 0.022533
Combined standard uncertainty (1σ) [dB] | [7 empty] | 0.5685 | 0.64609 | 0.64609 | 0.5685 | 0.64609 | 0.64609
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [7 empty] | 1.1143 | 1.2663 | 1.2663 | 1.1143 | 1.2663 | 1.2663
Total is a bit higjher than 37.842 as FF has been added (pointing error may also increase this a little)
Table 9.2.5.3-1: NFTR measurement uncertainty value derivation for EIRP accuracy measurements, FR1
Near field | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz |  |  |  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz
Stage 2: DUT measurement
A3-1 | Axes Intersection | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-2 | Axes Orthogonality | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-3 | Horizontal Pointing | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-4 | Probe Vertical Position | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-5 | Probe H/V pointing | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-6 | Measurement Distance | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-7 | Amplitude and Phase Drift | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-8 | Amplitude and Phase Noise | 0.02 | 0.02 | 0.02 | Gaussian | 1 | 1 | 0.02 | 0.02 | 0.02 | 0.0004 | 0.0004 | 0.0004
A3-9 | Leakage and Crosstalk | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-10 | Amplitude Non-Linearity | 0.04 | 0.04 | 0.04 | Gaussian | 1 | 1 | 0.04 | 0.04 | 0.04 | 0.0016 | 0.0016 | 0.0016
A3-11 | Amplitude and Phase Shift in rotary joints | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-12 | Channel Balance Amplitude and Phase | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-13 | Probe Polarization Amplitude and Phase | 0.0001 | 0.0001 | 0.0001 | Gaussian | 1 | 1 | 0.0001 | 0.0001 | 0.0001 | 1e-08 | 1e-08 | 1e-08
A3-14 | Probe Pattern Knowledge | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-15 | Multiple Reflections | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-16 | Room Scattering | 0.09 | 0.09 | 0.09 | Gaussian | 1 | 1 | 0.09 | 0.09 | 0.09 | 0.0081 | 0.0081 | 0.0081
A3-17 | DUT support Scattering | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-18 | Scan Area Truncation | 0.003 | 0.003 | 0.003 | Gaussian | 1 | 1 | 0.003 | 0.003 | 0.003 | 9e-06 | 9e-06 | 9e-06
A3-19 | Sampling Point Offset | 0.0058 | 0.0058 | 0.0058 | Gaussian | 1 | 1 | 0.0058 | 0.0058 | 0.0058 | 3.364e-05 | 3.364e-05 | 3.364e-05
A3-20 | Spherical Mode Truncation | 0.015 | 0.015 | 0.015 | Gaussian | 1 | 1 | 0.015 | 0.015 | 0.015 | 0.000225 | 0.000225 | 0.000225
A3-21 | Positioning | 0.03 | 0.03 | 0.03 | Rectangular | 1.73 | 1 | 0.017341 | 0.017341 | 0.017341 | 0.00030071 | 0.00030071 | 0.00030071
A3-22 | Probe Array Uniformity | 0.055 | 0.055 | 0.055 | Gaussian | 1 | 1 | 0.055 | 0.055 | 0.055 | 0.003025 | 0.003025 | 0.003025
A3-23 | Mismatch of receiver chain | 0.284 | 0.284 | 0.284 | U-Shaped | 1.41 | 1 | 0.20142 | 0.20142 | 0.20142 | 0.040569 | 0.040569 | 0.040569
A3-24 | Insertion loss of receiver chain | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-25 | Uncertainty of the absolute gain of the probe antenna | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
C1-1 | RF power measurement equipment (e.g. spectrum analyzer, power meter) | 0.14 | 0.26 | 0.26 | Gaussian | 1 | 1 | 0.14 | 0.26 | 0.26 | 0.0196 | 0.0676 | 0.0676
A3-26 | Measurement repeatability - positioning repeatability | 0.15 | 0.15 | 0.15 | Gaussian | 1 | 1 | 0.15 | 0.15 | 0.15 | 0.0225 | 0.0225 | 0.0225
A3-33 | Test system frequency flatness | 0.25 | 0.25 | 0.25 | Gaussian | 1 | 1 | 0.25 | 0.25 | 0.25 | 0.0625 | 0.0625 | 0.0625 | RK: added to be consistant with TRP anmd other methods
Stage 1: Calibration measurement | [10 empty] | 0 | 0 | 0
C1-3 | Uncertainty of the network analyzer | 0.13 | 0.2 | 0.2 | Gaussian | 1 | 1 | 0.13 | 0.2 | 0.2 | 0.0169 | 0.04 | 0.04
A3-27 | Mismatch of receiver chain | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-28 | Insertion loss of receiver chain | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-29 | Mismatch in the connection of the calibration antenna | 0.02 | 0.02 | 0.02 | U-Shaped | 1.41 | 1 | 0.014184 | 0.014184 | 0.014184 | 0.0002012 | 0.0002012 | 0.0002012
A3-30 | Influence of the calibration antenna feed cable | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-31 | Influence of the probe antenna cable | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
C1-4 | Uncertainty of the absolute gain of the reference antenna | 0.50229 | 0.43301 | 0.43301 | Rectangular | 1.7321 | 1 | 0.29 | 0.25 | 0.25 | 0.0841 | 0.0625 | 0.0625
A3-32 | Short term repeatability | 0.088 | 0.088 | 0.088 | Gaussian | 1 | 1 | 0.088 | 0.088 | 0.088 | 0.007744 | 0.007744 | 0.007744
Combined standard uncertainty (1σ) [dB] | [7 empty] | 0.5175 | 0.5633 | 0.5633 | 0.5175 | 0.5633 | 0.5633
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [7 empty] | 1.0143 | 1.1041 | 1.1041 | 1.0143 | 1.1041 | 1.1041
value is slightly higher than 37.842 due to addition of FF
Table 9.2.4.3-1: One Dimensional Compact Range MU value derivation for EIRP accuracy measurements, FR1
One Dimensional Compact Range Chamber | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz |  |  |  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz
Stage 2: DUT measurement
A4-1 | Misalignment DUT and pointing error | 0 | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A4-2a | Standing wave between DUT and test range antenna | 0.18 | 0.18 | 0.18 | U-shaped | 1.4142 | 1 | 0.12728 | 0.12728 | 0.12728 | 0.0162 | 0.0162 | 0.0162
A4-3a | Quiet zone ripple DUT | 0.0325 | 0.0325 | 0.0325 | Gaussian | 1 | 1 | 0.0325 | 0.0325 | 0.0325 | 0.0010563 | 0.0010563 | 0.0010563
A4-4a | Phase curvature AAS | 0.01 | 0.01 | 0.01 | Gaussian | 1 | 1 | 0.01 | 0.01 | 0.01 | 0.0001 | 0.0001 | 0.0001
A4-5a | Polarization mismatch between DUT and receiving antenna | 0.05 | 0.05 | 0.05 | Rectangular | 1.7321 | 1 | 0.028868 | 0.028868 | 0.028868 | 0.00083333 | 0.00083333 | 0.00083333
A4-6a | Mutual coupling between DUT and receiving antenna | 0 | 0 | 0 | Rectangular | 1.7321 | 1 | 0 | 0 | 0 | 0 | 0 | 0
C1-1 | RF power measurement equipment (e.g. spectrum analyzer, power meter) | 0.14 | 0.26 | 0.26 | Gaussian | 1 | 1 | 0.14 | 0.26 | 0.26 | 0.0196 | 0.0676 | 0.0676
A4-7 | Impedance mismatch in receiving chain | 0.0056 | 0.01 | 0.01 | U-shaped | 1.4142 | 1 | 0.0039598 | 0.0070711 | 0.0070711 | 1.568e-05 | 5e-05 | 5e-05
A4-8a | RF leakage (DUT connector terminated and test range antenna connector cable terminated) | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
Stage 1: Calibration measurement | [10 empty] | 0 | 0 | 0
A4-9 | Misalignment positioning system | 0 | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A4-10 | Pointing error between reference antenna and test range antenna | 0 | 0 | 0 | Rectangular | 1.7321 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A4-11 | Impedance mismatch in path to reference antenna | 0.05 | 0.05 | 0.05 | U-shaped | 1.4142 | 1 | 0.035355 | 0.035355 | 0.035355 | 0.00125 | 0.00125 | 0.00125
A4-12 | Impedance mismatch in path to compact probe | 0.03 | 0.03 | 0.03 | U-shaped | 1.4142 | 1 | 0.021213 | 0.021213 | 0.021213 | 0.00045 | 0.00045 | 0.00045
A4-2b | Standing wave between reference antenna and receiving antenna | 0.09 | 0.09 | 0.09 | U-shaped | 1.4142 | 1 | 0.06364 | 0.06364 | 0.06364 | 0.00405 | 0.00405 | 0.00405
A4-3b | Quiet zone ripple reference antenna | 0.178 | 0.178 | 0.178 | Gaussian | 1 | 1 | 0.178 | 0.178 | 0.178 | 0.031684 | 0.031684 | 0.031684
A4-4b | Phase curvature refernce antenna | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A4-5b | Polarization mismatch between reference antenna and receiving antenna | 0.05 | 0.05 | 0.05 | Rectangular | 1.7321 | 1 | 0.028868 | 0.028868 | 0.028868 | 0.00083333 | 0.00083333 | 0.00083333
A4-6a | Mutual coupling between reference antenna and receiving antenna | 0 | 0 | 0 | Rectangular | 1.7321 | 1 | 0 | 0 | 0 | 0 | 0 | 0
C1-1 | RF power measurement equipment (e.g. spectrum analyzer, power meter) | 0.14 | 0.26 | 0.26 | Gaussian | 1 | 1 | 0.14 | 0.26 | 0.26 | 0.0196 | 0.0676 | 0.0676
A4-13 | Influence of the reference antenna feed cable (flexing cables, adapters, attenuators, connector repeatability) | 0.082 | 0.082 | 0.082 | Rectangular | 1.7321 | 1 | 0.047343 | 0.047343 | 0.047343 | 0.0022413 | 0.0022413 | 0.0022413
A4-14 | Mismatch of receiver chain | 0.2 | 0.3 | 0.3 | U-shaped | 1.4142 | 1 | 0.14142 | 0.21213 | 0.21213 | 0.02 | 0.045 | 0.045
A4-15 | Insertion loss of receiver chain | 0.18 | 0.18 | 0.18 | Rectangular | 1.7321 | 1 | 0.10392 | 0.10392 | 0.10392 | 0.0108 | 0.0108 | 0.0108
C1-4 | Uncertainty of the absolute gain of the reference antenna | 0.50229 | 0.43301 | 0.43301 | Rectangular | 1.7321 | 1 | 0.29 | 0.25 | 0.25 | 0.0841 | 0.0625 | 0.0625
A4-8b | RF leakage (SGH connector terminated and test range antenna connector cable terminated. | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
Combined standard uncertainty (1σ) [dB] | [7 empty] | 0.46132 | 0.55879 | 0.55879 | 0.46132 | 0.55879 | 0.55879
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [7 empty] | 0.90418 | 1.0952 | 1.0952 | 0.90418 | 1.0952 | 1.0952
same as 37.842 should we add FF?
Table 9.2.6.3-1: PWS measurement uncertainty value derivation for EIRP accuracy measurements, FR1
Plane wave synthesiser | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz |  |  |  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz
Stage 2: DUT measurement
A7-1a | Misalignment DUT & pointing error | 0.1 | 0.1 | 0.1 | Rectangular | 1.7321 | 1 | 0.057735 | 0.057735 | 0.057735 | 0.0033333 | 0.0033333 | 0.0033333
C1-1 | RF power measurement equipment (e.g. spectrum analyzer, power meter) | 0.14 | 0.26 | 0.26 | Gaussian | 1 | 1 | 0.14 | 0.26 | 0.26 | 0.0196 | 0.0676 | 0.0676
A7-2a | Longitudinal position uncertainty (i.e. standing wave and imperfect field synthesis) for DUT antenna | 0.05 | 0.14 | 0.14 | Rectangular | 1.7321 | 1 | 0.028868 | 0.080829 | 0.080829 | 0.00083333 | 0.0065333 | 0.0065333
A7-3 | RF leakage (calibration antenna connector terminated) | 0.086 | 0.086 | 0.086 | Gaussian | 1 | 1 | 0.086 | 0.086 | 0.086 | 0.007396 | 0.007396 | 0.007396
A7-4a | QZ ripple with DUT | 0.42 | 0.43 | 0.43 | Rectangular | 1.7321 | 1 | 0.24249 | 0.24826 | 0.24826 | 0.0588 | 0.061633 | 0.061633
A7-5 | Miscellaneous Uncertainty | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A7-14 | System non-linearity | 0.062 | 0.062 | 0.062 | Rectangular | 1.7321 | 1 | 0.035796 | 0.035796 | 0.035796 | 0.0012813 | 0.0012813 | 0.0012813
A7-13 | Frequency Flatness | 0.13 | 0.13 | 0.13 | Rectangular | 1.7321 | 1 | 0.075056 | 0.075056 | 0.075056 | 0.0056333 | 0.0056333 | 0.0056333
Stage 1: Calibration measurement | [10 empty] | 0 | 0 | 0
C1-3 | Uncertainty of the network analyzer | 0.13 | 0.2 | 0.2 | Gaussian | 1 | 1 | 0.13 | 0.2 | 0.2 | 0.0169 | 0.04 | 0.04
A7-6 | Mismatch (i.e. reference antenna, network analyser and reference cable) | 0.127 | 0.325 | 0.325 | U-shaped | 1.4142 | 1 | 0.089803 | 0.22981 | 0.22981 | 0.0080645 | 0.052812 | 0.052812
A7-7 | Insertion loss variation | 0.18 | 0.18 | 0.18 | Rectangular | 1.7321 | 1 | 0.10392 | 0.10392 | 0.10392 | 0.0108 | 0.0108 | 0.0108
A7-3 | RF leakage (calibration antenna connector terminated) | 0.086 | 0.086 | 0.086 | Gaussian | 1 | 1 | 0.086 | 0.086 | 0.086 | 0.007396 | 0.007396 | 0.007396
A7-8 | Influence of the calibration antenna feed cable | 0.103 | 0.104 | 0.104 | Rectangular | 1.7321 | 1 | 0.059467 | 0.060044 | 0.060044 | 0.0035363 | 0.0036053 | 0.0036053
C1-4 | Uncertainty of the absolute gain of the reference antenna | 0.50229 | 0.43301 | 0.43301 | Rectangular | 1.7321 | 1 | 0.29 | 0.25 | 0.25 | 0.0841 | 0.0625 | 0.0625
A7-9 | Misalignment of positioning system | 0 | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A7-1b | Misalignment of calibration antenna & pointing error | 0.05 | 0.05 | 0.05 | Rectangular | 1.7321 | 1 | 0.028868 | 0.028868 | 0.028868 | 0.00083333 | 0.00083333 | 0.00083333
A7-10 | Rotary joints | 0 | 0 | 0 | U-shaped | 1.7321 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A7-2b | Longitudinal position uncertainty (i.e. standing wave and imperfect field synthesis) for calibration antenna | 0.12 | 0.12 | 0.12 | Rectangular | 1.7321 | 1 | 0.069282 | 0.069282 | 0.069282 | 0.0048 | 0.0048 | 0.0048
A7-4a | QZ ripple with calibration antenna | 0.2 | 0.2 | 0.2 | Rectangular | 1.7321 | 1 | 0.11547 | 0.11547 | 0.11547 | 0.013333 | 0.013333 | 0.013333
A7-11 | Switching uncertainty | 0.02 | 0.02 | 0.02 | Rectangular | 1.7321 | 1 | 0.011547 | 0.011547 | 0.011547 | 0.00013333 | 0.00013333 | 0.00013333
A7-12 | Field repeatability | 0.06 | 0.12 | 0.12 | Gaussian | 1 | 1 | 0.06 | 0.12 | 0.12 | 0.0036 | 0.0144 | 0.0144
Combined standard uncertainty (1σ) [dB] | [7 empty] | 0.50037 | 0.60334 | 0.60334 | 0.50037 | 0.60334 | 0.60334
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [7 empty] | 0.98073 | 1.1826 | 1.1826 | 0.98073 | 1.1826 | 1.1826
As per R& latest figures (28/1/20)
Checked 28/1/20
PWS corrections from R&S
Freq flaness added to CATR
Issues to be conformed:
-IAC/1D, should we add FF
-CATR pointing error (0 or 0.3)

## EIRP-Ex
Table 9.3.4-1: Test system specific measurement uncertainty values for the EIRP accuracy in Extreme test conditions
| chamber | Expanded uncertainty [dB] | [11 empty] | Review comments (RAN4#94)
|  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz
| Indoor anechoic | 2.46 | 2.5324 | 2.5324
| CATR | 2.5103 | 2.5814 | 2.5515
Near field
1D CATR
Plane wave sythesiser
| Common maximum accepted test system uncertainty | 2.5 | 2.6 | 2.6
Table 9.3.2.3-1: Indoor Anechoic Chamber measurement uncertainty value derivation for EIRP accuracy measurements in Extreme test conditions
Indoor anechoic | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz |  |  |  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz
Stage 2: DUT measurement
A1-1 | Positioning misalignment between the AAS BS and the reference antenna | 0.03 | 0.03 | 0.03 | Rectangular | 1.7321 | 1 | 0.017321 | 0.017321 | 0.017321 | 0.0003 | 0.0003 | 0.0003
A1-2 | Pointing misalignment between the AAS BS and the receiving antenna | 0.3 | 0.3 | 0.3 | Rectangular | 1.7321 | 1 | 0.17321 | 0.17321 | 0.17321 | 0.03 | 0.03 | 0.03
A1-17 | Quality of quiet zone (extreme) | 0.6 | 0.6 | 0.6 | Gaussian | 1 | 1 | 0.6 | 0.6 | 0.6 | 0.36 | 0.36 | 0.36
A1-4a | Polarization mismatch between the AAS BS and the receiving antenna | 0.01 | 0.01 | 0.01 | Rectangular | 1.7321 | 1 | 0.0057735 | 0.0057735 | 0.0057735 | 3.3333e-05 | 3.3333e-05 | 3.3333e-05
A1-5 | Mutual coupling between the AAS BS and the receiving antenna | 0 | 0 | 0 | Rectangular | 1.7321 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A1-6 | Phase curvature | 0.05 | 0.05 | 0.05 | Gaussian | 1 | 1 | 0.05 | 0.05 | 0.05 | 0.0025 | 0.0025 | 0.0025
C1-1 | RF power measurement equipment (e.g. spectrum analyzer, power meter) | 0.14 | 0.26 | 0.26 | Gaussian | 1 | 1 | 0.14 | 0.26 | 0.26 | 0.0196 | 0.0676 | 0.0676
A1-7 | Impedance mismatch in the receiving chain | 0.14 | 0.33 | 0.33 | U-shaped | 1.4142 | 1 | 0.098995 | 0.23335 | 0.23335 | 0.0098 | 0.05445 | 0.05445
A1-8 | Random uncertainty | 0.1 | 0.1 | 0.1 | Rectangular | 1.7321 | 1 | 0.057735 | 0.057735 | 0.057735 | 0.0033333 | 0.0033333 | 0.0033333
A1-19 | Radome loss variation | 0.4 | 0.4 | 0.4 | Rectangular | 1.7321 | 1 | 0.23094 | 0.23094 | 0.23094 | 0.053333 | 0.053333 | 0.053333
A1-18 | Wet radome loss variation | 0.95 | 0.95 | 0.95 | Gaussian | 1 | 1 | 0.95 | 0.95 | 0.95 | 0.9025 | 0.9025 | 0.9025
A1-20 | Change in absorber behaviour | 0.1 | 0.1 | 0.1 | Gaussian | 1 | 1 | 0.1 | 0.1 | 0.1 | 0.01 | 0.01 | 0.01
A1-16 | Frequency flatness | 0.25 | 0.25 | 0.25 | Gaussian | 1 | 1 | 0.25 | 0.25 | 0.25 | 0.0625 | 0.0625 | 0.0625
Stage 1: Calibration measurement | [10 empty] | 0 | 0 | 0
A1-9 | Impedance mismatch between the receiving antenna and the network analyzer | 0.05 | 0.05 | 0.05 | U-shaped | 1.4142 | 1 | 0.035355 | 0.035355 | 0.035355 | 0.00125 | 0.00125 | 0.00125
A1-10 | Positioning and pointing misalignment between the reference antenna and the receiving antenna | 0.01 | 0.01 | 0.01 | Rectangular | 1.7321 | 1 | 0.0057735 | 0.0057735 | 0.0057735 | 3.3333e-05 | 3.3333e-05 | 3.3333e-05
A1-11 | Impedance mismatch between the reference antenna and the network analyzer. | 0.05 | 0.05 | 0.05 | U-shaped | 1.4142 | 1 | 0.035355 | 0.035355 | 0.035355 | 0.00125 | 0.00125 | 0.00125
A1-3 | Quality of quiet zone | 0.1 | 0.1 | 0.1 | Gaussian | 1 | 1 | 0.1 | 0.1 | 0.1 | 0.01 | 0.01 | 0.01
A1-4b | Polarization mismatch between the reference antenna and the receiving antenna | 0.01 | 0.01 | 0.01 | Rectangular | 1.7321 | 1 | 0.0057735 | 0.0057735 | 0.0057735 | 3.3333e-05 | 3.3333e-05 | 3.3333e-05
A1-5 | Mutual coupling between the reference antenna and the receiving antenna | 0 | 0 | 0 | Rectangular | 1.7321 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A1-6 | Phase curvature | 0.05 | 0.05 | 0.05 | Gaussian | 1 | 1 | 0.05 | 0.05 | 0.05 | 0.0025 | 0.0025 | 0.0025
C1-3 | Uncertainty of the network analyzer | 0.13 | 0.2 | 0.2 | Gaussian | 1 | 1 | 0.13 | 0.2 | 0.2 | 0.0169 | 0.04 | 0.04
A1-12 | Influence of the reference antenna feed cable | 0.05 | 0.05 | 0.05 | Rectangular | 1.7321 | 1 | 0.028868 | 0.028868 | 0.028868 | 0.00083333 | 0.00083333 | 0.00083333
A1-13 | Reference antenna feed cable loss measurement uncertainty | 0.06 | 0.06 | 0.06 | Gaussian | 1 | 1 | 0.06 | 0.06 | 0.06 | 0.0036 | 0.0036 | 0.0036
A1-14 | Influence of the receiving antenna feed cable | 0.05 | 0.05 | 0.05 | Rectangular | 1.7321 | 1 | 0.028868 | 0.028868 | 0.028868 | 0.00083333 | 0.00083333 | 0.00083333
C1-4 | Uncertainty of the absolute gain of the reference antenna | 0.50229 | 0.43301 | 0.43301 | Rectangular | 1.7321 | 1 | 0.29 | 0.25 | 0.25 | 0.0841 | 0.0625 | 0.0625
A1-15 | Uncertainty of the absolute gain of the receiving antenna | 0 | 0 | 0 | Rectangular | 1.7321 | 1 | 0 | 0 | 0 | 0 | 0 | 0
Combined standard uncertainty (1σ) [dB] | [7 empty] | 1.2551 | 1.292 | 1.292 | 1.2551 | 1.292 | 1.292
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [7 empty] | 2.46 | 2.5324 | 2.5324 | 2.46 | 2.5324 | 2.5324
same as 37.842
Table 9.3.3.4-1: CATR MU value derivation for EIRP accuracy measurements in Extreme test conditions
CATR | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz |  |  |  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz
Stage 2: DUT measurement
A2-1a | Misalignment DUT & pointing error for EIRP | 0 | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 is used in 37.842 and 0.3 in 37.843
C1-1 | RF power measurement equipment (e.g. spectrum analyzer, power meter) | 0.14 | 0.26 | 0.26 | Gaussian | 1 | 1 | 0.14 | 0.26 | 0.26 | 0.0196 | 0.0676 | 0.0676
A2-2a | Standing wave between DUT and test range antenna | 0.21 | 0.21 | 0.21 | U-shaped | 1.4142 | 1 | 0.14849 | 0.14849 | 0.14849 | 0.02205 | 0.02205 | 0.02205
A2-3 | RF leakage (SGH connector terminated & test range antenna connector cable terminated) | 0.0012 | 0.0012 | 0.0012 | Gaussian | 1 | 1 | 0.0012 | 0.0012 | 0.0012 | 1.44e-06 | 1.44e-06 | 1.44e-06
A2-13 | Quality of quiet zone (extreme) | 0.6 | 0.6 | 0.6 | Gaussian | 1 | 1 | 0.6 | 0.6 | 0.6 | 0.36 | 0.36 | 0.36
A2-12 | Frequency flatness | 0.25 | 0.25 | 0.25 | Gaussian | 1 | 1 | 0.25 | 0.25 | 0.25 | 0.0625 | 0.0625 | 0.0625 | Not in 37.8942, added to be consistant with 37.843 and other methods.
A2-14 | Wet radome loss variation | 0.4 | 0.4 | 0.4 | Rectangular | 1.7321 | 1 | 0.23094 | 0.23094 | 0.23094 | 0.053333 | 0.053333 | 0.053333
A2-15 | Radome loss variation | 0.95 | 0.95 | 0.95 | Gaussian | 1 | 1 | 0.95 | 0.95 | 0.95 | 0.9025 | 0.9025 | 0.9025
A2-16 | Change in absorber behaviour | 0.1 | 0.1 | 0.1 | Gaussian | 1 | 1 | 0.1 | 0.1 | 0.1 | 0.01 | 0.01 | 0.01
Stage 1: Calibration measurement | [10 empty] | 0 | 0 | 0
C1-3 | Uncertainty of the network analyzer | 0.13 | 0.2 | 0.2 | Gaussian | 1 | 1 | 0.13 | 0.2 | 0 | 0.0169 | 0.04 | 0
A2-6 | Mismatch of receiver chain | 0.127 | 0.325 | 0.325 | U-shaped | 1.4142 | 1 | 0.089803 | 0.22981 | 0.22981 | 0.0080645 | 0.052812 | 0.052812
A2-3 | Insertion loss variation of receiver chain | 0.18 | 0.18 | 0.18 | Rectangular | 1.7321 | 1 | 0.10392 | 0.10392 | 0.10392 | 0.0108 | 0.0108 | 0.0108
A2-3 | RF leakage (SGH connector terminated & test range antenna connector cable terminated) | 0.0012 | 0.0012 | 0.0012 | Gaussian | 1 | 1 | 0.0012 | 0.0012 | 0.0012 | 1.44e-06 | 1.44e-06 | 1.44e-06
C1-4 | Influence of the calibration antenna feed cable: | 0.022 | 0.022 | 0.022 | U-shaped | 1.4142 | 1 | 0.015556 | 0.015556 | 0.015556 | 0.000242 | 0.000242 | 0.000242
C1-4 | Uncertainty of the absolute gain of the reference antenna | 0.50229 | 0.43301 | 0.43301 | Rectangular | 1.7321 | 1 | 0.29 | 0.25 | 0.25 | 0.0841 | 0.0625 | 0.0625
A2-1b | Misalignment positioning system | 0 | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A2-9 | Misalignment of calibration antenna and test range antenna | 0.5 | 0.5 | 0.5 | Exp. normal | 2 | 1 | 0.25 | 0.25 | 0.25 | 0.0625 | 0.0625 | 0.0625
A2-2b | Rotary Joints | 0.048 | 0.048 | 0.048 | U-shaped | 1.4142 | 1 | 0.033941 | 0.033941 | 0.033941 | 0.001152 | 0.001152 | 0.001152
A2-4b | Standing wave between calibration antenna and test range antenna | 0.09 | 0.09 | 0.09 | U-shaped | 1.4142 | 1 | 0.06364 | 0.06364 | 0.06364 | 0.00405 | 0.00405 | 0.00405
A2-11 | QZ ripple calibration antenna | 0.009 | 0.009 | 0.009 | Gaussian | 1 | 1 | 0.009 | 0.009 | 0.009 | 8.1e-05 | 8.1e-05 | 8.1e-05
A2-13 | Switching uncertainty | 0.26 | 0.26 | 0.26 | Rectangular | 1.7321 | 1 | 0.15011 | 0.15011 | 0.15011 | 0.022533 | 0.022533 | 0.022533
Combined standard uncertainty (1σ) [dB] | [7 empty] | 1.2808 | 1.3171 | 1.3018 | 1.2808 | 1.3171 | 1.3018
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [7 empty] | 2.5103 | 2.5814 | 2.5515 | 2.5103 | 2.5814 | 2.5515
Same as 37/843
Near field | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz |  |  |  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz
Stage 2: DUT measurement
Stage 1: Calibration measurement | [10 empty] | 0 | 0 | 0
One Dimensional Compact Range Chamber | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz |  |  |  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz
Stage 2: DUT measurement
Stage 1: Calibration measurement | [10 empty] | 0 | 0 | 0
Plane wave synthesiser | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz |  |  |  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz
Stage 2: DUT measurement
Stage 1: Calibration measurement | [10 empty] | 0 | 0 | 0
Checked 30/1/20
Same as TR 37.843

## DL RS
Table 6.4.5-1: Test system specific measurement uncertainty values for the OTA E-UTRA DL RS power test
| chamber | Expanded uncertainty [dB] | [11 empty] | Review comments (RAN4#94)
|  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz
| Indoor Anechoic Chamber | 1.1519 | 1.4364 | 1.4364
| Compact Antenna Test Range | 1.3461 | 1.5965 | 1.5965
| One Dimensional Compact Range Chamber | 1.1658 | 1.387 | 1.387
Near Field Test Range
PWS
| Common maximum accepted test system uncertainty | 1.3 | 1.5 | 1.5
Table 69.4.2.4-1: Indoor Anechoic Chamber measurement uncertainty value derivation for OTA E-UTRA DL RS power measurement
Indoor anechoic | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz |  |  |  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz
Stage 2: DUT measurement
A1-1 | Positioning misalignment between the AAS BS and the reference antenna | 0.03 | 0.03 | 0.03 | Rectangular | 1.7321 | 1 | 0.017321 | 0.017321 | 0.017321 | 0.0003 | 0.0003 | 0.0003
A1-2 | Pointing misalignment between the AAS BS and the receiving antenna | 0.3 | 0.3 | 0.3 | Rectangular | 1.7321 | 1 | 0.17321 | 0.17321 | 0.17321 | 0.03 | 0.03 | 0.03
A1-3 | Quality of quiet zone | 0.1 | 0.1 | 0.1 | Gaussian | 1 | 1 | 0.1 | 0.1 | 0.1 | 0.01 | 0.01 | 0.01
A1-4a | Polarization mismatch between the AAS BS and the receiving antenna | 0.01 | 0.01 | 0.01 | Rectangular | 1.7321 | 1 | 0.0057735 | 0.0057735 | 0.0057735 | 3.3333e-05 | 3.3333e-05 | 3.3333e-05
A1-5 | Mutual coupling between the AAS BS and the receiving antenna | 0 | 0 | 0 | Rectangular | 1.7321 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A1-6 | Phase curvature | 0.05 | 0.05 | 0.05 | Gaussian | 1 | 1 | 0.05 | 0.05 | 0.05 | 0.0025 | 0.0025 | 0.0025
C3-1 | DL-RS MU derived from conducted spec | 0.41 | 0.56 | 0.56 | Gaussian | 1 | 1 | 0.41 | 0.56 | 0.56 | 0.1681 | 0.3136 | 0.3136
A1-7 | Impedance mismatch in the receiving chain | 0.14 | 0.33 | 0.33 | U-shaped | 1.4142 | 1 | 0.098995 | 0.23335 | 0.23335 | 0.0098 | 0.05445 | 0.05445
A1-8 | Random uncertainty | 0.1 | 0.1 | 0.1 | Rectangular | 1.7321 | 1 | 0.057735 | 0.057735 | 0.057735 | 0.0033333 | 0.0033333 | 0.0033333
Stage 1: Calibration measurement | [10 empty] | 0 | 0 | 0
A1-9 | Impedance mismatch between the receiving antenna and the network analyzer | 0.05 | 0.05 | 0.05 | U-shaped | 1.4142 | 1 | 0.035355 | 0.035355 | 0.035355 | 0.00125 | 0.00125 | 0.00125
A1-10 | Positioning and pointing misalignment between the reference antenna and the receiving antenna | 0.01 | 0.01 | 0.01 | Rectangular | 1.7321 | 1 | 0.0057735 | 0.0057735 | 0.0057735 | 3.3333e-05 | 3.3333e-05 | 3.3333e-05
A1-11 | Impedance mismatch between the reference antenna and the network analyzer. | 0.05 | 0.05 | 0.05 | U-shaped | 1.4142 | 1 | 0.035355 | 0.035355 | 0.035355 | 0.00125 | 0.00125 | 0.00125
A1-3 | Quality of quiet zone | 0.1 | 0.1 | 0.1 | Gaussian | 1 | 1 | 0.1 | 0.1 | 0.1 | 0.01 | 0.01 | 0.01
A1-4b | Polarization mismatch between the reference antenna and the receiving antenna | 0.01 | 0.01 | 0.01 | Rectangular | 1.7321 | 1 | 0.0057735 | 0.0057735 | 0.0057735 | 3.3333e-05 | 3.3333e-05 | 3.3333e-05
A1-5 | Mutual coupling between the reference antenna and the receiving antenna | 0 | 0 | 0 | Rectangular | 1.7321 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A1-6 | Phase curvature | 0.05 | 0.05 | 0.05 | Gaussian | 1 | 1 | 0.05 | 0.05 | 0.05 | 0.0025 | 0.0025 | 0.0025
C1-3 | Uncertainty of the network analyzer | 0.13 | 0.2 | 0.2 | Gaussian | 1 | 1 | 0.13 | 0.2 | 0.2 | 0.0169 | 0.04 | 0.04
A1-12 | Influence of the reference antenna feed cable | 0.05 | 0.05 | 0.05 | Rectangular | 1.7321 | 1 | 0.028868 | 0.028868 | 0.028868 | 0.00083333 | 0.00083333 | 0.00083333
A1-13 | Reference antenna feed cable loss measurement uncertainty | 0.06 | 0.06 | 0.06 | Gaussian | 1 | 1 | 0.06 | 0.06 | 0.06 | 0.0036 | 0.0036 | 0.0036
A1-14 | Influence of the receiving antenna feed cable | 0.05 | 0.05 | 0.05 | Rectangular | 1.7321 | 1 | 0.028868 | 0.028868 | 0.028868 | 0.00083333 | 0.00083333 | 0.00083333
C1-4 | Uncertainty of the absolute gain of the reference antenna | 0.50229 | 0.43301 | 0.43301 | Rectangular | 1.7321 | 1 | 0.29 | 0.25 | 0.25 | 0.0841 | 0.0625 | 0.0625
A1-15 | Uncertainty of the absolute gain of the receiving antenna | 0 | 0 | 0 | Rectangular | 1.7321 | 1 | 0 | 0 | 0 | 0 | 0 | 0
Combined standard uncertainty (1σ) [dB] | [7 empty] | 0.58771 | 0.73284 | 0.73284 | 0.58771 | 0.73284 | 0.73284
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [7 empty] | 1.1519 | 1.4364 | 1.4364 | 1.1519 | 1.4364 | 1.4364
Table 6.4.3.4-1: CATR MU value derivation for OTA E-UTRA DL RS power measurement
CATR | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz |  |  |  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz
Stage 2: DUT measurement
A2-1a | Misalignment DUT & pointing error for EIRP | 0 | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 | 0 | 0 | 0 | 0
C3-1 | DL-RS MU derived from conducted spec | 0.41 | 0.56 | 0.56 | Gaussian | 1 | 1 | 0.41 | 0.56 | 0.56 | 0.1681 | 0.3136 | 0.3136
A2-2a | Standing wave between DUT and test range antenna | 0.21 | 0.21 | 0.21 | U-shaped | 1.4142 | 1 | 0.14849 | 0.14849 | 0.14849 | 0.02205 | 0.02205 | 0.02205
A2-3 | RF leakage (SGH connector terminated & test range antenna connector cable terminated) | 0.0012 | 0.0012 | 0.0012 | Gaussian | 1 | 1 | 0.0012 | 0.0012 | 0.0012 | 1.44e-06 | 1.44e-06 | 1.44e-06
A2-4a | QZ ripple DUT | 0.0928 | 0.0928 | 0.0928 | Gaussian | 1 | 1 | 0.0928 | 0.0928 | 0.0928 | 0.0086118 | 0.0086118 | 0.0086118
A2-12 | Frequency flatness | 0.25 | 0.25 | 0.25 | Gaussian | 1 | 1 | 0.25 | 0.25 | 0.25 | 0.0625 | 0.0625 | 0.0625
Stage 1: Calibration measurement | [10 empty] | 0 | 0 | 0
6 | Network Analyzer | 0.13 | 0.2 | 0.2 | Gaussian | 1 | 1 | 0.13 | 0.2 | 0.2 | 0.0169 | 0.04 | 0.04
7 | Uncertainty of return loss (S11) measurement of SGH and test receiver (VNA) ports | 0.127 | 0.325 | 0.325 | U-shaped | 1.4142 | 1 | 0.089803 | 0.22981 | 0.22981 | 0.0080645 | 0.052812 | 0.052812
8 | Insertion loss variation in receiver chain | 0.18 | 0.18 | 0.18 | Rectangular | 1.7321 | 1 | 0.10392 | 0.10392 | 0.10392 | 0.0108 | 0.0108 | 0.0108
9 | RF leakage, test range antenna cable connector terminated. | 0.0012 | 0.0012 | 0.0012 | Gaussian | 1 | 1 | 0.0012 | 0.0012 | 0.0012 | 1.44e-06 | 1.44e-06 | 1.44e-06
10 | Influence of the calibration antenna feed cable | 0.022 | 0.022 | 0.022 | U-shaped | 1.4142 | 1 | 0.015556 | 0.015556 | 0.015556 | 0.000242 | 0.000242 | 0.000242
11 | SGH Calibration uncertainty | 0.50229 | 0.43301 | 0.43301 | Rectangular | 1.7321 | 1 | 0.29 | 0.25 | 0.25 | 0.0841 | 0.0625 | 0.0625
12 | Misalignment positioning system | 0 | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 | 0 | 0 | 0 | 0
13 | Misalignment SGH and pointing error | 0.5 | 0.5 | 0.5 | Exp. normal | 2 | 1 | 0.25 | 0.25 | 0.25 | 0.0625 | 0.0625 | 0.0625
14 | Rotary joints | 0.048 | 0.048 | 0.048 | U-shaped | 1.4142 | 1 | 0.033941 | 0.033941 | 0.033941 | 0.001152 | 0.001152 | 0.001152
15 | Standing wave between SGH and test range antenna | 0.09 | 0.09 | 0.09 | U-shaped | 1.4142 | 1 | 0.06364 | 0.06364 | 0.06364 | 0.00405 | 0.00405 | 0.00405
16 | QZ ripple with SGH | 0.009 | 0.009 | 0.009 | Gaussian | 1 | 1 | 0.009 | 0.009 | 0.009 | 8.1e-05 | 8.1e-05 | 8.1e-05
17 | Switching uncertainty | 0.26 | 0.26 | 0.26 | Rectangular | 1.7321 | 1 | 0.15011 | 0.15011 | 0.15011 | 0.022533 | 0.022533 | 0.022533
Combined standard uncertainty (1σ) [dB] | [7 empty] | 0.6868 | 0.81452 | 0.81452 | 0.6868 | 0.81452 | 0.81452
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [7 empty] | 1.3461 | 1.5965 | 1.5965 | 1.3461 | 1.5965 | 1.5965
Table 9.4.4.3-1: NFTR measurement uncertainty value derivation for OTA E-UTRA DL RS power measurement
Near field | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz |  |  |  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz
Stage 2: DUT measurement
A3-1 | Axes Intersection | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-2 | Axes Orthogonality | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-3 | Horizontal Pointing | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-4 | Probe Vertical Position | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-5 | Probe H/V pointing | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-6 | Measurement Distance | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-7 | Amplitude and Phase Drift | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-8 | Amplitude and Phase Noise | 0.02 | 0.02 | 0.02 | Gaussian | 1 | 1 | 0.02 | 0.02 | 0.02 | 0.0004 | 0.0004 | 0.0004
A3-9 | Leakage and Crosstalk | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-10 | Amplitude Non-Linearity | 0.04 | 0.04 | 0.04 | Gaussian | 1 | 1 | 0.04 | 0.04 | 0.04 | 0.0016 | 0.0016 | 0.0016
A3-11 | Amplitude and Phase Shift in rotary joints | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-12 | Channel Balance Amplitude and Phase | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-13 | Probe Polarization Amplitude and Phase | 0.0001 | 0.0001 | 0.0001 | Gaussian | 1 | 1 | 0.0001 | 0.0001 | 0.0001 | 1e-08 | 1e-08 | 1e-08
A3-14 | Probe Pattern Knowledge | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-15 | Multiple Reflections | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-16 | Room Scattering | 0.09 | 0.09 | 0.09 | Gaussian | 1 | 1 | 0.09 | 0.09 | 0.09 | 0.0081 | 0.0081 | 0.0081
A3-17 | DUT support Scattering | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-18 | Scan Area Truncation | 0.003 | 0.003 | 0.003 | Gaussian | 1 | 1 | 0.003 | 0.003 | 0.003 | 9e-06 | 9e-06 | 9e-06
A3-19 | Sampling Point Offset | 0.0058 | 0.0058 | 0.0058 | Gaussian | 1 | 1 | 0.0058 | 0.0058 | 0.0058 | 3.364e-05 | 3.364e-05 | 3.364e-05
A3-20 | Spherical Mode Truncation | 0.015 | 0.015 | 0.015 | Gaussian | 1 | 1 | 0.015 | 0.015 | 0.015 | 0.000225 | 0.000225 | 0.000225
A3-21 | Positioning | 0.03 | 0.03 | 0.03 | Rectangular | 1.73 | 1 | 0.017341 | 0.017341 | 0.017341 | 0.00030071 | 0.00030071 | 0.00030071
A3-22 | Probe Array Uniformity | 0.055 | 0.055 | 0.055 | Gaussian | 1 | 1 | 0.055 | 0.055 | 0.055 | 0.003025 | 0.003025 | 0.003025
A3-23 | Mismatch of receiver chain | 0.284 | 0.284 | 0.284 | U-Shaped | 1.41 | 1 | 0.20142 | 0.20142 | 0.20142 | 0.040569 | 0.040569 | 0.040569
A3-24 | Insertion loss of receiver chain | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-25 | Uncertainty of the absolute gain of the probe antenna | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
C3-1 | DL-RS MU derived from conducted spec | 0.41 | 0.56 | 0.56 | Gaussian | 1 | 1 | 0.41 | 0.56 | 0.56 | 0.1681 | 0.3136 | 0.3136
A3-26 | Measurement repeatability - positioning repeatability | 0.15 | 0.15 | 0.15 | Gaussian | 1 | 1 | 0.15 | 0.15 | 0.15 | 0.0225 | 0.0225 | 0.0225
Stage 1: Calibration measurement | [10 empty] | 0 | 0 | 0
C1-3 | Uncertainty of the network analyzer | 0.13 | 0.2 | 0.2 | Gaussian | 1 | 1 | 0.13 | 0.2 | 0.2 | 0.0169 | 0.04 | 0.04
A3-27 | Mismatch of receiver chain | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-28 | Insertion loss of receiver chain | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-29 | Mismatch in the connection of the calibration antenna | 0.02 | 0.02 | 0.02 | U-Shaped | 1.41 | 1 | 0.014184 | 0.014184 | 0.014184 | 0.0002012 | 0.0002012 | 0.0002012
A3-30 | Influence of the calibration antenna feed cable | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-31 | Influence of the probe antenna cable | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
C1-4 | Uncertainty of the absolute gain of the reference antenna | 0.50229 | 0.43301 | 0.43301 | Rectangular | 1.7321 | 1 | 0.29 | 0.25 | 0.25 | 0.0841 | 0.0625 | 0.0625
A3-32 | Short term repeatability | 0.088 | 0.088 | 0.088 | Gaussian | 1 | 1 | 0.088 | 0.088 | 0.088 | 0.007744 | 0.007744 | 0.007744
Combined standard uncertainty (1σ) [dB] | [7 empty] | 0.59482 | 0.70768 | 0.70768 | 0.59482 | 0.70768 | 0.70768
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [7 empty] | 1.1658 | 1.387 | 1.387 | 1.1658 | 1.387 | 1.387
One Dimensional Compact Range Chamber | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz |  |  |  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz
Stage 2: DUT measurement
Stage 1: Calibration measurement | [10 empty] | 0 | 0 | 0
Plane wave synthesiser | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz |  |  |  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz
Stage 2: DUT measurement
Stage 1: Calibration measurement | [10 empty] | 0 | 0 | 0
Checked 28/1/20
Need to update with final EIRP values
same open questions

## power dynamics
| chamber | Expanded uncertainty [dB] | [11 empty] | Review comments (RAN4#94)
|  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz
| Indoor anechoic | 0.40801 | 0.40801 | 0.40801
| CATR | 0.392 | 0.392 | 0.392
| Near field | 0.392 | 0.392 | 0
1D CATR
Plane wave sythesiser
| Common maximum accepted test system uncertainty | 0.4 | 0.4 | 0.4
Table 9.5.2.3-1: Indoor Anechoic Chamber measurement uncertainty value derivation for OTA total power dynamic range measurement
Indoor anechoic | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz |  |  |  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz
Stage 2: DUT measurement
C3-2 | Total power dynamic range conducted unbcertainty | 0.2 | 0.2 | 0.2 | Gaussian | 1 | 1 | 0.2 | 0.2 | 0.2 | 0.04 | 0.04 | 0.04
A1-8 | Random uncertainty | 0.1 | 0.1 | 0.1 | Rectangular | 1.7321 | 1 | 0.057735 | 0.057735 | 0.057735 | 0.0033333 | 0.0033333 | 0.0033333
Stage 1: Calibration measurement | [10 empty] | 0 | 0 | 0
Combined standard uncertainty (1σ) [dB] | [7 empty] | 0.20817 | 0.20817 | 0.20817 | 0.20817 | 0.20817 | 0.20817
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [7 empty] | 0.40801 | 0.40801 | 0.40801 | 0.40801 | 0.40801 | 0.40801
Table 9.5.3.3-1: CATR MU value derivation for OTA total power dynamic range measurement
CATR | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz |  |  |  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz
Stage 2: DUT measurement
C3-2 | Total power dynamic range conducted unbcertainty | 0.2 | 0.2 | 0.2 | Gaussian | 1 | 1 | 0.2 | 0.2 | 0.2 | 0.04 | 0.04 | 0.04
Stage 1: Calibration measurement | [10 empty] | 0 | 0 | 0
Combined standard uncertainty (1σ) [dB] | [7 empty] | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [7 empty] | 0.392 | 0.392 | 0.392 | 0.392 | 0.392 | 0.392
Table 9.5.3.4-1: NFTR MU value derivation for OTA E-UTRA total power dynamic range measurement
Near field | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz |  |  |  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz
Stage 2: DUT measurement
C3-2 | Total power dynamic range conducted unbcertainty | 0.2 | 0.2 | 0.2 | Gaussian | 1 | 1 | 0.2 | 0.2 | 0.2 | 0.04 | 0.04 | 0.04
Stage 1: Calibration measurement | [10 empty] | 0 | 0 | 0
Combined standard uncertainty (1σ) [dB] | [7 empty] | 0.2 | 0.2 |  | 0.2 | 0.2 | 0.2
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [7 empty] | 0.392 | 0.392 |  | 0.392 | 0.392 | 0.392
One Dimensional Compact Range Chamber | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz |  |  |  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz
Stage 2: DUT measurement
Stage 1: Calibration measurement | [10 empty] | 0 | 0 | 0
Plane wave synthesiser | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz |  |  |  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz
Stage 2: DUT measurement
Stage 1: Calibration measurement | [10 empty] | 0 | 0 | 0
Checked 28/1/20
Need to update with final EIRP values
same open questions

## EVM
| chamber | Expanded uncertainty [dB] |  |  | Total Evm on top of existing 1% EVM uncertainty [%] | [8 empty] | Review comments (RAN4#94)
|  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz
| Indoor anechoic | 0.24663 | 0.24663 | 0.24663 | 1.0288 | 1.0288 | 1.0288
| CATR | 0.34321 | 0.34321 | 0.34321 | 1.0403 | 1.0403 | 1.0403
Near field
1D CATR
Plane wave sythesiser
| Common maximum accepted test system uncertainty |  |  |  | 1 | 1 | 1
Table 9.7.2.3-1: IAC MU value derivation for power uncertainty aspects of OTA EVM
Indoor anechoic | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz |  |  |  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz
Stage 2: DUT measurement
A1-3 | Quality of quiet zone | 0.1 | 0.1 | 0.1 | Gaussian | 1 | 1 | 0.1 | 0.1 | 0.1 | 0.01 | 0.01 | 0.01
A1-6 | Phase curvature | 0.05 | 0.05 | 0.05 | Gaussian | 1 | 1 | 0.05 | 0.05 | 0.05 | 0.0025 | 0.0025 | 0.0025
A1-8 | Random uncertainty | 0.1 | 0.1 | 0.1 | Rectangular | 1.7321 | 1 | 0.057735 | 0.057735 | 0.057735 | 0.0033333 | 0.0033333 | 0.0033333
Stage 1: Calibration measurement | [10 empty] | 0 | 0 | 0
Combined standard uncertainty (1σ) [dB] | [7 empty] | 0.12583 | 0.12583 | 0.12583 | 0.12583 | 0.12583 | 0.12583
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [7 empty] | 0.24663 | 0.24663 | 0.24663 | 0.24663 | 0.24663 | 0.24663
Table 9.7.3.34-1: CATR MU value derivation for power variation aspects for OTA E-UTRA DL RS power measurement
CATR | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz |  |  |  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz
Stage 2: DUT measurement
A2-2a | Standing wave between DUT and test range antenna | 0.21 | 0.21 | 0.21 | U-shaped | 1.4142 | 1 | 0.14849 | 0.14849 | 0.14849 | 0.02205 | 0.02205 | 0.02205
A2-4a | QZ ripple DUT | 0.0928 | 0.0928 | 0.0928 | Gaussian | 1 | 1 | 0.0928 | 0.0928 | 0.0928 | 0.0086118 | 0.0086118 | 0.0086118
Stage 1: Calibration measurement | [10 empty] | 0 | 0 | 0
Combined standard uncertainty (1σ) [dB] | [7 empty] | 0.17511 | 0.17511 | 0.17511 | 0.17511 | 0.17511 | 0.17511
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [7 empty] | 0.34321 | 0.34321 | 0.34321 | 0.34321 | 0.34321 | 0.34321
Near field | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz |  |  |  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz
Stage 2: DUT measurement
Stage 1: Calibration measurement | [10 empty] | 0 | 0 | 0
One Dimensional Compact Range Chamber | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz |  |  |  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz
Stage 2: DUT measurement
Plane wave synthesiser | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz |  |  |  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz
Stage 2: DUT measurement
Stage 1: Calibration measurement | [10 empty] | 0 | 0 | 0

## In-band TRP
| chamber | Expanded uncertainty [dB] | [12 empty] | Review comments (RAN4#94)
|  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz
| Indoor anechoic | 1.1484 | 1.2964 | 1.2964
| CATR | 1.3854 | 1.5104 | 1.5104
| Near field | 1.2611 | 1.3343 | 1.3343
1D CATR
| Plane wave sythesiser | 1.2346 | 1.4003 | 1.4003
| Reverberation Chamber | 1.3653 | 1.4619 | 1.4619
| Common maximum accepted test system uncertainty | 1.4 | 1.5 | 1.5
Table 11.2.2.3-1: IAC MU value derivation for EIRP measurement of OTA BS output power
Indoor anechoic | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz |  |  |  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz
Stage 2: DUT measurement
A1-1 | Positioning misalignment between the AAS BS and the reference antenna | 0.03 | 0.03 | 0.03 | Rectangular | 1.7321 | 1 | 0.017321 | 0.017321 | 0.017321 | 0.0003 | 0.0003 | 0.0003
A1-2 | Pointing misalignment between the AAS BS and the receiving antenna | 0.3 | 0.3 | 0.3 | Rectangular | 1.7321 | 1 | 0.17321 | 0.17321 | 0.17321 | 0.03 | 0.03 | 0.03
A1-3 | Quality of quiet zone | 0.1 | 0.1 | 0.1 | Gaussian | 1 | 1 | 0.1 | 0.1 | 0.1 | 0.01 | 0.01 | 0.01
A1-4a | Polarization mismatch between the AAS BS and the receiving antenna | 0.01 | 0.01 | 0.01 | Rectangular | 1.7321 | 1 | 0.0057735 | 0.0057735 | 0.0057735 | 3.3333e-05 | 3.3333e-05 | 3.3333e-05
A1-5 | Mutual coupling between the AAS BS and the receiving antenna | 0 | 0 | 0 | Rectangular | 1.7321 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A1-6 | Phase curvature | 0.05 | 0.05 | 0.05 | Gaussian | 1 | 1 | 0.05 | 0.05 | 0.05 | 0.0025 | 0.0025 | 0.0025
C1-1 | Uncertainty of the RF Power Measurement Equipment | 0.14 | 0.26 | 0.26 | Gaussian | 1 | 1 | 0.14 | 0.26 | 0.26 | 0.0196 | 0.0676 | 0.0676
A1-7 | Impedance mismatch in the receiving chain | 0.14 | 0.33 | 0.33 | U-shaped | 1.4142 | 1 | 0.098995 | 0.23335 | 0.23335 | 0.0098 | 0.05445 | 0.05445
A1-8 | Random uncertainty | 0.1 | 0.1 | 0.1 | Rectangular | 1.7321 | 1 | 0.057735 | 0.057735 | 0.057735 | 0.0033333 | 0.0033333 | 0.0033333
Stage 1: Calibration measurement | [10 empty] | 0 | 0 | 0
A1-9 | Impedance mismatch between the receiving antenna and the network analyzer | 0.05 | 0.05 | 0.05 | U-shaped | 1.4142 | 1 | 0.035355 | 0.035355 | 0.035355 | 0.00125 | 0.00125 | 0.00125
A1-10 | Positioning and pointing misalignment between the reference antenna and the receiving antenna | 0.01 | 0.01 | 0.01 | Rectangular | 1.7321 | 1 | 0.0057735 | 0.0057735 | 0.0057735 | 3.3333e-05 | 3.3333e-05 | 3.3333e-05
A1-11 | Impedance mismatch between the reference antenna and the network analyzer. | 0.05 | 0.05 | 0.05 | U-shaped | 1.4142 | 1 | 0.035355 | 0.035355 | 0.035355 | 0.00125 | 0.00125 | 0.00125
A1-3 | Quality of quiet zone | 0.1 | 0.1 | 0.1 | Gaussian | 1 | 1 | 0.1 | 0.1 | 0.1 | 0.01 | 0.01 | 0.01
A1-4b | Polarization mismatch for reference antenna | 0.01 | 0.01 | 0.01 | Rectangular | 1.7321 | 1 | 0.0057735 | 0.0057735 | 0.0057735 | 3.3333e-05 | 3.3333e-05 | 3.3333e-05
A1-5 | Mutual coupling between the reference antenna and the receiving antenna | 0 | 0 | 0 | Rectangular | 1.7321 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A1-6 | Phase curvature | 0.05 | 0.05 | 0.05 | Gaussian | 1 | 1 | 0.05 | 0.05 | 0.05 | 0.0025 | 0.0025 | 0.0025
C1-3 | Uncertainty of the network analyzer | 0.13 | 0.2 | 0.2 | Gaussian | 1 | 1 | 0.13 | 0.2 | 0.2 | 0.0169 | 0.04 | 0.04
A1-12 | Influence of the reference antenna feed cable | 0.05 | 0.05 | 0.05 | Rectangular | 1.7321 | 1 | 0.028868 | 0.028868 | 0.028868 | 0.00083333 | 0.00083333 | 0.00083333
A1-13 | Reference antenna feed cable loss measurement uncertainty | 0.06 | 0.06 | 0.06 | Gaussian | 1 | 1 | 0.06 | 0.06 | 0.06 | 0.0036 | 0.0036 | 0.0036
A1-14 | Influence of the receiving antenna feed cable | 0.05 | 0.05 | 0.05 | Rectangular | 1.7321 | 1 | 0.028868 | 0.028868 | 0.028868 | 0.00083333 | 0.00083333 | 0.00083333
C1-4 | Uncertainty of the absolute gain of the reference antenna | 0.50229 | 0.43301 | 0.43301 | Rectangular | 1.7321 | 1 | 0.29 | 0.25 | 0.25 | 0.0841 | 0.0625 | 0.0625
A1-15 | Uncertainty of the absolute gain of the receiving antenna | 0 | 0 | 0 | Rectangular | 1.7321 | 1 | 0 | 0 | 0 | 0 | 0 | 0
Combined standard uncertainty (1σ) [dB] | [7 empty] | 0.44373 | 0.53949 | 0.53949 | 0.44373 | 0.53949 | 0.53949
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [7 empty] | 0.86972 | 1.0574 | 1.0574 | 0.86972 | 1.0574 | 1.0574
TRP Summation error | [7 empty] | 0.75 | 0.75 | 0.75
Total MU | [7 empty] | 1.1484 | 1.2964 | 1.2964 | same as 37.843
Table 11.2.3.3-1: CATR MU value derivation for OTA BS output power measurement, FR1
CATR | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz |  |  |  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz
Stage 2: DUT measurement
A2-18 | Misalignment DUT & pointing error for TRP | 0.3 | 0.3 | 0.3 | Rectangular | 1.7321 | 1 | 0.17321 | 0.17321 | 0.17321 | 0.03 | 0.03 | 0.03
C1-1 | RF power measurement equipment (e.g. spectrum analyzer, power meter) | 0.14 | 0.26 | 0.26 | Gaussian | 1 | 1 | 0.14 | 0.26 | 0.26 | 0.0196 | 0.0676 | 0.0676
A2-2a | Standing wave between DUT and test range antenna | 0.21 | 0.21 | 0.21 | U-shaped | 1.4142 | 1 | 0.14849 | 0.14849 | 0.14849 | 0.02205 | 0.02205 | 0.02205
A2-3 | RF leakage, test range antenna cable connector terminated. | 0.0012 | 0.0012 | 0.0012 | Gaussian | 1 | 1 | 0.0012 | 0.0012 | 0.0012 | 1.44e-06 | 1.44e-06 | 1.44e-06
A2-4a | QZ ripple with DUT | 0.0928 | 0.0928 | 0.0928 | Gaussian | 1 | 1 | 0.0928 | 0.0928 | 0.0928 | 0.0086118 | 0.0086118 | 0.0086118
A2-12 | Test system frequency flatness | 0.25 | 0.25 | 0.25 | Gaussian | 1 | 1 | 0.25 | 0.25 | 0.25 | 0.0625 | 0.0625 | 0.0625
Stage 1: Calibration measurement | [10 empty] | 0 | 0 | 0
C1-3 | Network Analyzer | 0.13 | 0.2 | 0.2 | Gaussian | 1 | 1 | 0.13 | 0.2 | 0.2 | 0.0169 | 0.04 | 0.04
A2-5 | Uncertainty of return loss (S11) measurement of SGH and test receiver (VNA) ports | 0.127 | 0.325 | 0.325 | U-shaped | 1.4142 | 1 | 0.089803 | 0.22981 | 0.22981 | 0.0080645 | 0.052812 | 0.052812
A2-6 | Insertion loss variation in receiver chain | 0.18 | 0.18 | 0.18 | Rectangular | 1.7321 | 1 | 0.10392 | 0.10392 | 0.10392 | 0.0108 | 0.0108 | 0.0108
A2-3 | RF leakage, test range antenna cable connector terminated. | 0.0012 | 0.0012 | 0.0012 | Gaussian | 1 | 1 | 0.0012 | 0.0012 | 0.0012 | 1.44e-06 | 1.44e-06 | 1.44e-06
A2-7 | Influence of the calibration antenna feed cable | 0.022 | 0.022 | 0.022 | U-shaped | 1.4142 | 1 | 0.015556 | 0.015556 | 0.015556 | 0.000242 | 0.000242 | 0.000242
C1-4 | SGH Calibration uncertainty | 0.50229 | 0.43301 | 0.43301 | Rectangular | 1.7321 | 1 | 0.29 | 0.25 | 0.25 | 0.0841 | 0.0625 | 0.0625
A2-8 | Misalignment positioning system | 0 | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A2-1b | Misalignment SGH and pointing error | 0.5 | 0.5 | 0.5 | Exp. normal | 2 | 1 | 0.25 | 0.25 | 0.25 | 0.0625 | 0.0625 | 0.0625
A2-9 | Rotary joints | 0.048 | 0.048 | 0.048 | U-shaped | 1.4142 | 1 | 0.033941 | 0.033941 | 0.033941 | 0.001152 | 0.001152 | 0.001152
A2-2b | Standing wave between SGH and test range antenna | 0.09 | 0.09 | 0.09 | U-shaped | 1.4142 | 1 | 0.06364 | 0.06364 | 0.06364 | 0.00405 | 0.00405 | 0.00405
A2-4b | QZ ripple with SGH | 0.009 | 0.009 | 0.009 | Gaussian | 1 | 1 | 0.009 | 0.009 | 0.009 | 8.1e-05 | 8.1e-05 | 8.1e-05
A2-11 | Switching uncertainty | 0.26 | 0.26 | 0.26 | Rectangular | 1.7321 | 1 | 0.15011 | 0.15011 | 0.15011 | 0.022533 | 0.022533 | 0.022533
Combined standard uncertainty (1σ) [dB] | [7 empty] | 0.5943 | 0.66891 | 0.66891 | 0.5943 | 0.66891 | 0.66891
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [7 empty] | 1.1648 | 1.3111 | 1.3111 | 1.1648 | 1.3111 | 1.3111
TRP Summation error | [7 empty] | 0.75 | 0.75 | 0.75
Total MU | [7 empty] | 1.3854 | 1.5104 | 1.5104 | pointimg error use 0 or 0.3, otherwise sa TR37.843
Table 11.2.4.3-1: NFTR MU value derivation for power density pattern measurement
Near field | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz |  |  |  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz
Stage 2: DUT measurement
A3-1 | Axes Intersection | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-2 | Axes Orthogonality | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-3 | Horizontal Pointing | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-4 | Probe Vertical Position | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-5 | Probe H/V pointing | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-6 | Measurement Distance | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-7 | Amplitude and Phase Drift | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-8 | Amplitude and Phase Noise | 0.02 | 0.02 | 0.02 | Gaussian | 1 | 1 | 0.02 | 0.02 | 0.02 | 0.0004 | 0.0004 | 0.0004
A3-9 | Leakage and Crosstalk | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-10 | Amplitude Non-Linearity | 0.04 | 0.04 | 0.04 | Gaussian | 1 | 1 | 0.04 | 0.04 | 0.04 | 0.0016 | 0.0016 | 0.0016
A3-11 | Amplitude and Phase Shift in rotary joints | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-12 | Channel Balance Amplitude and Phase | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-13 | Probe Polarization Amplitude and Phase | 0.0001 | 0.0001 | 0.0001 | Gaussian | 1 | 1 | 0.0001 | 0.0001 | 0.0001 | 1e-08 | 1e-08 | 1e-08
A3-14 | Probe Pattern Knowledge | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-15 | Multiple Reflections | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-16 | Room Scattering | 0.09 | 0.09 | 0.09 | Gaussian | 1 | 1 | 0.09 | 0.09 | 0.09 | 0.0081 | 0.0081 | 0.0081
A3-17 | DUT support Scattering | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-21 | Positioning | 0.03 | 0.03 | 0.03 | Rectangular | 1.73 | 1 | 0.017341 | 0.017341 | 0.017341 | 0.00030071 | 0.00030071 | 0.00030071
A3-22 | Probe Array Uniformity | 0.055 | 0.055 | 0.055 | Gaussian | 1 | 1 | 0.055 | 0.055 | 0.055 | 0.003025 | 0.003025 | 0.003025
A3-23 | Mismatch of receiver chain | 0.284 | 0.284 | 0.284 | U-Shaped | 1.41 | 1 | 0.20142 | 0.20142 | 0.20142 | 0.040569 | 0.040569 | 0.040569
A3-24 | Insertion loss of receiver chain | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-25 | Uncertainty of the absolute gain of the probe antenna | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
C1-1 | RF power measurement equipment (e.g. spectrum analyzer, power meter) | 0.14 | 0.26 | 0.26 | Gaussian | 1 | 1 | 0.14 | 0.26 | 0.26 | 0.0196 | 0.0676 | 0.0676 |  | RK:this figure is wrong in 37.843, 0.41/0.56 is used?
A3-26 | Measurement repeatability - positioning repeatability | 0.15 | 0.15 | 0.15 | Gaussian | 1 | 1 | 0.15 | 0.15 | 0.15 | 0.0225 | 0.0225 | 0.0225
A3-33 | Test system frequency flatness | 0.25 | 0.25 | 0.25 | Gaussian | 1 | 1 | 0.25 | 0.25 | 0.25 | 0.0625 | 0.0625 | 0.0625
Stage 1: Calibration measurement
C1-3 | Network analyzer | 0.13 | 0.2 | 0.2 | Gaussian | 1 | 1 | 0.13 | 0.2 | 0.2 | 0.0169 | 0.04 | 0.04
A3-27 | Mismatch of receiver chain | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-28 | Insertion loss of receiver chain | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-29 | Mismatch in the connection of the calibration antenna | 0.02 | 0.02 | 0.02 | U-Shaped | 1.41 | 1 | 0.014184 | 0.014184 | 0.014184 | 0.0002012 | 0.0002012 | 0.0002012
A3-30 | Influence of the calibration antenna feed cable | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-31 | Influence of the probe antenna cable | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
C1-4 | Reference antenna | 0.50229 | 0.43301 | 0.43301 | Rectangular | 1.7321 | 1 | 0.29 | 0.25 | 0.25 | 0.0841 | 0.0625 | 0.0625
A3-32 | Short term repeatability | 0.088 | 0.088 | 0.088 | Gaussian | 1 | 1 | 0.088 | 0.088 | 0.088 | 0.007744 | 0.007744 | 0.007744
Combined standard uncertainty (1σ) [dB] | [7 empty] | 0.51724 | 0.56306 | 0.56306 | 0.51724 | 0.56306 | 0.56306 | Value in 37.843 has not been recalculated with frequency flatness (which is added in square brackets) so is wrong anyway. Correction of TE and 3 missing parameters also adjusts value. Final corrected value is a bit smaller so is ok.
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [7 empty] | 1.0138 | 1.1036 | 1.1036 | 1.0138 | 1.1036 | 1.1036
TRP Summation error | [7 empty] | 0.75 | 0.75 | 0.75
Total MU | [7 empty] | 1.2611 | 1.3343 | 1.3343
One Dimensional Compact Range Chamber | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz |  |  |  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz
Stage 2: DUT measurement
Stage 1: Calibration measurement | [10 empty] | 0 | 0 | 0
Table 11.2.6.3-1: PWS MU value derivation for OTA BS output power
Plane wave synthesiser | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz |  |  |  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz
Stage 2: DUT measurement
A7-1a | Misalignment DUT & pointing error | 0.1 | 0.1 | 0.1 | Rectangular | 1.7321 | 1 | 0.057735 | 0.057735 | 0.057735 | 0.0033333 | 0.0033333 | 0.0033333
C1-1 | RF power measurement equipment (e.g. spectrum analyzer, power meter) | 0.14 | 0.26 | 0.26 | Gaussian | 1 | 1 | 0.14 | 0.26 | 0.26 | 0.0196 | 0.0676 | 0.0676
A7-2a | Longitudinal position uncertainty (i.e. standing wave and imperfect field synthesis) for DUT antenna | 0.05 | 0.14 | 0.14 | Rectangular | 1.7321 | 1 | 0.028868 | 0.080829 | 0.080829 | 0.00083333 | 0.0065333 | 0.0065333
A7-3 | RF leakage (calibration antenna connector terminated) | 0.086 | 0.086 | 0.086 | Gaussian | 1 | 1 | 0.086 | 0.086 | 0.086 | 0.007396 | 0.007396 | 0.007396
A7-4a | QZ ripple with DUT | 0.42 | 0.43 | 0.43 | Rectangular | 1.7321 | 1 | 0.24249 | 0.24826 | 0.24826 | 0.0588 | 0.061633 | 0.061633
A7-5 | Miscellaneous Uncertainty | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A7-14 | System non-linearity | 0.062 | 0.062 | 0.062 | Rectangular | 1.7321 | 1 | 0.035796 | 0.035796 | 0.035796 | 0.0012813 | 0.0012813 | 0.0012813
A7-13 | Frequency Flatness | 0.13 | 0.13 | 0.13 | Rectangular | 1.7321 | 1 | 0.075056 | 0.075056 | 0.075056 | 0.0056333 | 0.0056333 | 0.0056333
Stage 1: Calibration measurement | [10 empty] | 0 | 0 | 0
C1-3 | Uncertainty of the network analyzer | 0.13 | 0.2 | 0.2 | Gaussian | 1 | 1 | 0.13 | 0.2 | 0.2 | 0.0169 | 0.04 | 0.04
C1-3 | Mismatch (i.e. reference antenna, network analyser and reference cable) | 0.127 | 0.325 | 0.325 | U-shaped | 1.4142 | 1 | 0.089803 | 0.22981 | 0.22981 | 0.0080645 | 0.052812 | 0.052812
A7-6 | Insertion loss variation | 0.18 | 0.18 | 0.18 | Rectangular | 1.7321 | 1 | 0.10392 | 0.10392 | 0.10392 | 0.0108 | 0.0108 | 0.0108
A7-3 | RF leakage (calibration antenna connector terminated) | 0.086 | 0.086 | 0.086 | Gaussian | 1 | 1 | 0.086 | 0.086 | 0.086 | 0.007396 | 0.007396 | 0.007396
A7-3 | Influence of the calibration antenna feed cable | 0.103 | 0.104 | 0.104 | Rectangular | 1.7321 | 1 | 0.059467 | 0.060044 | 0.060044 | 0.0035363 | 0.0036053 | 0.0036053
C1-4 | Uncertainty of the absolute gain of the reference antenna | 0.50229 | 0.43301 | 0.43301 | Rectangular | 1.7321 | 1 | 0.29 | 0.25 | 0.25 | 0.0841 | 0.0625 | 0.0625
C1-4 | Misalignment of positioning system | 0 | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A7-9 | Misalignment of calibration antenna & pointing error | 0.05 | 0.05 | 0.05 | Rectangular | 1.7321 | 1 | 0.028868 | 0.028868 | 0.028868 | 0.00083333 | 0.00083333 | 0.00083333
A7-1b | Rotary joints | 0 | 0 | 0 | U-shaped | 1.7321 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A7-10 | Longitudinal position uncertainty (i.e. standing wave and imperfect field synthesis) for calibration antenna | 0.12 | 0.12 | 0.12 | Rectangular | 1.7321 | 1 | 0.069282 | 0.069282 | 0.069282 | 0.0048 | 0.0048 | 0.0048
A7-2b | QZ ripple with calibration antenna | 0.2 | 0.2 | 0.2 | Rectangular | 1.7321 | 1 | 0.11547 | 0.11547 | 0.11547 | 0.013333 | 0.013333 | 0.013333
A7-4a | Switching uncertainty | 0.02 | 0.02 | 0.02 | Rectangular | 1.7321 | 1 | 0.011547 | 0.011547 | 0.011547 | 0.00013333 | 0.00013333 | 0.00013333
A7-11 | Field repeatability | 0.06 | 0.12 | 0.12 | Gaussian | 1 | 1 | 0.06 | 0.12 | 0.12 | 0.0036 | 0.0144 | 0.0144
Combined standard uncertainty (1σ) [dB] | [7 empty] | 0.50037 | 0.60334 | 0.60334 | 0.50037 | 0.60334 | 0.60334
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [7 empty] | 0.98073 | 1.1826 | 1.1826 | 0.98073 | 1.1826 | 1.1826
TRP Summation error | [7 empty] | 0.75 | 0.75 | 0.75 | As per RS comments,
Total MU | [7 empty] | 1.2346 | 1.4003 | 1.4003
Table 11.2.5.3-1: Reverberation chamber MU value derivation for OTA BS output power
Reverberation Chamber | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz |  |  |  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz
Stage 2: DUT measurement
C1-1 | RF power measurement equipment (e.g. spectrum analyzer, power meter) | 0.14 | 0.26 | 0.26 | Gaussian | 1 | 1 | 0.14 | 0.26 | 0.26 | 0.0196 | 0.0676 | 0.0676
A6-1 | Impedance mismatch in the receiving chain | 0.2 | 0.2 | 0.2 | U-shaped | 1.4142 | 1 | 0.14142 | 0.14142 | 0.14142 | 0.02 | 0.02 | 0.02
A6-2 | Random uncertainty | 0.1 | 0.1 | 0.1 | Rectangular | 1.7321 | 1 | 0.057735 | 0.057735 | 0.057735 | 0.0033333 | 0.0033333 | 0.0033333
Stage 1: Calibration measurement
A6-3 | Reference antenna radiation efficiency | 0.2 | 0.2 | 0.2 | Gaussian | 1 | 1 | 0.2 | 0.2 | 0.2 | 0.04 | 0.04 | 0.04
A6-4 | Mean value estimation of reference antenna radiation efficiency | 0.15 | 0.15 | 0.15 | Gaussian | 1 | 1 | 0.15 | 0.15 | 0.15 | 0.0225 | 0.0225 | 0.0225
C1-3 | Uncertainty of the network analyzer | 0.13 | 0.2 | 0.2 | Gaussian | 1 | 1 | 0.13 | 0.2 | 0.2 | 0.0169 | 0.04 | 0.04
A6-5 | Influence of the reference antenna feed cable | 0.2 | 0.2 | 0.2 | Gaussian | 1 | 1 | 0.2 | 0.2 | 0.2 | 0.04 | 0.04 | 0.04
A6-6 | Mean value estimation of transfer function | 0.27 | 0.27 | 0.27 | Gaussian | 1 | 1 | 0.27 | 0.27 | 0.27 | 0.0729 | 0.0729 | 0.0729
A6-7 | Uniformity of transfer function | 0.5 | 0.5 | 0.5 | Gaussian | 1 | 1 | 0.5 | 0.5 | 0.5 | 0.25 | 0.25 | 0.25 |  | RK: mismatch in Rx chain had error, in uncertainty value but std uncertainty was ok, corrected to 0.2 U-shaped. Small error in NA uncertainty. Final balue close to TR 37.843
Combined standard uncertainty (1σ) [dB] | [7 empty] | 0.69659 | 0.74588 | 0.74588 | 0.69659 | 0.74588 | 0.74588
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [7 empty] | 1.3653 | 1.4619 | 1.4619 | 1.3653 | 1.4619 | 1.4619
0.19799
Checked 28/1/20
Reverb corrections confoirmed with Ericsson
PWS corrections from R&S
Issues to be conformed:
-IAC should we add FF?
-CATR pointing error
-Near field corrections.
check 11/2/20
NF - removed A3-18-20 as per email 11/2/20
- removed A3-18-20 as per email 11/2/20

## ACLR-abs
Table 11.3.7-1: Test system specific MU values for the absolute OTA ACLR, Gaussian test conditions, FR1
| chamber | Expanded uncertainty [dB] | [11 empty] | Review comments (RAN4#94)
|  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz
| Indoor anechoic | 1.1484 | 1.2964 | 1.2964
| CATR | 1.3854 | 1.5104 | 1.5104
| Near field | 1.2611 | 1.3343 | 1.3343
1D CATR
| Plane wave sythesiser | 1.2346 | 1.4003 | 1.4003
| Reverberation Chamber | 1.3653 | 1.4619 | 1.4619
| Common maximum accepted test system uncertainty | 2.2 | 2.7 | 2.7
Table 11.3.2.3-1: IAC MU value derivation for the absolute ACLR measurementEIRP measurement of the OTA ACLR
Indoor anechoic | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz |  |  |  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz
Stage 2: DUT measurement
A1-1 | Positioning misalignment between the AAS BS and the reference antenna | 0.03 | 0.03 | 0.03 | Rectangular | 1.7321 | 1 | 0.017321 | 0.017321 | 0.017321 | 0.0003 | 0.0003 | 0.0003
A1-2 | Pointing misalignment between the AAS BS and the receiving antenna | 0.3 | 0.3 | 0.3 | Rectangular | 1.7321 | 1 | 0.17321 | 0.17321 | 0.17321 | 0.03 | 0.03 | 0.03
A1-3 | Quality of quiet zone | 0.1 | 0.1 | 0.1 | Gaussian | 1 | 1 | 0.1 | 0.1 | 0.1 | 0.01 | 0.01 | 0.01
A1-4a | Polarization mismatch between the AAS BS and the receiving antenna | 0.01 | 0.01 | 0.01 | Rectangular | 1.7321 | 1 | 0.0057735 | 0.0057735 | 0.0057735 | 3.3333e-05 | 3.3333e-05 | 3.3333e-05
A1-5 | Mutual coupling between the AAS BS and the receiving antenna | 0 | 0 | 0 | Rectangular | 1.7321 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A1-6 | Phase curvature | 0.05 | 0.05 | 0.05 | Gaussian | 1 | 1 | 0.05 | 0.05 | 0.05 | 0.0025 | 0.0025 | 0.0025
C1-1 | RF power measurement equipment (e.g. spectrum analyzer, power meter) | 0.14 | 0.26 | 0.26 | Gaussian | 1 | 1 | 0.14 | 0.26 | 0.26 | 0.0196 | 0.0676 | 0.0676
A1-7 | Impedance mismatch in the receiving chain | 0.14 | 0.33 | 0.33 | U-shaped | 1.4142 | 1 | 0.098995 | 0.23335 | 0.23335 | 0.0098 | 0.05445 | 0.05445
A1-8 | Random uncertainty | 0.1 | 0.1 | 0.1 | Rectangular | 1.7321 | 1 | 0.057735 | 0.057735 | 0.057735 | 0.0033333 | 0.0033333 | 0.0033333
Stage 1: Calibration measurement | [10 empty] | 0 | 0 | 0
A1-9 | Impedance mismatch between the receiving antenna and the network analyzer | 0.05 | 0.05 | 0.05 | U-shaped | 1.4142 | 1 | 0.035355 | 0.035355 | 0.035355 | 0.00125 | 0.00125 | 0.00125
A1-10 | Positioning and pointing misalignment between the reference antenna and the receiving antenna | 0.01 | 0.01 | 0.01 | Rectangular | 1.7321 | 1 | 0.0057735 | 0.0057735 | 0.0057735 | 3.3333e-05 | 3.3333e-05 | 3.3333e-05
A1-11 | Impedance mismatch between the reference antenna and the network analyzer. | 0.05 | 0.05 | 0.05 | U-shaped | 1.4142 | 1 | 0.035355 | 0.035355 | 0.035355 | 0.00125 | 0.00125 | 0.00125
A1-3 | Quality of quiet zone | 0.1 | 0.1 | 0.1 | Gaussian | 1 | 1 | 0.1 | 0.1 | 0.1 | 0.01 | 0.01 | 0.01
A1-4b | Polarization mismatch between the reference antenna and the receiving antenna | 0.01 | 0.01 | 0.01 | Rectangular | 1.7321 | 1 | 0.0057735 | 0.0057735 | 0.0057735 | 3.3333e-05 | 3.3333e-05 | 3.3333e-05
A1-5 | Mutual coupling between the reference antenna and the receiving antenna | 0 | 0 | 0 | Rectangular | 1.7321 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A1-6 | Phase curvature | 0.05 | 0.05 | 0.05 | Gaussian | 1 | 1 | 0.05 | 0.05 | 0.05 | 0.0025 | 0.0025 | 0.0025
C1-3 | Uncertainty of the network analyzer | 0.13 | 0.2 | 0.2 | Gaussian | 1 | 1 | 0.13 | 0.2 | 0.2 | 0.0169 | 0.04 | 0.04
A1-12 | Influence of the reference antenna feed cable | 0.05 | 0.05 | 0.05 | Rectangular | 1.7321 | 1 | 0.028868 | 0.028868 | 0.028868 | 0.00083333 | 0.00083333 | 0.00083333
A1-13 | Reference antenna feed cable loss measurement uncertainty | 0.06 | 0.06 | 0.06 | Gaussian | 1 | 1 | 0.06 | 0.06 | 0.06 | 0.0036 | 0.0036 | 0.0036
A1-14 | Influence of the receiving antenna feed cable | 0.05 | 0.05 | 0.05 | Rectangular | 1.7321 | 1 | 0.028868 | 0.028868 | 0.028868 | 0.00083333 | 0.00083333 | 0.00083333
C1-4 | Uncertainty of the absolute gain of the reference antenna | 0.50229 | 0.43301 | 0.43301 | Rectangular | 1.7321 | 1 | 0.29 | 0.25 | 0.25 | 0.0841 | 0.0625 | 0.0625
A1-15 | Uncertainty of the absolute gain of the receiving antenna | 0 | 0 | 0 | Rectangular | 1.7321 | 1 | 0 | 0 | 0 | 0 | 0 | 0
Combined standard uncertainty (1σ) [dB] | [7 empty] | 0.44373 | 0.53949 | 0.53949 | 0.44373 | 0.53949 | 0.53949
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [7 empty] | 0.86972 | 1.0574 | 1.0574 | 0.86972 | 1.0574 | 1.0574
TRP Summation error | [7 empty] | 0.75 | 0.75 | 0.75
Total MU | [7 empty] | 1.1484 | 1.2964 | 1.2964 | Same as 37.843
Table 11.3.3.3-1: CATR MU value derivation for the EIRP measurement of the absolute OTA ACLR, FR1
CATR | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz |  |  |  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz
Stage 2: DUT measurement
A2-18 | Misalignment DUT & pointing error for TRP | 0.3 | 0.3 | 0.3 | Rectangular | 1.7321 | 1 | 0.17321 | 0.17321 | 0.17321 | 0.03 | 0.03 | 0.03
C1-1 | RF power measurement equipment (e.g. spectrum analyzer, power meter) | 0.14 | 0.26 | 0.26 | Gaussian | 1 | 1 | 0.14 | 0.26 | 0.26 | 0.0196 | 0.0676 | 0.0676
A2-2a | Standing wave between DUT and test range antenna | 0.21 | 0.21 | 0.21 | U-shaped | 1.4142 | 1 | 0.14849 | 0.14849 | 0.14849 | 0.02205 | 0.02205 | 0.02205
A2-3 | RF leakage (SGH connector terminated & test range antenna connector cable terminated) | 0.0012 | 0.0012 | 0.0012 | Gaussian | 1 | 1 | 0.0012 | 0.0012 | 0.0012 | 1.44e-06 | 1.44e-06 | 1.44e-06
A2-4a | QZ ripple DUT | 0.0928 | 0.0928 | 0.0928 | Gaussian | 1 | 1 | 0.0928 | 0.0928 | 0.0928 | 0.0086118 | 0.0086118 | 0.0086118
A2-12 | Frequency flatness | 0.25 | 0.25 | 0.25 | Gaussian | 1 | 1 | 0.25 | 0.25 | 0.25 | 0.0625 | 0.0625 | 0.0625
Stage 1: Calibration measurement | [10 empty] | 0 | 0 | 0
C1-3 | Uncertainty of the network analyzer | 0.13 | 0.2 | 0.2 | Gaussian | 1 | 1 | 0.13 | 0.2 | 0.2 | 0.0169 | 0.04 | 0.04
A2-5 | Mismatch of receiver chain | 0.127 | 0.325 | 0.325 | U-shaped | 1.4142 | 1 | 0.089803 | 0.22981 | 0.22981 | 0.0080645 | 0.052812 | 0.052812
A2-6 | Insertion loss variation of receiver chain | 0.18 | 0.18 | 0.18 | Rectangular | 1.7321 | 1 | 0.10392 | 0.10392 | 0.10392 | 0.0108 | 0.0108 | 0.0108
A2-3 | RF leakage (SGH connector terminated & test range antenna connector cable terminated) | 0.0012 | 0.0012 | 0.0012 | Gaussian | 1 | 1 | 0.0012 | 0.0012 | 0.0012 | 1.44e-06 | 1.44e-06 | 1.44e-06
A2-7 | Influence of the calibration antenna feed cable: | 0.022 | 0.022 | 0.022 | U-shaped | 1.4142 | 1 | 0.015556 | 0.015556 | 0.015556 | 0.000242 | 0.000242 | 0.000242
C1-4 | Uncertainty of the absolute gain of the reference antenna | 0.50229 | 0.43301 | 0.43301 | Rectangular | 1.7321 | 1 | 0.29 | 0.25 | 0.25 | 0.0841 | 0.0625 | 0.0625
A2-8 | Misalignment positioning system | 0 | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A2-1b | Misalignment of calibration antenna and test range antenna | 0.5 | 0.5 | 0.5 | Exp. normal | 2 | 1 | 0.25 | 0.25 | 0.25 | 0.0625 | 0.0625 | 0.0625
A2-9 | Rotary Joints | 0.048 | 0.048 | 0.048 | U-shaped | 1.4142 | 1 | 0.033941 | 0.033941 | 0.033941 | 0.001152 | 0.001152 | 0.001152
A2-2b | Standing wave between calibration antenna and test range antenna | 0.09 | 0.09 | 0.09 | U-shaped | 1.4142 | 1 | 0.06364 | 0.06364 | 0.06364 | 0.00405 | 0.00405 | 0.00405
A2-4b | QZ ripple calibration antenna | 0.009 | 0.009 | 0.009 | Gaussian | 1 | 1 | 0.009 | 0.009 | 0.009 | 8.1e-05 | 8.1e-05 | 8.1e-05
A2-11 | Switching uncertainty | 0.26 | 0.26 | 0.26 | Rectangular | 1.7321 | 1 | 0.15011 | 0.15011 | 0.15011 | 0.022533 | 0.022533 | 0.022533
Combined standard uncertainty (1σ) [dB] | [7 empty] | 0.5943 | 0.66891 | 0.66891 | 0.5943 | 0.66891 | 0.66891
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [7 empty] | 1.1648 | 1.3111 | 1.3111 | 1.1648 | 1.3111 | 1.3111
TRP Summation error | [7 empty] | 0.75 | 0.75 | 0.75 |  |  |  | Nok: As there is a significant difference between the new and old MU values, further checking is needed. A suggestion is to keep these values in [].
Total MU | [7 empty] | 1.3854 | 1.5104 | 1.5104 | Calculation is done wrong in 37.843, this nnumber is much lower
Table 11.3.4.3-1: NFTR MU value derivation for absolute ACLR measurement
Near field | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz |  |  |  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz
Stage 2: DUT measurement
A3-1 | Axes Intersection | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-2 | Axes Orthogonality | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-3 | Horizontal Pointing | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-4 | Probe Vertical Position | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-5 | Probe H/V pointing | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-6 | Measurement Distance | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-7 | Amplitude and Phase Drift | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-8 | Amplitude and Phase Noise | 0.02 | 0.02 | 0.02 | Gaussian | 1 | 1 | 0.02 | 0.02 | 0.02 | 0.0004 | 0.0004 | 0.0004
A3-9 | Leakage and Crosstalk | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-10 | Amplitude Non-Linearity | 0.04 | 0.04 | 0.04 | Gaussian | 1 | 1 | 0.04 | 0.04 | 0.04 | 0.0016 | 0.0016 | 0.0016
A3-11 | Amplitude and Phase Shift in rotary joints | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-12 | Channel Balance Amplitude and Phase | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-13 | Probe Polarization Amplitude and Phase | 0.0001 | 0.0001 | 0.0001 | Gaussian | 1 | 1 | 0.0001 | 0.0001 | 0.0001 | 1e-08 | 1e-08 | 1e-08
A3-14 | Probe Pattern Knowledge | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-15 | Multiple Reflections | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-16 | Room Scattering | 0.09 | 0.09 | 0.09 | Gaussian | 1 | 1 | 0.09 | 0.09 | 0.09 | 0.0081 | 0.0081 | 0.0081
A3-17 | DUT support Scattering | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-21 | Positioning | 0.03 | 0.03 | 0.03 | Rectangular | 1.73 | 1 | 0.017341 | 0.017341 | 0.017341 | 0.00030071 | 0.00030071 | 0.00030071 | These UID numbers are wrong in TRP, do not line up with EIRP, they are correct here.
A3-22 | Probe Array Uniformity | 0.055 | 0.055 | 0.055 | Gaussian | 1 | 1 | 0.055 | 0.055 | 0.055 | 0.003025 | 0.003025 | 0.003025
A3-23 | Mismatch of receiver chain | 0.284 | 0.284 | 0.284 | U-Shaped | 1.41 | 1 | 0.20142 | 0.20142 | 0.20142 | 0.040569 | 0.040569 | 0.040569
A3-24 | Insertion loss of receiver chain | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-25 | Uncertainty of the absolute gain of the probe antenna | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
C1-1 | RF power measurement equipment (e.g. spectrum analyzer, power meter) | 0.14 | 0.26 | 0.26 | Gaussian | 1 | 1 | 0.14 | 0.26 | 0.26 | 0.0196 | 0.0676 | 0.0676 | The near filed calcs used error derived from the conducted measurement here, we should either use in all or just use the TE?
A3-26 | Measurement repeatability - positioning repeatability | 0.15 | 0.15 | 0.15 | Gaussian | 1 | 1 | 0.15 | 0.15 | 0.15 | 0.0225 | 0.0225 | 0.0225
A3-33 | Test system frequency flatness | 0.25 | 0.25 | 0.25 | Gaussian | 1 | 1 | 0.25 | 0.25 | 0.25 | 0.0625 | 0.0625 | 0.0625
Stage 1: Calibration measurement | [10 empty] | 0 | 0 | 0
C1-3 | Network analyzer | 0.13 | 0.2 | 0.2 | Gaussian | 1 | 1 | 0.13 | 0.2 | 0.2 | 0.0169 | 0.04 | 0.04
A3-27 | Mismatch of receiver chain | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-28 | Insertion loss of receiver chain | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-29 | Mismatch in the connection of the calibration antenna | 0.02 | 0.02 | 0.02 | U-Shaped | 1.41 | 1 | 0.014184 | 0.014184 | 0.014184 | 0.0002012 | 0.0002012 | 0.0002012
A3-30 | Influence of the calibration antenna feed cable | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-31 | Influence of the probe antenna cable | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
C1-4 | Reference antenna | 0.50229 | 0.43301 | 0.43301 | Rectangular | 1.7321 | 1 | 0.29 | 0.25 | 0.25 | 0.0841 | 0.0625 | 0.0625
A3-32 | Short term repeatability | 0.088 | 0.088 | 0.088 | Gaussian | 1 | 1 | 0.088 | 0.088 | 0.088 | 0.007744 | 0.007744 | 0.007744
Combined standard uncertainty (1σ) [dB] | [7 empty] | 0.51724 | 0.56306 | 0.56306 | 0.51724 | 0.56306 | 0.56306
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [7 empty] | 1.0138 | 1.1036 | 1.1036 | 1.0138 | 1.1036 | 1.1036
TRP Summation error | [7 empty] | 0.75 | 0.75 | 0.75
Total MU | [7 empty] | 1.2611 | 1.3343 | 1.3343 | different from TR
The calculation in 37.843 is for relative?
Table 11.3.6.3-2: PWS MU value derivation for absolute ACLR measurement
Plane wave synthesiser | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz |  |  |  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz
Stage 2: DUT measurement
A7-1a | Misalignment DUT & pointing error | 0.1 | 0.1 | 0.1 | Rectangular | 1.7321 | 1 | 0.057735 | 0.057735 | 0.057735 | 0.0033333 | 0.0033333 | 0.0033333
C1-1 | RF power measurement equipment (e.g. spectrum analyzer, power meter) | 0.14 | 0.26 | 0.26 | Gaussian | 1 | 1 | 0.14 | 0.26 | 0.26 | 0.0196 | 0.0676 | 0.0676
A7-2a | Longitudinal position uncertainty (i.e. standing wave and imperfect field synthesis) for DUT antenna | 0.05 | 0.14 | 0.14 | Rectangular | 1.7321 | 1 | 0.028868 | 0.080829 | 0.080829 | 0.00083333 | 0.0065333 | 0.0065333
A7-3 | RF leakage (calibration antenna connector terminated) | 0.086 | 0.086 | 0.086 | Gaussian | 1 | 1 | 0.086 | 0.086 | 0.086 | 0.007396 | 0.007396 | 0.007396
A7-4a | QZ ripple with DUT | 0.42 | 0.43 | 0.43 | Rectangular | 1.7321 | 1 | 0.24249 | 0.24826 | 0.24826 | 0.0588 | 0.061633 | 0.061633
A7-5 | Miscellaneous Uncertainty | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A7-14 | System non-linearity | 0.062 | 0.062 | 0.062 | Rectangular | 1.7321 | 1 | 0.035796 | 0.035796 | 0.035796 | 0.0012813 | 0.0012813 | 0.0012813
A7-13 | Frequency Flatness | 0.13 | 0.13 | 0.13 | Rectangular | 1.7321 | 1 | 0.075056 | 0.075056 | 0.075056 | 0.0056333 | 0.0056333 | 0.0056333
Stage 1: Calibration measurement | [10 empty] | 0 | 0 | 0
C1-3 | Uncertainty of the network analyzer | 0.13 | 0.2 | 0.2 | Gaussian | 1 | 1 | 0.13 | 0.2 | 0.2 | 0.0169 | 0.04 | 0.04
A7-6 | Mismatch (i.e. reference antenna, network analyser and reference cable) | 0.127 | 0.325 | 0.325 | U-shaped | 1.4142 | 1 | 0.089803 | 0.22981 | 0.22981 | 0.0080645 | 0.052812 | 0.052812
A7-7 | Insertion loss variation | 0.18 | 0.18 | 0.18 | Rectangular | 1.7321 | 1 | 0.10392 | 0.10392 | 0.10392 | 0.0108 | 0.0108 | 0.0108
A7-3 | RF leakage (calibration antenna connector terminated) | 0.086 | 0.086 | 0.086 | Gaussian | 1 | 1 | 0.086 | 0.086 | 0.086 | 0.007396 | 0.007396 | 0.007396
A7-8 | Influence of the calibration antenna feed cable | 0.103 | 0.104 | 0.104 | Rectangular | 1.7321 | 1 | 0.059467 | 0.060044 | 0.060044 | 0.0035363 | 0.0036053 | 0.0036053
C1-4 | Uncertainty of the absolute gain of the reference antenna | 0.50229 | 0.43301 | 0.43301 | Rectangular | 1.7321 | 1 | 0.29 | 0.25 | 0.25 | 0.0841 | 0.0625 | 0.0625
A7-9 | Misalignment of positioning system | 0 | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A7-1b | Misalignment of calibration antenna & pointing error | 0.05 | 0.05 | 0.05 | Rectangular | 1.7321 | 1 | 0.028868 | 0.028868 | 0.028868 | 0.00083333 | 0.00083333 | 0.00083333
A7-10 | Rotary joints | 0 | 0 | 0 | U-shaped | 1.7321 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A7-2b | Longitudinal position uncertainty (i.e. standing wave and imperfect field synthesis) for calibration antenna | 0.12 | 0.12 | 0.12 | Rectangular | 1.7321 | 1 | 0.069282 | 0.069282 | 0.069282 | 0.0048 | 0.0048 | 0.0048
A7-4a | QZ ripple with calibration antenna | 0.2 | 0.2 | 0.2 | Rectangular | 1.7321 | 1 | 0.11547 | 0.11547 | 0.11547 | 0.013333 | 0.013333 | 0.013333
A7-11 | Switching uncertainty | 0.02 | 0.02 | 0.02 | Rectangular | 1.7321 | 1 | 0.011547 | 0.011547 | 0.011547 | 0.00013333 | 0.00013333 | 0.00013333
A7-12 | Field repeatability | 0.06 | 0.12 | 0.12 | Gaussian | 1 | 1 | 0.06 | 0.12 | 0.12 | 0.0036 | 0.0144 | 0.0144
Combined standard uncertainty (1σ) [dB] | [7 empty] | 0.50037 | 0.60334 | 0.60334 | 0.50037 | 0.60334 | 0.60334
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [7 empty] | 0.98073 | 1.1826 | 1.1826 | 0.98073 | 1.1826 | 1.1826
TRP Summation error | [7 empty] | 0.75 | 0.75 | 0.75
Total MU | [7 empty] | 1.2346 | 1.4003 | 1.4003 | same as in-band TRP, now coppied the table over (not sure if its needed in final TR?) In line with R&S update
Table 11.3.5.3-2: Reverberation chamber MU value derivation for absolute ACLR measurement
Reverberation Chamber | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz |  |  |  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz
Stage 2: DUT measurement
C1-1 | RF power measurement equipment (e.g. spectrum analyzer, power meter) | 0.14 | 0.26 | 0.26 | Gaussian | 1 | 1 | 0.14 | 0.26 | 0.26 | 0.0196 | 0.0676 | 0.0676
A6-1 | Impedance mismatch in the receiving chain | 0.2 | 0.2 | 0.2 | U-shaped | 1.4142 | 1 | 0.14142 | 0.14142 | 0.14142 | 0.02 | 0.02 | 0.02
A6-2 | Random uncertainty | 0.1 | 0.1 | 0.1 | Rectangular | 1.7321 | 1 | 0.057735 | 0.057735 | 0.057735 | 0.0033333 | 0.0033333 | 0.0033333
Stage 1: Calibration measurement
A6-3 | Reference antenna radiation efficiency | 0.2 | 0.2 | 0.2 | Gaussian | 1 | 1 | 0.2 | 0.2 | 0.2 | 0.04 | 0.04 | 0.04
A6-4 | Mean value estimation of reference antenna radiation efficiency | 0.15 | 0.15 | 0.15 | Gaussian | 1 | 1 | 0.15 | 0.15 | 0.15 | 0.0225 | 0.0225 | 0.0225
C1-3 | Uncertainty of the network analyzer | 0.13 | 0.2 | 0.2 | Gaussian | 1 | 1 | 0.13 | 0.2 | 0.2 | 0.0169 | 0.04 | 0.04
A6-5 | Influence of the reference antenna feed cable | 0.2 | 0.2 | 0.2 | Gaussian | 1 | 1 | 0.2 | 0.2 | 0.2 | 0.04 | 0.04 | 0.04
A6-6 | Mean value estimation of transfer function | 0.27 | 0.27 | 0.27 | Gaussian | 1 | 1 | 0.27 | 0.27 | 0.27 | 0.0729 | 0.0729 | 0.0729
A6-7 | Uniformity of transfer function | 0.5 | 0.5 | 0.5 | Gaussian | 1 | 1 | 0.5 | 0.5 | 0.5 | 0.25 | 0.25 | 0.25
Combined standard uncertainty (1σ) [dB] | [7 empty] | 0.69659 | 0.74588 | 0.74588 | 0.69659 | 0.74588 | 0.74588
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [7 empty] | 1.3653 | 1.4619 | 1.4619 | 1.3653 | 1.4619 | 1.4619
small error in NA, otherwise sae as 37.843
Checked 28/1/20
IAC shoudl we add FF?
CATR-pointing error?
CATR calcualtion was done worng in 37.843
12/2/20
NF relative shoudl be absolute - mmoved to this sheet

## ACLR-rel
Table 11.3.7-2: Test system specific MU values for the relative OTA ACLR, Gaussian test conditions, FR1
| chamber | Expanded uncertainty [dB] | [11 empty] | Review comments (RAN4#94)
|  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz
Indoor anechoic
| CATR | 1.1181 | 1.302 | 1.302
Near field
1D CATR
| Plane wave sythesiser | 1.0179 | 1.2215 | 1.2215
| Reverberation Chamber | 1.3653 | 1.4619 | 1.4619
| Common maximum accepted test system uncertainty | 1 | 1.2 | 1.2
Indoor anechoic | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz |  |  |  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz
Stage 2: DUT measurement
Stage 1: Calibration measurement | [10 empty] | 0 | 0 | 0
no contribution in 37.843
Table 11.3.3.3-2: CATR MU value derivation for the EIRP measurement of the relative OTA ACLR, FR1
CATR | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz |  |  |  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz
Stage 2: DUT measurement
A2-18 | Misalignment DUT & pointing error for TRP | 0.3 | 0.3 | 0.3 | Rectangular | 1.7321 | 1 | 0.17321 | 0.17321 | 0.17321 | 0.03 | 0.03 | 0.03
C1-1 | RF power measurement equipment (e.g. spectrum analyzer, power meter) | 0.14 | 0.26 | 0.26 | Gaussian | 1 | 1 | 0.14 | 0.26 | 0.26 | 0.0196 | 0.0676 | 0.0676
A2-4a | QZ ripple DUT | 0.0928 | 0.0928 | 0.0928 | Gaussian | 1 | 1 | 0.0928 | 0.0928 | 0.0928 | 0.0086118 | 0.0086118 | 0.0086118
A2-12 | Frequency flatness | 0.25 | 0.25 | 0.25 | Gaussian | 1 | 1 | 0.25 | 0.25 | 0.25 | 0.0625 | 0.0625 | 0.0625
Stage 1: Calibration measurement | [10 empty] | 0 | 0 | 0
C1-3 | Uncertainty of the network analyzer | 0.13 | 0.2 | 0.2 | Gaussian | 1 | 1 | 0.13 | 0.2 | 0.2 | 0.0169 | 0.04 | 0.04
A2-5 | Mismatch of receiver chain | 0.127 | 0.325 | 0.325 | U-shaped | 1.4142 | 1 | 0.089803 | 0.22981 | 0.22981 | 0.0080645 | 0.052812 | 0.052812
A2-6 | Insertion loss variation of receiver chain | 0.18 | 0.18 | 0.18 | Rectangular | 1.7321 | 1 | 0.10392 | 0.10392 | 0.10392 | 0.0108 | 0.0108 | 0.0108
A2-11 | Switching uncertainty | 0.26 | 0.26 | 0.26 | Rectangular | 1.7321 | 1 | 0.15011 | 0.15011 | 0.15011 | 0.022533 | 0.022533 | 0.022533
Combined standard uncertainty (1σ) [dB] | [7 empty] | 0.4231 | 0.54301 | 0.54301 | 0.4231 | 0.54301 | 0.54301
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [7 empty] | 0.82927 | 1.0643 | 1.0643 | 0.82927 | 1.0643 | 1.0643
TRP Summation error | [7 empty] | 0.75 | 0.75 | 0.75
Total MU | [7 empty] | 1.1181 | 1.302 | 1.302 | TE error and issue with the pointing error, final va;ues very similar to 37.843
Table 11.3.4.3-1: NFTR MU value derivation for relative ACLR measurement
Near field | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz |  |  |  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz
Stage 2: DUT measurement
Stage 1: Calibration measurement
One Dimensional Compact Range Chamber | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz |  |  |  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz
Stage 2: DUT measurement
Stage 1: Calibration measurement | [10 empty] | 0 | 0 | 0
Table 11.3.6.3-1: PWS MU value derivation for relative ACLR measurement
Plane wave synthesiser | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz |  |  |  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz
Stage 2: DUT measurement
A7-1a | Misalignment DUT & pointing error | 0.1 | 0.1 | 0.1 | Rectangular | 1.7321 | 1 | 0.057735 | 0.057735 | 0.057735 | 0.0033333 | 0.0033333 | 0.0033333
C1-1 | RF power measurement equipment (e.g. spectrum analyzer, power meter) | 0.14 | 0.26 | 0.26 | Gaussian | 1 | 1 | 0.14 | 0.26 | 0.26 | 0.0196 | 0.0676 | 0.0676
A7-4a | QZ ripple with DUT | 0.42 | 0.43 | 0.43 | Rectangular | 1.7321 | 1 | 0.24249 | 0.24826 | 0.24826 | 0.0588 | 0.061633 | 0.061633
A7-13 | Frequency Flatness | 0.13 | 0.13 | 0.13 | Rectangular | 1.7321 | 1 | 0.075056 | 0.075056 | 0.075056 | 0.0056333 | 0.0056333 | 0.0056333
Stage 1: Calibration measurement | [10 empty] | 0 | 0 | 0
C1-3 | Uncertainty of the network analyzer | 0.13 | 0.2 | 0.2 | Gaussian | 1 | 1 | 0.13 | 0.2 | 0.2 | 0.0169 | 0.04 | 0.04
A7-6 | Mismatch (i.e. reference antenna, network analyser and reference cable) | 0.127 | 0.325 | 0.325 | U-shaped | 1.4142 | 1 | 0.089803 | 0.22981 | 0.22981 | 0.0080645 | 0.052812 | 0.052812
A7-7 | Insertion loss variation | 0.18 | 0.18 | 0.18 | Rectangular | 1.7321 | 1 | 0.10392 | 0.10392 | 0.10392 | 0.0108 | 0.0108 | 0.0108
A7-11 | Switching uncertainty | 0.02 | 0.02 | 0.02 | Rectangular | 1.7321 | 1 | 0.011547 | 0.011547 | 0.011547 | 0.00013333 | 0.00013333 | 0.00013333
Combined standard uncertainty (1σ) [dB] | [7 empty] | 0.35109 | 0.49188 | 0.49188 | 0.35109 | 0.49188 | 0.49188
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [7 empty] | 0.68814 | 0.96408 | 0.96408 | 0.68814 | 0.96408 | 0.96408
TRP Summation error | [7 empty] | 0.75 | 0.75 | 0.75
Total MU | [7 empty] | 1.0179 | 1.2215 | 1.2215 | Same as R&S update
Table 11.3.5.3-1: Reverberation chamber MU value derivation for relative ACLR measurement
Reverberation Chamber | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz |  |  |  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz
Stage 2: DUT measurement
C1-1 | RF power measurement equipment (e.g. spectrum analyzer, power meter) | 0.14 | 0.26 | 0.26 | Gaussian | 1 | 1 | 0.14 | 0.26 | 0.26 | 0.0196 | 0.0676 | 0.0676
A6-1 | Impedance mismatch in the receiving chain | 0.2 | 0.2 | 0.2 | U-shaped | 1.4142 | 1 | 0.14142 | 0.14142 | 0.14142 | 0.02 | 0.02 | 0.02
A6-2 | Random uncertainty | 0.1 | 0.1 | 0.1 | Rectangular | 1.7321 | 1 | 0.057735 | 0.057735 | 0.057735 | 0.0033333 | 0.0033333 | 0.0033333
Stage 1: Calibration measurement
A6-3 | Reference antenna radiation efficiency | 0.2 | 0.2 | 0.2 | Gaussian | 1 | 1 | 0.2 | 0.2 | 0.2 | 0.04 | 0.04 | 0.04
A6-4 | Mean value estimation of reference antenna radiation efficiency | 0.15 | 0.15 | 0.15 | Gaussian | 1 | 1 | 0.15 | 0.15 | 0.15 | 0.0225 | 0.0225 | 0.0225
C1-3 | Uncertainty of the network analyzer | 0.13 | 0.2 | 0.2 | Gaussian | 1 | 1 | 0.13 | 0.2 | 0.2 | 0.0169 | 0.04 | 0.04
A6-5 | Influence of the reference antenna feed cable | 0.2 | 0.2 | 0.2 | Gaussian | 1 | 1 | 0.2 | 0.2 | 0.2 | 0.04 | 0.04 | 0.04
A6-6 | Mean value estimation of transfer function | 0.27 | 0.27 | 0.27 | Gaussian | 1 | 1 | 0.27 | 0.27 | 0.27 | 0.0729 | 0.0729 | 0.0729
A6-7 | Uniformity of transfer function | 0.5 | 0.5 | 0.5 | Gaussian | 1 | 1 | 0.5 | 0.5 | 0.5 | 0.25 | 0.25 | 0.25
Combined standard uncertainty (1σ) [dB] | [7 empty] | 0.69659 | 0.74588 | 0.74588 | 0.69659 | 0.74588 | 0.74588
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [7 empty] | 1.3653 | 1.4619 | 1.4619 | 1.3653 | 1.4619 | 1.4619
Same as TR 37.843 (small rounding errors)
checked 28/1/20
CATR : ponting error issue?
NF: Used TE derived from conducted measurement. Other use the TE values direct. We should use same method for all. Error is still high there seems to be no cancelation /simplification of errors for relative? so value is high
Reverb chamber - error is high.
12/2/20
NF relative shoudl be absolute - moved to absolute sheet

## OBUE
Table 11.4.7-1: Test system specific MU values for the OTA OBUE measurement, FR1
| chamber | Expanded uncertainty [dB] | [11 empty] | Review comments (RAN4#94)
|  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz
| Indoor anechoic | 1.1484 | 1.2964 | 1.2964
| CATR | 1.3854 | 1.5104 | 1.5104
| Near field | 1.2611 | 1.3343 | 1.3343
1D CATR
| Plane wave sythesiser | 1.2346 | 1.4003 | 1.4003
| Reverberation Chamber | 1.3974 | 1.4619 | 1.4619
| Common maximum accepted test system uncertainty | 1.8 | 2 | 2 | *note 38.817-02 has an error in this section the values are 1.4 and 1.5 but 38.141-2 is ok
Table 11.4.2.3-1: IAC MU value derivation for OTA OBUE measurement, FR1
Indoor anechoic | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz |  |  |  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz
Stage 2: DUT measurement
A1-1 | Positioning misalignment between the AAS BS and the reference antenna | 0.03 | 0.03 | 0.03 | Rectangular | 1.7321 | 1 | 0.017321 | 0.017321 | 0.017321 | 0.0003 | 0.0003 | 0.0003
A1-2 | Pointing misalignment between the AAS BS and the receiving antenna | 0.3 | 0.3 | 0.3 | Rectangular | 1.7321 | 1 | 0.17321 | 0.17321 | 0.17321 | 0.03 | 0.03 | 0.03
A1-3 | Quality of quiet zone | 0.1 | 0.1 | 0.1 | Gaussian | 1 | 1 | 0.1 | 0.1 | 0.1 | 0.01 | 0.01 | 0.01
A1-4a | Polarization mismatch between the AAS BS and the receiving antenna | 0.01 | 0.01 | 0.01 | Rectangular | 1.7321 | 1 | 0.0057735 | 0.0057735 | 0.0057735 | 3.3333e-05 | 3.3333e-05 | 3.3333e-05
A1-5 | Mutual coupling between the AAS BS and the receiving antenna | 0 | 0 | 0 | Rectangular | 1.7321 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A1-6 | Phase curvature | 0.05 | 0.05 | 0.05 | Gaussian | 1 | 1 | 0.05 | 0.05 | 0.05 | 0.0025 | 0.0025 | 0.0025
C1-1 | RF power measurement equipment (e.g. spectrum analyzer, power meter) | 0.14 | 0.26 | 0.26 | Gaussian | 1 | 1 | 0.14 | 0.26 | 0.26 | 0.0196 | 0.0676 | 0.0676
A1-7 | Impedance mismatch in the receiving chain | 0.14 | 0.33 | 0.33 | U-shaped | 1.4142 | 1 | 0.098995 | 0.23335 | 0.23335 | 0.0098 | 0.05445 | 0.05445
A1-8 | Random uncertainty | 0.1 | 0.1 | 0.1 | Rectangular | 1.7321 | 1 | 0.057735 | 0.057735 | 0.057735 | 0.0033333 | 0.0033333 | 0.0033333
Stage 1: Calibration measurement | [10 empty] | 0 | 0 | 0
A1-9 | Impedance mismatch between the receiving antenna and the network analyzer | 0.05 | 0.05 | 0.05 | U-shaped | 1.4142 | 1 | 0.035355 | 0.035355 | 0.035355 | 0.00125 | 0.00125 | 0.00125
A1-10 | Positioning and pointing misalignment between the reference antenna and the receiving antenna | 0.01 | 0.01 | 0.01 | Rectangular | 1.7321 | 1 | 0.0057735 | 0.0057735 | 0.0057735 | 3.3333e-05 | 3.3333e-05 | 3.3333e-05
A1-11 | Impedance mismatch between the reference antenna and the network analyzer. | 0.05 | 0.05 | 0.05 | U-shaped | 1.4142 | 1 | 0.035355 | 0.035355 | 0.035355 | 0.00125 | 0.00125 | 0.00125
A1-3 | Quality of quiet zone | 0.1 | 0.1 | 0.1 | Gaussian | 1 | 1 | 0.1 | 0.1 | 0.1 | 0.01 | 0.01 | 0.01
A1-4b | Polarization mismatch between the reference antenna and the receiving antenna | 0.01 | 0.01 | 0.01 | Rectangular | 1.7321 | 1 | 0.0057735 | 0.0057735 | 0.0057735 | 3.3333e-05 | 3.3333e-05 | 3.3333e-05
A1-5 | Mutual coupling between the reference antenna and the receiving antenna | 0 | 0 | 0 | Rectangular | 1.7321 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A1-6 | Phase curvature | 0.05 | 0.05 | 0.05 | Gaussian | 1 | 1 | 0.05 | 0.05 | 0.05 | 0.0025 | 0.0025 | 0.0025
C1-3 | Uncertainty of the network analyzer | 0.13 | 0.2 | 0.2 | Gaussian | 1 | 1 | 0.13 | 0.2 | 0.2 | 0.0169 | 0.04 | 0.04
A1-12 | Influence of the reference antenna feed cable | 0.05 | 0.05 | 0.05 | Rectangular | 1.7321 | 1 | 0.028868 | 0.028868 | 0.028868 | 0.00083333 | 0.00083333 | 0.00083333
A1-13 | Reference antenna feed cable loss measurement uncertainty | 0.06 | 0.06 | 0.06 | Gaussian | 1 | 1 | 0.06 | 0.06 | 0.06 | 0.0036 | 0.0036 | 0.0036
A1-14 | Influence of the receiving antenna feed cable | 0.05 | 0.05 | 0.05 | Rectangular | 1.7321 | 1 | 0.028868 | 0.028868 | 0.028868 | 0.00083333 | 0.00083333 | 0.00083333
C1-4 | Uncertainty of the absolute gain of the reference antenna | 0.50229 | 0.43301 | 0.43301 | Rectangular | 1.7321 | 1 | 0.29 | 0.25 | 0.25 | 0.0841 | 0.0625 | 0.0625
A1-15 | Uncertainty of the absolute gain of the receiving antenna | 0 | 0 | 0 | Rectangular | 1.7321 | 1 | 0 | 0 | 0 | 0 | 0 | 0
Combined standard uncertainty (1σ) [dB] | [7 empty] | 0.44373 | 0.53949 | 0.53949 | 0.44373 | 0.53949 | 0.53949
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [7 empty] | 0.86972 | 1.0574 | 1.0574 | 0.86972 | 1.0574 | 1.0574
TRP Summation error | [7 empty] | 0.75 | 0.75 | 0.75
Total MU | [7 empty] | 1.1484 | 1.2964 | 1.2964 | same as TR 37.843
Table 11.4.3.3-1: CATR MU value derivation for OTA OBUE measurement, FR1
CATR | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz |  |  |  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz
Stage 2: DUT measurement
A2-18 | Misalignment DUT & pointing error for TRP | 0.3 | 0.3 | 0.3 | Rectangular | 1.7321 | 1 | 0.17321 | 0.17321 | 0.17321 | 0.03 | 0.03 | 0.03
C1-1 | RF power measurement equipment (e.g. spectrum analyzer, power meter) | 0.14 | 0.26 | 0.26 | Gaussian | 1 | 1 | 0.14 | 0.26 | 0.26 | 0.0196 | 0.0676 | 0.0676
A2-2a | Standing wave between DUT and test range antenna | 0.21 | 0.21 | 0.21 | U-shaped | 1.4142 | 1 | 0.14849 | 0.14849 | 0.14849 | 0.02205 | 0.02205 | 0.02205
A2-3 | RF leakage (SGH connector terminated & test range antenna connector cable terminated) | 0.0012 | 0.0012 | 0.0012 | Gaussian | 1 | 1 | 0.0012 | 0.0012 | 0.0012 | 1.44e-06 | 1.44e-06 | 1.44e-06
A2-4a | QZ ripple DUT | 0.0928 | 0.0928 | 0.0928 | Gaussian | 1 | 1 | 0.0928 | 0.0928 | 0.0928 | 0.0086118 | 0.0086118 | 0.0086118
A2-12 | Frequency flatness | 0.25 | 0.25 | 0.25 | Gaussian | 1 | 1 | 0.25 | 0.25 | 0.25 | 0.0625 | 0.0625 | 0.0625
Stage 1: Calibration measurement | [10 empty] | 0 | 0 | 0
C1-3 | Uncertainty of the network analyzer | 0.13 | 0.2 | 0.2 | Gaussian | 1 | 1 | 0.13 | 0.2 | 0.2 | 0.0169 | 0.04 | 0.04
A2-5 | Mismatch of receiver chain | 0.127 | 0.325 | 0.325 | U-shaped | 1.4142 | 1 | 0.089803 | 0.22981 | 0.22981 | 0.0080645 | 0.052812 | 0.052812
A2-6 | Insertion loss variation of receiver chain | 0.18 | 0.18 | 0.18 | Rectangular | 1.7321 | 1 | 0.10392 | 0.10392 | 0.10392 | 0.0108 | 0.0108 | 0.0108
A2-3 | RF leakage (SGH connector terminated & test range antenna connector cable terminated) | 0.0012 | 0.0012 | 0.0012 | Gaussian | 1 | 1 | 0.0012 | 0.0012 | 0.0012 | 1.44e-06 | 1.44e-06 | 1.44e-06
A2-7 | Influence of the calibration antenna feed cable: | 0.022 | 0.022 | 0.022 | U-shaped | 1.4142 | 1 | 0.015556 | 0.015556 | 0.015556 | 0.000242 | 0.000242 | 0.000242
C1-4 | Uncertainty of the absolute gain of the reference antenna | 0.50229 | 0.43301 | 0.43301 | Rectangular | 1.7321 | 1 | 0.29 | 0.25 | 0.25 | 0.0841 | 0.0625 | 0.0625
A2-8 | Misalignment positioning system | 0 | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A2-1b | Misalignment of calibration antenna and test range antenna | 0.5 | 0.5 | 0.5 | Exp. normal | 2 | 1 | 0.25 | 0.25 | 0.25 | 0.0625 | 0.0625 | 0.0625
A2-9 | Rotary Joints | 0.048 | 0.048 | 0.048 | U-shaped | 1.4142 | 1 | 0.033941 | 0.033941 | 0.033941 | 0.001152 | 0.001152 | 0.001152
A2-2b | Standing wave between calibration antenna and test range antenna | 0.09 | 0.09 | 0.09 | U-shaped | 1.4142 | 1 | 0.06364 | 0.06364 | 0.06364 | 0.00405 | 0.00405 | 0.00405
A2-4b | QZ ripple calibration antenna | 0.009 | 0.009 | 0.009 | Gaussian | 1 | 1 | 0.009 | 0.009 | 0.009 | 8.1e-05 | 8.1e-05 | 8.1e-05
A2-11 | Switching uncertainty | 0.26 | 0.26 | 0.26 | Rectangular | 1.7321 | 1 | 0.15011 | 0.15011 | 0.15011 | 0.022533 | 0.022533 | 0.022533
Combined standard uncertainty (1σ) [dB] | [7 empty] | 0.5943 | 0.66891 | 0.66891 | 0.5943 | 0.66891 | 0.66891
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [7 empty] | 1.1648 | 1.3111 | 1.3111 | 1.1648 | 1.3111 | 1.3111
TRP Summation error | [7 empty] | 0.75 | 0.75 | 0.75 |  |  |  | Nok: As there is a significant difference between the new and old MU values, further checking is needed. A suggestion is to keep these values in [].
Total MU | [7 empty] | 1.3854 | 1.5104 | 1.5104 | calculation is wrong in 37.843, other than pointing error all contributions same as 37.843
Reference to In-band TRP is used in TR
Near field | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz |  |  |  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz
Stage 2: DUT measurement
A3-1 | Axes Intersection | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-2 | Axes Orthogonality | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-3 | Horizontal Pointing | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-4 | Probe Vertical Position | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-5 | Probe H/V pointing | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-6 | Measurement Distance | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-7 | Amplitude and Phase Drift | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-8 | Amplitude and Phase Noise | 0.02 | 0.02 | 0.02 | Gaussian | 1 | 1 | 0.02 | 0.02 | 0.02 | 0.0004 | 0.0004 | 0.0004
A3-9 | Leakage and Crosstalk | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-10 | Amplitude Non-Linearity | 0.04 | 0.04 | 0.04 | Gaussian | 1 | 1 | 0.04 | 0.04 | 0.04 | 0.0016 | 0.0016 | 0.0016
A3-11 | Amplitude and Phase Shift in rotary joints | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-12 | Channel Balance Amplitude and Phase | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-13 | Probe Polarization Amplitude and Phase | 0.0001 | 0.0001 | 0.0001 | Gaussian | 1 | 1 | 0.0001 | 0.0001 | 0.0001 | 1e-08 | 1e-08 | 1e-08
A3-14 | Probe Pattern Knowledge | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-15 | Multiple Reflections | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-16 | Room Scattering | 0.09 | 0.09 | 0.09 | Gaussian | 1 | 1 | 0.09 | 0.09 | 0.09 | 0.0081 | 0.0081 | 0.0081
A3-17 | DUT support Scattering | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-21 | Positioning | 0.03 | 0.03 | 0.03 | Rectangular | 1.73 | 1 | 0.017341 | 0.017341 | 0.017341 | 0.00030071 | 0.00030071 | 0.00030071
A3-22 | Probe Array Uniformity | 0.055 | 0.055 | 0.055 | Gaussian | 1 | 1 | 0.055 | 0.055 | 0.055 | 0.003025 | 0.003025 | 0.003025
A3-23 | Mismatch of receiver chain | 0.284 | 0.284 | 0.284 | U-Shaped | 1.41 | 1 | 0.20142 | 0.20142 | 0.20142 | 0.040569 | 0.040569 | 0.040569
A3-24 | Insertion loss of receiver chain | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-25 | Uncertainty of the absolute gain of the probe antenna | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
C1-1 | RF power measurement equipment (e.g. spectrum analyzer, power meter) | 0.14 | 0.26 | 0.26 | Gaussian | 1 | 1 | 0.14 | 0.26 | 0.26 | 0.0196 | 0.0676 | 0.0676
A3-26 | Measurement repeatability - positioning repeatability | 0.15 | 0.15 | 0.15 | Gaussian | 1 | 1 | 0.15 | 0.15 | 0.15 | 0.0225 | 0.0225 | 0.0225
A3-33 | Test system frequency flatness | 0.25 | 0.25 | 0.25 | Gaussian | 1 | 1 | 0.25 | 0.25 | 0.25 | 0.0625 | 0.0625 | 0.0625
Stage 1: Calibration measurement
C1-3 | Uncertainty of the network analyzer | 0.13 | 0.2 | 0.2 | Gaussian | 1 | 1 | 0.13 | 0.2 | 0.2 | 0.0169 | 0.04 | 0.04
A3-27 | Mismatch of receiver chain | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-28 | Insertion loss of receiver chain | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-29 | Mismatch in the connection of the calibration antenna | 0.02 | 0.02 | 0.02 | U-Shaped | 1.41 | 1 | 0.014184 | 0.014184 | 0.014184 | 0.0002012 | 0.0002012 | 0.0002012
A3-30 | Influence of the calibration antenna feed cable | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-31 | Influence of the probe antenna cable | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
C1-4 | Uncertainty of the absolute gain of the reference antenna | 0.50229 | 0.43301 | 0.43301 | Rectangular | 1.7321 | 1 | 0.29 | 0.25 | 0.25 | 0.0841 | 0.0625 | 0.0625
A3-32 | Short term repeatability | 0.088 | 0.088 | 0.088 | Gaussian | 1 | 1 | 0.088 | 0.088 | 0.088 | 0.007744 | 0.007744 | 0.007744
Combined standard uncertainty (1σ) [dB] | [7 empty] | 0.51724 | 0.56306 | 0.56306 | 0.51724 | 0.56306 | 0.56306 | Nok: As there is a significant difference between the new and old MU values, further checking is needed. A suggestion is to keep these values in [].
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [7 empty] | 1.0138 | 1.1036 | 1.1036 | 1.0138 | 1.1036 | 1.1036
TRP Summation error | [7 empty] | 0.75 | 0.75 | 0.75 | same as in-band power (as stated in 37.843, although some correctiuons to in-band power mena value is a little different)
Total MU | [7 empty] | 1.2611 | 1.3343 | 1.3343
Reference to In-band TRP is used in TR
Plane wave synthesiser | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz |  |  |  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz
Stage 2: DUT measurement
A7-1a | Misalignment DUT & pointing error | 0.1 | 0.1 | 0.1 | Rectangular | 1.7321 | 1 | 0.057735 | 0.057735 | 0.057735 | 0.0033333 | 0.0033333 | 0.0033333
C1-1 | RF power measurement equipment (e.g. spectrum analyzer, power meter) | 0.14 | 0.26 | 0.26 | Gaussian | 1 | 1 | 0.14 | 0.26 | 0.26 | 0.0196 | 0.0676 | 0.0676
A7-2a | Longitudinal position uncertainty (i.e. standing wave and imperfect field synthesis) for DUT antenna | 0.05 | 0.14 | 0.14 | Rectangular | 1.7321 | 1 | 0.028868 | 0.080829 | 0.080829 | 0.00083333 | 0.0065333 | 0.0065333
A7-3 | RF leakage (calibration antenna connector terminated) | 0.086 | 0.086 | 0.086 | Gaussian | 1 | 1 | 0.086 | 0.086 | 0.086 | 0.007396 | 0.007396 | 0.007396
A7-4a | QZ ripple with DUT | 0.42 | 0.43 | 0.43 | Rectangular | 1.7321 | 1 | 0.24249 | 0.24826 | 0.24826 | 0.0588 | 0.061633 | 0.061633
A7-5 | Miscellaneous Uncertainty | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A7-14 | System non-linearity | 0.062 | 0.062 | 0.062 | Rectangular | 1.7321 | 1 | 0.035796 | 0.035796 | 0.035796 | 0.0012813 | 0.0012813 | 0.0012813
A7-13 | Frequency Flatness | 0.13 | 0.13 | 0.13 | Rectangular | 1.7321 | 1 | 0.075056 | 0.075056 | 0.075056 | 0.0056333 | 0.0056333 | 0.0056333
Stage 1: Calibration measurement | [10 empty] | 0 | 0 | 0
C1-3 | Uncertainty of the network analyzer | 0.13 | 0.2 | 0.2 | Gaussian | 1 | 1 | 0.13 | 0.2 | 0.2 | 0.0169 | 0.04 | 0.04
A7-6 | Mismatch (i.e. reference antenna, network analyser and reference cable) | 0.127 | 0.325 | 0.325 | U-shaped | 1.4142 | 1 | 0.089803 | 0.22981 | 0.22981 | 0.0080645 | 0.052812 | 0.052812
A7-7 | Insertion loss variation | 0.18 | 0.18 | 0.18 | Rectangular | 1.7321 | 1 | 0.10392 | 0.10392 | 0.10392 | 0.0108 | 0.0108 | 0.0108
A7-3 | RF leakage (calibration antenna connector terminated) | 0.086 | 0.086 | 0.086 | Gaussian | 1 | 1 | 0.086 | 0.086 | 0.086 | 0.007396 | 0.007396 | 0.007396
A7-8 | Influence of the calibration antenna feed cable | 0.103 | 0.104 | 0.104 | Rectangular | 1.7321 | 1 | 0.059467 | 0.060044 | 0.060044 | 0.0035363 | 0.0036053 | 0.0036053
C1-4 | Uncertainty of the absolute gain of the reference antenna | 0.50229 | 0.43301 | 0.43301 | Rectangular | 1.7321 | 1 | 0.29 | 0.25 | 0.25 | 0.0841 | 0.0625 | 0.0625
A7-9 | Misalignment of positioning system | 0 | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A7-1b | Misalignment of calibration antenna & pointing error | 0.05 | 0.05 | 0.05 | Rectangular | 1.7321 | 1 | 0.028868 | 0.028868 | 0.028868 | 0.00083333 | 0.00083333 | 0.00083333
A7-10 | Rotary joints | 0 | 0 | 0 | U-shaped | 1.7321 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A7-2b | Longitudinal position uncertainty (i.e. standing wave and imperfect field synthesis) for calibration antenna | 0.12 | 0.12 | 0.12 | Rectangular | 1.7321 | 1 | 0.069282 | 0.069282 | 0.069282 | 0.0048 | 0.0048 | 0.0048
A7-4a | QZ ripple with calibration antenna | 0.2 | 0.2 | 0.2 | Rectangular | 1.7321 | 1 | 0.11547 | 0.11547 | 0.11547 | 0.013333 | 0.013333 | 0.013333
A7-11 | Switching uncertainty | 0.02 | 0.02 | 0.02 | Rectangular | 1.7321 | 1 | 0.011547 | 0.011547 | 0.011547 | 0.00013333 | 0.00013333 | 0.00013333
A7-12 | Field repeatability | 0.06 | 0.12 | 0.12 | Gaussian | 1 | 1 | 0.06 | 0.12 | 0.12 | 0.0036 | 0.0144 | 0.0144
Combined standard uncertainty (1σ) [dB] | [7 empty] | 0.50037 | 0.60334 | 0.60334 | 0.50037 | 0.60334 | 0.60334
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [7 empty] | 0.98073 | 1.1826 | 1.1826 | 0.98073 | 1.1826 | 1.1826
TRP Summation error | [7 empty] | 0.75 | 0.75 | 0.75
Total MU | [7 empty] | 1.2346 | 1.4003 | 1.4003 | same as current R&S sheet
Table 11.4.5.3-1: Reverberation chamber MU value derivation for OTA OBUE
Reverberation Chamber | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz |  |  |  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz
Stage 2: DUT measurement
C1-1 | RF power measurement equipment (e.g. spectrum analyzer, power meter) | 0.14 | 0.26 | 0.26 | Gaussian | 1 | 1 | 0.14 | 0.26 | 0.26 | 0.0196 | 0.0676 | 0.0676
A6-1 | Impedance mismatch in the receiving chain | 0.2 | 0.2 | 0.2 | U-shaped | 1.4142 | 1 | 0.14142 | 0.14142 | 0.14142 | 0.02 | 0.02 | 0.02
A6-2 | Random uncertainty | 0.1 | 0.1 | 0.1 | Rectangular | 1.7321 | 1 | 0.057735 | 0.057735 | 0.057735 | 0.0033333 | 0.0033333 | 0.0033333
Stage 1: Calibration measurement
A6-3 | Reference antenna radiation efficiency | 0.2 | 0.2 | 0.2 | Gaussian | 1 | 1 | 0.2 | 0.2 | 0.2 | 0.04 | 0.04 | 0.04
A6-4 | Mean value estimation of reference antenna radiation efficiency | 0.15 | 0.15 | 0.15 | Gaussian | 1 | 1 | 0.15 | 0.15 | 0.15 | 0.0225 | 0.0225 | 0.0225
C1-3 | Uncertainty of the Network Analyzer | 0.2 | 0.2 | 0.2 | Gaussian | 1 | 1 | 0.2 | 0.2 | 0.2 | 0.04 | 0.04 | 0.04
A6-5 | Influence of the reference antenna feed cable | 0.2 | 0.2 | 0.2 | Gaussian | 1 | 1 | 0.2 | 0.2 | 0.2 | 0.04 | 0.04 | 0.04
A6-6 | Mean value estimation of transfer function | 0.27 | 0.27 | 0.27 | Gaussian | 1 | 1 | 0.27 | 0.27 | 0.27 | 0.0729 | 0.0729 | 0.0729
A6-7 | Uniformity of transfer function | 0.5 | 0.5 | 0.5 | Gaussian | 1 | 1 | 0.5 | 0.5 | 0.5 | 0.25 | 0.25 | 0.25
Combined standard uncertainty (1σ) [dB] | [7 empty] | 0.71297 | 0.74588 | 0.74588 | 0.71297 | 0.74588 | 0.74588
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [7 empty] | 1.3974 | 1.4619 | 1.4619 | 1.3974 | 1.4619 | 1.4619
same as 37.843
check 28/1/20
IAC-shoudl FF be included?
CATR- pointing error? calc error in 37.843-
check 11/2/20
NF - removed A3-18-20 as per email 11/2/20

## SEM
| chamber | Expanded uncertainty [dB] | [11 empty] | Review comments (RAN4#94)
|  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz
| Indoor anechoic | 1.1484 | 1.2964 | 1.2964
| CATR | 1.3854 | 1.5104 | 1.5104
| Near field | 1.2611 | 1.3343 | 1.3343
1D CATR
| Plane wave sythesiser | 1.2346 | 1.4003 | 1.4003
| Reverberation Chamber | 1.3974 | 1.4619 | 1.4619
| Common maximum accepted test system uncertainty | 1.8 | 2 | 2 | *note 38.817-02 has an error in this section the values are 1.4 and 1.5 but 38.141-2 is ok
Table 11.5.2.3-1: IAC MU value derivation for for OTA SEM
Indoor anechoic | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz |  |  |  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz
Stage 2: DUT measurement
A1-1 | Positioning misalignment between the AAS BS and the reference antenna | 0.03 | 0.03 | 0.03 | Rectangular | 1.7321 | 1 | 0.017321 | 0.017321 | 0.017321 | 0.0003 | 0.0003 | 0.0003
A1-2 | Pointing misalignment between the AAS BS and the receiving antenna | 0.3 | 0.3 | 0.3 | Rectangular | 1.7321 | 1 | 0.17321 | 0.17321 | 0.17321 | 0.03 | 0.03 | 0.03
A1-3 | Quality of quiet zone | 0.1 | 0.1 | 0.1 | Gaussian | 1 | 1 | 0.1 | 0.1 | 0.1 | 0.01 | 0.01 | 0.01
A1-4a | Polarization mismatch between the AAS BS and the receiving antenna | 0.01 | 0.01 | 0.01 | Rectangular | 1.7321 | 1 | 0.0057735 | 0.0057735 | 0.0057735 | 3.3333e-05 | 3.3333e-05 | 3.3333e-05
A1-5 | Mutual coupling between the AAS BS and the receiving antenna | 0 | 0 | 0 | Rectangular | 1.7321 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A1-6 | Phase curvature | 0.05 | 0.05 | 0.05 | Gaussian | 1 | 1 | 0.05 | 0.05 | 0.05 | 0.0025 | 0.0025 | 0.0025
C1-1 | RF power measurement equipment (e.g. spectrum analyzer, power meter) | 0.14 | 0.26 | 0.26 | Gaussian | 1 | 1 | 0.14 | 0.26 | 0.26 | 0.0196 | 0.0676 | 0.0676
A1-7 | Impedance mismatch in the receiving chain | 0.14 | 0.33 | 0.33 | U-shaped | 1.4142 | 1 | 0.098995 | 0.23335 | 0.23335 | 0.0098 | 0.05445 | 0.05445
A1-8 | Random uncertainty | 0.1 | 0.1 | 0.1 | Rectangular | 1.7321 | 1 | 0.057735 | 0.057735 | 0.057735 | 0.0033333 | 0.0033333 | 0.0033333
Stage 1: Calibration measurement | [10 empty] | 0 | 0 | 0
A1-9 | Impedance mismatch between the receiving antenna and the network analyzer | 0.05 | 0.05 | 0.05 | U-shaped | 1.4142 | 1 | 0.035355 | 0.035355 | 0.035355 | 0.00125 | 0.00125 | 0.00125
A1-10 | Positioning and pointing misalignment between the reference antenna and the receiving antenna | 0.01 | 0.01 | 0.01 | Rectangular | 1.7321 | 1 | 0.0057735 | 0.0057735 | 0.0057735 | 3.3333e-05 | 3.3333e-05 | 3.3333e-05
A1-11 | Impedance mismatch between the reference antenna and the network analyzer. | 0.05 | 0.05 | 0.05 | U-shaped | 1.4142 | 1 | 0.035355 | 0.035355 | 0.035355 | 0.00125 | 0.00125 | 0.00125
A1-3 | Quality of quiet zone | 0.1 | 0.1 | 0.1 | Gaussian | 1 | 1 | 0.1 | 0.1 | 0.1 | 0.01 | 0.01 | 0.01
A1-4b | Polarization mismatch between the reference antenna and the receiving antenna | 0.01 | 0.01 | 0.01 | Rectangular | 1.7321 | 1 | 0.0057735 | 0.0057735 | 0.0057735 | 3.3333e-05 | 3.3333e-05 | 3.3333e-05
A1-5 | Mutual coupling between the reference antenna and the receiving antenna | 0 | 0 | 0 | Rectangular | 1.7321 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A1-6 | Phase curvature | 0.05 | 0.05 | 0.05 | Gaussian | 1 | 1 | 0.05 | 0.05 | 0.05 | 0.0025 | 0.0025 | 0.0025
C1-3 | Uncertainty of the network analyzer | 0.13 | 0.2 | 0.2 | Gaussian | 1 | 1 | 0.13 | 0.2 | 0.2 | 0.0169 | 0.04 | 0.04
A1-12 | Influence of the reference antenna feed cable | 0.05 | 0.05 | 0.05 | Rectangular | 1.7321 | 1 | 0.028868 | 0.028868 | 0.028868 | 0.00083333 | 0.00083333 | 0.00083333
A1-13 | Reference antenna feed cable loss measurement uncertainty | 0.06 | 0.06 | 0.06 | Gaussian | 1 | 1 | 0.06 | 0.06 | 0.06 | 0.0036 | 0.0036 | 0.0036
A1-14 | Influence of the receiving antenna feed cable | 0.05 | 0.05 | 0.05 | Rectangular | 1.7321 | 1 | 0.028868 | 0.028868 | 0.028868 | 0.00083333 | 0.00083333 | 0.00083333
C1-4 | Uncertainty of the absolute gain of the reference antenna | 0.50229 | 0.43301 | 0.43301 | Rectangular | 1.7321 | 1 | 0.29 | 0.25 | 0.25 | 0.0841 | 0.0625 | 0.0625
A1-15 | Uncertainty of the absolute gain of the receiving antenna | 0 | 0 | 0 | Rectangular | 1.7321 | 1 | 0 | 0 | 0 | 0 | 0 | 0
Combined standard uncertainty (1σ) [dB] | [7 empty] | 0.44373 | 0.53949 | 0.53949 | 0.44373 | 0.53949 | 0.53949
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [7 empty] | 0.86972 | 1.0574 | 1.0574 | 0.86972 | 1.0574 | 1.0574
TRP Summation error | [7 empty] | 0.75 | 0.75 | 0.75
Total MU | [7 empty] | 1.1484 | 1.2964 | 1.2964
Table 11.5.3.3-1: CATR MU value derivation for OTA SEM
CATR | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz |  |  |  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz
Stage 2: DUT measurement
A2-18 | Misalignment DUT & pointing error for TRP | 0.3 | 0.3 | 0.3 | Rectangular | 1.7321 | 1 | 0.17321 | 0.17321 | 0.17321 | 0.03 | 0.03 | 0.03
C1-1 | RF power measurement equipment (e.g. spectrum analyzer, power meter) | 0.14 | 0.26 | 0.26 | Gaussian | 1 | 1 | 0.14 | 0.26 | 0.26 | 0.0196 | 0.0676 | 0.0676
A2-2a | Standing wave between DUT and test range antenna | 0.21 | 0.21 | 0.21 | U-shaped | 1.4142 | 1 | 0.14849 | 0.14849 | 0.14849 | 0.02205 | 0.02205 | 0.02205
A2-3 | RF leakage (SGH connector terminated & test range antenna connector cable terminated) | 0.0012 | 0.0012 | 0.0012 | Gaussian | 1 | 1 | 0.0012 | 0.0012 | 0.0012 | 1.44e-06 | 1.44e-06 | 1.44e-06
A2-4a | QZ ripple DUT | 0.0928 | 0.0928 | 0.0928 | Gaussian | 1 | 1 | 0.0928 | 0.0928 | 0.0928 | 0.0086118 | 0.0086118 | 0.0086118
A2-12 | Frequency flatness | 0.25 | 0.25 | 0.25 | Gaussian | 1 | 1 | 0.25 | 0.25 | 0.25 | 0.0625 | 0.0625 | 0.0625
Stage 1: Calibration measurement | [10 empty] | 0 | 0 | 0
C1-3 | Uncertainty of the network analyzer | 0.13 | 0.2 | 0.2 | Gaussian | 1 | 1 | 0.13 | 0.2 | 0.2 | 0.0169 | 0.04 | 0.04
A2-5 | Mismatch of receiver chain | 0.127 | 0.325 | 0.325 | U-shaped | 1.4142 | 1 | 0.089803 | 0.22981 | 0.22981 | 0.0080645 | 0.052812 | 0.052812
A2-6 | Insertion loss variation of receiver chain | 0.18 | 0.18 | 0.18 | Rectangular | 1.7321 | 1 | 0.10392 | 0.10392 | 0.10392 | 0.0108 | 0.0108 | 0.0108
A2-3 | RF leakage (SGH connector terminated & test range antenna connector cable terminated) | 0.0012 | 0.0012 | 0.0012 | Gaussian | 1 | 1 | 0.0012 | 0.0012 | 0.0012 | 1.44e-06 | 1.44e-06 | 1.44e-06
A2-7 | Influence of the calibration antenna feed cable: | 0.022 | 0.022 | 0.022 | U-shaped | 1.4142 | 1 | 0.015556 | 0.015556 | 0.015556 | 0.000242 | 0.000242 | 0.000242
C1-4 | Uncertainty of the absolute gain of the reference antenna | 0.50229 | 0.43301 | 0.43301 | Rectangular | 1.7321 | 1 | 0.29 | 0.25 | 0.25 | 0.0841 | 0.0625 | 0.0625
A2-8 | Misalignment positioning system | 0 | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A2-1b | Misalignment of calibration antenna and test range antenna | 0.5 | 0.5 | 0.5 | Exp. normal | 2 | 1 | 0.25 | 0.25 | 0.25 | 0.0625 | 0.0625 | 0.0625
A2-9 | Rotary Joints | 0.048 | 0.048 | 0.048 | U-shaped | 1.4142 | 1 | 0.033941 | 0.033941 | 0.033941 | 0.001152 | 0.001152 | 0.001152
A2-2b | Standing wave between calibration antenna and test range antenna | 0.09 | 0.09 | 0.09 | U-shaped | 1.4142 | 1 | 0.06364 | 0.06364 | 0.06364 | 0.00405 | 0.00405 | 0.00405
A2-4b | QZ ripple calibration antenna | 0.009 | 0.009 | 0.009 | Gaussian | 1 | 1 | 0.009 | 0.009 | 0.009 | 8.1e-05 | 8.1e-05 | 8.1e-05
A2-11 | Switching uncertainty | 0.26 | 0.26 | 0.26 | Rectangular | 1.7321 | 1 | 0.15011 | 0.15011 | 0.15011 | 0.022533 | 0.022533 | 0.022533
Combined standard uncertainty (1σ) [dB] | [7 empty] | 0.5943 | 0.66891 | 0.66891 | 0.5943 | 0.66891 | 0.66891
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [7 empty] | 1.1648 | 1.3111 | 1.3111 | 1.1648 | 1.3111 | 1.3111
TRP Summation error | [7 empty] | 0.75 | 0.75 | 0.75
Total MU | [7 empty] | 1.3854 | 1.5104 | 1.5104
Reference to In-band TRP is used in TR
Near field | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz |  |  |  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz
Stage 2: DUT measurement
A3-1 | Axes Intersection | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-2 | Axes Orthogonality | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-3 | Horizontal Pointing | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-4 | Probe Vertical Position | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-5 | Probe H/V pointing | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-6 | Measurement Distance | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-7 | Amplitude and Phase Drift | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-8 | Amplitude and Phase Noise | 0.02 | 0.02 | 0.02 | Gaussian | 1 | 1 | 0.02 | 0.02 | 0.02 | 0.0004 | 0.0004 | 0.0004
A3-9 | Leakage and Crosstalk | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-10 | Amplitude Non-Linearity | 0.04 | 0.04 | 0.04 | Gaussian | 1 | 1 | 0.04 | 0.04 | 0.04 | 0.0016 | 0.0016 | 0.0016
A3-11 | Amplitude and Phase Shift in rotary joints | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-12 | Channel Balance Amplitude and Phase | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-13 | Probe Polarization Amplitude and Phase | 0.0001 | 0.0001 | 0.0001 | Gaussian | 1 | 1 | 0.0001 | 0.0001 | 0.0001 | 1e-08 | 1e-08 | 1e-08
A3-14 | Probe Pattern Knowledge | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-15 | Multiple Reflections | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-16 | Room Scattering | 0.09 | 0.09 | 0.09 | Gaussian | 1 | 1 | 0.09 | 0.09 | 0.09 | 0.0081 | 0.0081 | 0.0081
A3-17 | DUT support Scattering | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-21 | Positioning | 0.03 | 0.03 | 0.03 | Rectangular | 1.73 | 1 | 0.017341 | 0.017341 | 0.017341 | 0.00030071 | 0.00030071 | 0.00030071
A3-22 | Probe Array Uniformity | 0.055 | 0.055 | 0.055 | Gaussian | 1 | 1 | 0.055 | 0.055 | 0.055 | 0.003025 | 0.003025 | 0.003025
A3-23 | Mismatch of receiver chain | 0.284 | 0.284 | 0.284 | U-Shaped | 1.41 | 1 | 0.20142 | 0.20142 | 0.20142 | 0.040569 | 0.040569 | 0.040569
A3-24 | Insertion loss of receiver chain | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-25 | Uncertainty of the absolute gain of the probe antenna | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
C1-1 | RF power measurement equipment (e.g. spectrum analyzer, power meter) | 0.14 | 0.26 | 0.26 | Gaussian | 1 | 1 | 0.14 | 0.26 | 0.26 | 0.0196 | 0.0676 | 0.0676
A3-26 | Measurement repeatability - positioning repeatability | 0.15 | 0.15 | 0.15 | Gaussian | 1 | 1 | 0.15 | 0.15 | 0.15 | 0.0225 | 0.0225 | 0.0225
A3-33 | Test system frequency flatness | 0.25 | 0.25 | 0.25 | Gaussian | 1 | 1 | 0.25 | 0.25 | 0.25 | 0.0625 | 0.0625 | 0.0625
Stage 1: Calibration measurement
C1-3 | Uncertainty of the network analyzer | 0.13 | 0.2 | 0.2 | Gaussian | 1 | 1 | 0.13 | 0.2 | 0.2 | 0.0169 | 0.04 | 0.04
A3-27 | Mismatch of receiver chain | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-28 | Insertion loss of receiver chain | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-29 | Mismatch in the connection of the calibration antenna | 0.02 | 0.02 | 0.02 | U-Shaped | 1.41 | 1 | 0.014184 | 0.014184 | 0.014184 | 0.0002012 | 0.0002012 | 0.0002012
A3-30 | Influence of the calibration antenna feed cable | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A3-31 | Influence of the probe antenna cable | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
C1-4 | Uncertainty of the absolute gain of the reference antenna | 0.50229 | 0.43301 | 0.43301 | Rectangular | 1.7321 | 1 | 0.29 | 0.25 | 0.25 | 0.0841 | 0.0625 | 0.0625
A3-32 | Short term repeatability | 0.088 | 0.088 | 0.088 | Gaussian | 1 | 1 | 0.088 | 0.088 | 0.088 | 0.007744 | 0.007744 | 0.007744
Combined standard uncertainty (1σ) [dB] | [7 empty] | 0.51724 | 0.56306 | 0.56306 | 0.51724 | 0.56306 | 0.56306
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [7 empty] | 1.0138 | 1.1036 | 1.1036 | 1.0138 | 1.1036 | 1.1036
TRP Summation error | [7 empty] | 0.75 | 0.75 | 0.75
Total MU | [7 empty] | 1.2611 | 1.3343 | 1.3343
Stage 1: Calibration measurement | [10 empty] | 0 | 0 | 0
Reference to In-band TRP is used in TR
Plane wave synthesiser | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz |  |  |  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz
Stage 2: DUT measurement
A7-1a | Misalignment DUT & pointing error | 0.1 | 0.1 | 0.1 | Rectangular | 1.7321 | 1 | 0.057735 | 0.057735 | 0.057735 | 0.0033333 | 0.0033333 | 0.0033333
C1-1 | RF power measurement equipment (e.g. spectrum analyzer, power meter) | 0.14 | 0.26 | 0.26 | Gaussian | 1 | 1 | 0.14 | 0.26 | 0.26 | 0.0196 | 0.0676 | 0.0676
A7-2a | Longitudinal position uncertainty (i.e. standing wave and imperfect field synthesis) for DUT antenna | 0.05 | 0.14 | 0.14 | Rectangular | 1.7321 | 1 | 0.028868 | 0.080829 | 0.080829 | 0.00083333 | 0.0065333 | 0.0065333
A7-3 | RF leakage (calibration antenna connector terminated) | 0.086 | 0.086 | 0.086 | Gaussian | 1 | 1 | 0.086 | 0.086 | 0.086 | 0.007396 | 0.007396 | 0.007396
A7-4a | QZ ripple with DUT | 0.42 | 0.43 | 0.43 | Rectangular | 1.7321 | 1 | 0.24249 | 0.24826 | 0.24826 | 0.0588 | 0.061633 | 0.061633
A7-5 | Miscellaneous Uncertainty | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A7-14 | System non-linearity | 0.062 | 0.062 | 0.062 | Rectangular | 1.7321 | 1 | 0.035796 | 0.035796 | 0.035796 | 0.0012813 | 0.0012813 | 0.0012813
A7-13 | Frequency Flatness | 0.13 | 0.13 | 0.13 | Rectangular | 1.7321 | 1 | 0.075056 | 0.075056 | 0.075056 | 0.0056333 | 0.0056333 | 0.0056333
Stage 1: Calibration measurement | [10 empty] | 0 | 0 | 0
C1-3 | Uncertainty of the network analyzer | 0.13 | 0.2 | 0.2 | Gaussian | 1 | 1 | 0.13 | 0.2 | 0.2 | 0.0169 | 0.04 | 0.04
A7-6 | Mismatch (i.e. reference antenna, network analyser and reference cable) | 0.127 | 0.325 | 0.325 | U-shaped | 1.4142 | 1 | 0.089803 | 0.22981 | 0.22981 | 0.0080645 | 0.052812 | 0.052812
A7-7 | Insertion loss variation | 0.18 | 0.18 | 0.18 | Rectangular | 1.7321 | 1 | 0.10392 | 0.10392 | 0.10392 | 0.0108 | 0.0108 | 0.0108
A7-3 | RF leakage (calibration antenna connector terminated) | 0.086 | 0.086 | 0.086 | Gaussian | 1 | 1 | 0.086 | 0.086 | 0.086 | 0.007396 | 0.007396 | 0.007396
A7-8 | Influence of the calibration antenna feed cable | 0.103 | 0.104 | 0.104 | Rectangular | 1.7321 | 1 | 0.059467 | 0.060044 | 0.060044 | 0.0035363 | 0.0036053 | 0.0036053
C1-4 | Uncertainty of the absolute gain of the reference antenna | 0.50229 | 0.43301 | 0.43301 | Rectangular | 1.7321 | 1 | 0.29 | 0.25 | 0.25 | 0.0841 | 0.0625 | 0.0625
A7-9 | Misalignment of positioning system | 0 | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A7-1b | Misalignment of calibration antenna & pointing error | 0.05 | 0.05 | 0.05 | Rectangular | 1.7321 | 1 | 0.028868 | 0.028868 | 0.028868 | 0.00083333 | 0.00083333 | 0.00083333
A7-10 | Rotary joints | 0 | 0 | 0 | U-shaped | 1.7321 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A7-2b | Longitudinal position uncertainty (i.e. standing wave and imperfect field synthesis) for calibration antenna | 0.12 | 0.12 | 0.12 | Rectangular | 1.7321 | 1 | 0.069282 | 0.069282 | 0.069282 | 0.0048 | 0.0048 | 0.0048
A7-4a | QZ ripple with calibration antenna | 0.2 | 0.2 | 0.2 | Rectangular | 1.7321 | 1 | 0.11547 | 0.11547 | 0.11547 | 0.013333 | 0.013333 | 0.013333
A7-11 | Switching uncertainty | 0.02 | 0.02 | 0.02 | Rectangular | 1.7321 | 1 | 0.011547 | 0.011547 | 0.011547 | 0.00013333 | 0.00013333 | 0.00013333
A7-12 | Field repeatability | 0.06 | 0.12 | 0.12 | Gaussian | 1 | 1 | 0.06 | 0.12 | 0.12 | 0.0036 | 0.0144 | 0.0144
Combined standard uncertainty (1σ) [dB] | [7 empty] | 0.50037 | 0.60334 | 0.60334 | 0.50037 | 0.60334 | 0.60334
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [7 empty] | 0.98073 | 1.1826 | 1.1826 | 0.98073 | 1.1826 | 1.1826
TRP Summation error | [7 empty] | 0.75 | 0.75 | 0.75
Total MU | [7 empty] | 1.2346 | 1.4003 | 1.4003
Table 11.5.5.3-1: Reverberation chamber MU value derivation for OTA SEM
Reverberation Chamber | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz |  |  |  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz
Stage 2: DUT measurement
C1-1 | RF power measurement equipment (e.g. spectrum analyzer, power meter) | 0.14 | 0.26 | 0.26 | Gaussian | 1 | 1 | 0.14 | 0.26 | 0.26 | 0.0196 | 0.0676 | 0.0676
A6-1 | Impedance mismatch in the receiving chain | 0.2 | 0.2 | 0.2 | U-shaped | 1.4142 | 1 | 0.14142 | 0.14142 | 0.14142 | 0.02 | 0.02 | 0.02
A6-2 | Random uncertainty | 0.1 | 0.1 | 0.1 | Rectangular | 1.7321 | 1 | 0.057735 | 0.057735 | 0.057735 | 0.0033333 | 0.0033333 | 0.0033333
Stage 1: Calibration measurement
A6-3 | Reference antenna radiation efficiency | 0.2 | 0.2 | 0.2 | Gaussian | 1 | 1 | 0.2 | 0.2 | 0.2 | 0.04 | 0.04 | 0.04
A6-4 | Mean value estimation of reference antenna radiation efficiency | 0.15 | 0.15 | 0.15 | Gaussian | 1 | 1 | 0.15 | 0.15 | 0.15 | 0.0225 | 0.0225 | 0.0225
C1-3 | Uncertainty of the Network Analyzer | 0.2 | 0.2 | 0.2 | Gaussian | 1 | 1 | 0.2 | 0.2 | 0.2 | 0.04 | 0.04 | 0.04
A6-5 | Influence of the reference antenna feed cable | 0.2 | 0.2 | 0.2 | Gaussian | 1 | 1 | 0.2 | 0.2 | 0.2 | 0.04 | 0.04 | 0.04
A6-6 | Mean value estimation of transfer function | 0.27 | 0.27 | 0.27 | Gaussian | 1 | 1 | 0.27 | 0.27 | 0.27 | 0.0729 | 0.0729 | 0.0729
A6-7 | Uniformity of transfer function | 0.5 | 0.5 | 0.5 | Gaussian | 1 | 1 | 0.5 | 0.5 | 0.5 | 0.25 | 0.25 | 0.25
Combined standard uncertainty (1σ) [dB] | [7 empty] | 0.71297 | 0.74588 | 0.74588 | 0.71297 | 0.74588 | 0.74588
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [7 empty] | 1.3974 | 1.4619 | 1.4619 | 1.3974 | 1.4619 | 1.4619
Notes
Thsi is same as OBUE - thay can probably be merged
check 11/2/20
NF - removed A3-18-20 as per email 11/2/20

## OOB EM
| chamber | Expanded uncertainty [dB] | [17 empty] | Review comments (RAN4#94)
|  | 30MHz<f≤6 GHz | 6GHz<f ≤19GHz | 19GHz<f ≤26GHz
| General Directional Chamber | 2.2939 | 4.1358 | 4.1358
| Reverberation Chamber | 3.2895 | 3.3763 | 3.3763
| Common maximum accepted test system uncertainty | 2.3 | 4.2 | 4.2
Table 12.2.2.3-1: General directional chamber MU value derivation for the TX spurious emissions, FR1
General Directional Chamber | [14 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | 30MHz<f≤6 GHz | 6GHz<f ≤19GHz | 19GHz<f ≤26GHz |  |  |  | 30MHz<f≤6 GHz | 6GHz<f ≤19GHz | 19GHz<f ≤26GHz
Stage 2: DUT measurement
A5-1 | Positioning misalignment between the AAS BS and the reference antenna | 0.03 | 0.03 | 0.03 | Rectangular | 1.7321 | 1 | 0.017321 | 0.017321 | 0.017321 |  |  |  |  | 0.0003 | 0.0003 | 0.0003
A5-2 | Pointing misalignment between the AAS BS and the receiving antenna | 0 | 0 | 0 | Rectangular | 1.7321 | 1 | 0 | 0 | 0 |  |  |  |  | 0 | 0 | 0
A5-3 | Quality of quiet zone | 0.1 | 0.1 | 0.1 | Gaussian | 1 | 1 | 0.1 | 0.1 | 0.1 |  |  |  |  | 0.01 | 0.01 | 0.01
A5-4a | Polarization mismatch between the AAS BS and the receiving antenna | 0.01 | 0.01 | 0.01 | Rectangular | 1.7321 | 1 | 0.0057735 | 0.0057735 | 0.0057735 |  |  |  |  | 3.3333e-05 | 3.3333e-05 | 3.3333e-05
A5-5a | Mutual coupling between the AAS BS and the receiving antenna | 0 | 0 | 0 | Rectangular | 1.7321 | 1 | 0 | 0 | 0 |  |  |  |  | 0 | 0 | 0
A5-6a | Phase curvature | 0.05 | 0.05 | 0.05 | Gaussian | 1 | 1 | 0.05 | 0.05 | 0.05 |  |  |  |  | 0.0025 | 0.0025 | 0.0025
C3-3 | Transmitter mandatory spurious emissions - Conducted Uncertainty (minus miss match) | 1 | 2 | 2 | Gaussian | 1 | 1 | 1 | 2 | 2 |  |  |  |  | 1 | 4 | 4
A5-7 | Impedance mismatch in the receiving chain | 0.2 | 0.45 | 0.45 | U-shaped | 1.4142 | 1 | 0.14142 | 0.3182 | 0.3182 |  |  |  |  | 0.02 | 0.10125 | 0.10125
A5-8 | Random uncertainty | 0.1 | 0.1 | 0.1 | Rectangular | 1.7321 | 1 | 0.057735 | 0.057735 | 0.057735 |  |  |  |  | 0.0033333 | 0.0033333 | 0.0033333
A5-17 | Measurement antenna frequency variation | 0.1 | 0.1 | 0.1 | Rectangular | 1.7321 | 1 | 0.057735 | 0.057735 | 0.057735 |  |  |  |  | 0.0033333 | 0.0033333 | 0.0033333
A5-18 | FSPL estimation error | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 |  |  |  |  | 0 | 0 | 0
A5-16 | Test system frequency flatness | 0.25 | 0.25 | 0.25 | Gaussian | 1 | 1 | 0.25 | 0.25 | 0.25 |  |  |  |  | 0.0625 | 0.0625 | 0.0625
Stage 1: Calibration measurement | [14 empty] | 0 | 0 | 0
A5-9 | Impedance mismatch between the receiving antenna and the network analyzer | 0.05 | 0.05 | 0.05 | U-shaped | 1.4142 | 1 | 0.035355 | 0.035355 | 0.035355 |  |  |  |  | 0.00125 | 0.00125 | 0.00125
A5-10 | Positioning and pointing misalignment between the reference antenna and the receiving antenna | 0.01 | 0.01 | 0.01 | Rectangular | 1.7321 | 1 | 0.0057735 | 0.0057735 | 0.0057735 |  |  |  |  | 3.3333e-05 | 3.3333e-05 | 3.3333e-05
A5-11 | Impedance mismatch between the reference antenna and the network analyzer. | 0.05 | 0.05 | 0.05 | U-shaped | 1.4142 | 1 | 0.035355 | 0.035355 | 0.035355 |  |  |  |  | 0.00125 | 0.00125 | 0.00125
A5-3 | Quality of quiet zone | 0.1 | 0.1 | 0.1 | Gaussian | 1 | 1 | 0.1 | 0.1 | 0.1 |  |  |  |  | 0.01 | 0.01 | 0.01
A5-4b | Polarization mismatch for reference antenna | 0.01 | 0.01 | 0.01 | Rectangular | 1.7321 | 1 | 0.0057735 | 0.0057735 | 0.0057735 |  |  |  |  | 3.3333e-05 | 3.3333e-05 | 3.3333e-05
A5-5b | Mutual coupling between the reference antenna and the receiving antenna | 0 | 0 | 0 | Rectangular | 1.7321 | 1 | 0 | 0 | 0 |  |  |  |  | 0 | 0 | 0
A5-6b | Phase curvature | 0.05 | 0.05 | 0.05 | Gaussian | 1 | 1 | 0.05 | 0.05 | 0.05 |  |  |  |  | 0.0025 | 0.0025 | 0.0025
C1-3 | Uncertainty of the network analyzer | 0.13 | 0.2 | 0.2 | Gaussian | 1 | 1 | 0.13 | 0.2 | 0.2 |  |  |  |  | 0.0169 | 0.04 | 0.04
A5-12 | Influence of the reference antenna feed cable | 0.05 | 0.05 | 0.05 | Rectangular | 1.7321 | 1 | 0.028868 | 0.028868 | 0.028868 |  |  |  |  | 0.00083333 | 0.00083333 | 0.00083333
A5-13 | Reference antenna feed cable loss measurement uncertainty | 0.06 | 0.06 | 0.06 | Gaussian | 1 | 1 | 0.06 | 0.06 | 0.06 |  |  |  |  | 0.0036 | 0.0036 | 0.0036
A5-14 | Influence of the receiving antenna feed cable | 0.05 | 0.05 | 0.05 | Rectangular | 1.7321 | 1 | 0.028868 | 0.028868 | 0.028868 |  |  |  |  | 0.00083333 | 0.00083333 | 0.00083333
C1-4 | Uncertainty of the absolute gain of the reference antenna | 0.50229 | 0.43301 | 0.43301 | Rectangular | 1.7321 | 1 | 0.29 | 0.25 | 0.25 |  |  |  |  | 0.0841 | 0.0625 | 0.0625
A5-15 | Uncertainty of the absolute gain of the receiving antenna | 0 | 0 | 0 | Rectangular | 1.7321 | 1 | 0 | 0 | 0 |  |  |  |  | 0 | 0 | 0
Combined standard uncertainty (1σ) [dB] | [7 empty] | 1.106 | 2.0751 | 2.0751 |  |  |  |  | 1.106 | 2.0751 | 2.0751
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [7 empty] | 2.1678 | 4.0672 | 4.0672 |  |  |  |  | 2.1678 | 4.0672 | 4.0672
TRP Summation error | [7 empty] | 0.75 | 0.75 | 0.75
Total MU | [7 empty] | 2.2939 | 4.1358 | 4.1358
same as 37.843
Table 12.2.3.4-1: Reverberation Chamber MUperpoint value derivation for TX spurious emissions, 380 MHz – 198 GHz
Reverberation Chamber (erros are same definitions as for in-bandvalues are consistant but wider freq range) | [14 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  |  |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | 380 MHz < f ≦ 3 GHz | 3 GHz < f ≦ 6 GHz | 6 GHz < f ≦ 12.75 GHz | 12.75 GHz < f ≦ 19 GHz | 19GHz<f ≤26GHz |  |  |  | 380 MHz < f ≦ 3 GHz | 3 GHz < f ≦ 6 GHz | 6 GHz < f ≦ 12.75 GHz | 12.75 GHz < f ≦ 19 GHz | 19GHz<f ≤26GHz
Stage 2: DUT measurement
C1-1 | RF power measurement equipment (e.g. spectrum analyzer, power meter) | 0.14 | 0.26 | 0.26 | 0.37 | 0.37 | Gaussian | 1 | 1 | 0.14 | 0.26 | 0.26 | 0.37 | 0.37 | 0.0196 | 0.0676 | 0.0676 | 0.1369 | 0.1369
A6-1 | Impedance mismatch in the receiving chain | 0.2 | 0.2 | 0.45 | 0.45 | 0.45 | U-shaped | 1.4142 | 1 | 0.14142 | 0.14142 | 0.3182 | 0.3182 | 0.3182 | 0.02 | 0.02 | 0.10125 | 0.10125 | 0.10125
A6-2 | Random uncertainty | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | Rectangular | 1.7321 | 1 | 0.057735 | 0.057735 | 0.057735 | 0.057735 | 0.057735 | 0.0033333 | 0.0033333 | 0.0033333 | 0.0033333 | 0.0033333
Stage 1: Calibration measurement | [14 empty] | 0 | 0 | 0 | 0 | 0
A6-3 | Reference antenna radiation efficiency | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | Gaussian | 1 | 1 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25
A6-4 | Mean value estimation of reference antenna radiation efficiency | 0.27 | 0.27 | 0.27 | 0.27 | 0.27 | Gaussian | 1 | 1 | 0.27 | 0.27 | 0.27 | 0.27 | 0.27 | 0.0729 | 0.0729 | 0.0729 | 0.0729 | 0.0729
C1-3 | Uncertainty of the network analyzer | 0.13 | 0.2 | 0.2 | 0.2 | 0.2 | Gaussian | 1 | 1 | 0.13 | 0.2 | 0.2 | 0.2 | 0.2 | 0.0169 | 0.04 | 0.04 | 0.04 | 0.04
A6-5 | Influence of the reference antenna feed cable | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | Gaussian | 1 | 1 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04
A6-6 | Mean value estimation of transfer function | 0.27 | 0.27 | 0.27 | 0.27 | 0.27 | Gaussian | 1 | 1 | 0.27 | 0.27 | 0.27 | 0.27 | 0.27 | 0.0729 | 0.0729 | 0.0729 | 0.0729 | 0.0729
A6-7 | Uniformity of transfer function | 1.5 | 1.5 | 1.5 | 1.5 | 1.5 | Gaussian | 1 | 1 | 1.5 | 1.5 | 1.5 | 1.5 | 1.5 | 2.25 | 2.25 | 2.25 | 2.25 | 2.25
Combined standard uncertainty (1σ) [dB] | [9 empty] | 1.657 | 1.6783 | 1.7023 | 1.7226 | 1.7226 | 1.657 | 1.6783 | 1.7023 | 1.7226 | 1.7226
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [9 empty] | 3.2477 | 3.2895 | 3.3366 | 3.3763 | 3.3763 | 3.2477 | 3.2895 | 3.3366 | 3.3763 | 3.3763
same as TR 37.843
check 28/1/20
Directional chamber uses the conducted uncertainty value not the TE value, Reverb chamber uses TE value. Should they be the same or are systems sufficiently different that there is no need?
Reverb values are high eiter way.

## RX EM
| chamber | Expanded uncertainty [dB] | [17 empty] | Review comments (RAN4#94)
|  | 30MHz<f≤6 GHz | 6GHz<f ≤19GHz | 19GHz<f ≤26GHz
| General Directional Chamber | 2.5023 | 4.2549 | 4.2549
Reverberation Chamber
| Common maximum accepted test system uncertainty | 2.5 | 4.2 | 4.2
Table 12.3.2.3-1: IAC MU value derivation for RX spurious emissions
General Directional Chamber | [14 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | 30MHz<f≤6 GHz | 6GHz<f ≤19GHz | 19GHz<f ≤26GHz |  |  |  | 30MHz<f≤6 GHz | 6GHz<f ≤19GHz | 19GHz<f ≤26GHz
Stage 2: DUT measurement
A5-1 | Positioning misalignment between the AAS BS and the reference antenna | 0.03 | 0.03 | 0.03 | Rectangular | 1.7321 | 1 | 0.017321 | 0.017321 | 0.017321 |  |  |  |  | 0.0003 | 0.0003 | 0.0003
A5-2 | Pointing misalignment between the AAS BS and the receiving antenna | 0 | 0 | 0 | Rectangular | 1.7321 | 1 | 0 | 0 | 0 |  |  |  |  | 0 | 0 | 0
A5-3 | Quality of quiet zone | 0.1 | 0.1 | 0.1 | Gaussian | 1 | 1 | 0.1 | 0.1 | 0.1 |  |  |  |  | 0.01 | 0.01 | 0.01
A5-4a | Polarization mismatch between the AAS BS and the receiving antenna | 0.01 | 0.01 | 0.01 | Rectangular | 1.7321 | 1 | 0.0057735 | 0.0057735 | 0.0057735 |  |  |  |  | 3.3333e-05 | 3.3333e-05 | 3.3333e-05
A5-5a | Mutual coupling between the AAS BS and the receiving antenna | 0 | 0 | 0 | Rectangular | 1.7321 | 1 | 0 | 0 | 0 |  |  |  |  | 0 | 0 | 0
A5-6a | Phase curvature | 0.05 | 0.05 | 0.05 | Gaussian | 1 | 1 | 0.05 | 0.05 | 0.05 |  |  |  |  | 0.0025 | 0.0025 | 0.0025
C3-4 | Receiver spurious emissions - Conducted Uncertainty (minus missmatch) | 1 | 2 | 2 | Gaussian | 1 | 1 | 1 | 2 | 2 |  |  |  |  | 1 | 4 | 4
A5-7 | Impedance mismatch in the receiving chain | 0.2 | 0.45 | 0.45 | U-shaped | 1.4142 | 1 | 0.14142 | 0.3182 | 0.3182 |  |  |  |  | 0.02 | 0.10125 | 0.10125
A5-8 | Random uncertainty | 0.1 | 0.1 | 0.1 | Rectangular | 1.7321 | 1 | 0.057735 | 0.057735 | 0.057735 |  |  |  |  | 0.0033333 | 0.0033333 | 0.0033333
A5-17 | Measurement antenna frequency variation | 0.1 | 0.1 | 0.1 | Rectangular | 1.7321 | 1 | 0.057735 | 0.057735 | 0.057735 |  |  |  |  | 0.0033333 | 0.0033333 | 0.0033333
A5-18 | FSPL estimation error | 0 | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 | 0 |  |  |  |  | 0 | 0 | 0
A5-16 | Test system frequency flatness | 0.25 | 0.25 | 0.25 | Gaussian | 1 | 1 | 0.25 | 0.25 | 0.25 |  |  |  |  | 0.0625 | 0.0625 | 0.0625
A5-19 | Measurement system dynamic range uncertainty | 0.51 | 0.51 | 0.51 | Gaussian | 1 | 1 | 0.51 | 0.51 | 0.51 |  |  |  |  | 0.2601 | 0.2601 | 0.2601
Stage 1: Calibration measurement | [14 empty] | 0 | 0 | 0
A5-9 | Impedance mismatch between the receiving antenna and the network analyzer | 0.05 | 0.05 | 0.05 | U-shaped | 1.4142 | 1 | 0.035355 | 0.035355 | 0.035355 |  |  |  |  | 0.00125 | 0.00125 | 0.00125
A5-10 | Positioning and pointing misalignment between the reference antenna and the receiving antenna | 0.01 | 0.01 | 0.01 | Rectangular | 1.7321 | 1 | 0.0057735 | 0.0057735 | 0.0057735 |  |  |  |  | 3.3333e-05 | 3.3333e-05 | 3.3333e-05
A5-11 | Impedance mismatch between the reference antenna and the network analyzer. | 0.05 | 0.05 | 0.05 | U-shaped | 1.4142 | 1 | 0.035355 | 0.035355 | 0.035355 |  |  |  |  | 0.00125 | 0.00125 | 0.00125
A5-3 | Quality of quiet zone | 0.1 | 0.1 | 0.1 | Gaussian | 1 | 1 | 0.1 | 0.1 | 0.1 |  |  |  |  | 0.01 | 0.01 | 0.01
A5-4b | Polarization mismatch for reference antenna | 0.01 | 0.01 | 0.01 | Rectangular | 1.7321 | 1 | 0.0057735 | 0.0057735 | 0.0057735 |  |  |  |  | 3.3333e-05 | 3.3333e-05 | 3.3333e-05
A5-5b | Mutual coupling between the reference antenna and the receiving antenna | 0 | 0 | 0 | Rectangular | 1.7321 | 1 | 0 | 0 | 0 |  |  |  |  | 0 | 0 | 0
A5-6b | Phase curvature | 0.05 | 0.05 | 0.05 | Gaussian | 1 | 1 | 0.05 | 0.05 | 0.05 |  |  |  |  | 0.0025 | 0.0025 | 0.0025
C1-3 | Uncertainty of the network analyzer | 0.13 | 0.2 | 0.2 | Gaussian | 1 | 1 | 0.13 | 0.2 | 0.2 |  |  |  |  | 0.0169 | 0.04 | 0.04
A5-12 | Influence of the reference antenna feed cable | 0.05 | 0.05 | 0.05 | Rectangular | 1.7321 | 1 | 0.028868 | 0.028868 | 0.028868 |  |  |  |  | 0.00083333 | 0.00083333 | 0.00083333
A5-13 | Reference antenna feed cable loss measurement uncertainty | 0.06 | 0.06 | 0.06 | Gaussian | 1 | 1 | 0.06 | 0.06 | 0.06 |  |  |  |  | 0.0036 | 0.0036 | 0.0036
A5-14 | Influence of the receiving antenna feed cable | 0.05 | 0.05 | 0.05 | Rectangular | 1.7321 | 1 | 0.028868 | 0.028868 | 0.028868 |  |  |  |  | 0.00083333 | 0.00083333 | 0.00083333
C1-4 | Uncertainty of the absolute gain of the reference antenna | 0.50229 | 0.43301 | 0.43301 | Rectangular | 1.7321 | 1 | 0.29 | 0.25 | 0.25 |  |  |  |  | 0.0841 | 0.0625 | 0.0625
A5-15 | Uncertainty of the absolute gain of the receiving antenna | 0 | 0 | 0 | Rectangular | 1.7321 | 1 | 0 | 0 | 0 |  |  |  |  | 0 | 0 | 0
Combined standard uncertainty (1σ) [dB] | [7 empty] | 1.218 | 2.1369 | 2.1369 |  |  |  |  | 1.218 | 2.1369 | 2.1369
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [7 empty] | 2.3872 | 4.1883 | 4.1883 | same as 37.843 |  |  |  | 2.3872 | 4.1883 | 4.1883
TRP Summation error | [7 empty] | 0.75 | 0.75 | 0.75
Total MU | [7 empty] | 2.5023 | 4.2549 | 4.2549
Reverberation Chamber | [14 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  |  |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | 380 MHz < f ≦ 3 GHz | 3 GHz < f ≦ 6 GHz | 6 GHz < f ≦ 12.75 GHz | 12.75 GHz < f ≦ 19 GHz | 19GHz<f ≤26GHz |  |  |  | 380 MHz < f ≦ 3 GHz | 3 GHz < f ≦ 6 GHz | 6 GHz < f ≦ 12.75 GHz | 12.75 GHz < f ≦ 19 GHz | 19GHz<f ≤26GHz
Stage 2: DUT measurement
Stage 1: Calibration measurement | [14 empty] | 0 | 0 | 0 | 0 | 0
Combined standard uncertainty (1σ) [dB] | [9 empty] | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [9 empty] | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
check 28/1/20
Uses conducted uncertainty not TE

## COEX EM
Table 12.3.3-1: Test system specific MU values for the Additional (coexistence) spurious emissions measurement
| chamber | Expanded uncertainty [dB] | [11 empty] | Review comments (RAN4#94)
|  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz
| CATR | 2.6153 | 3.052 | 3.4661
| Common maximum accepted test system uncertainty | 2.6 | 3 | 3.5 | * note. error in 38.817-02 3 to 4.2GHz is captured as 2.6dB, ok in 38.141-2. 4.2 to 6 GHZ as 3dB in TR but as 3.5dB in TS
Table 12.4.2.3-1: CATR MU value derivation for additional (co-existanceexistence) OTA TX spurious emissions
CATR | [10 empty] | rms calculations
UID | Uncertainty source | Uncertainty value |  |  | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB]
|  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz |  |  |  | f≤3 GHz | 3<f≤4.2 GHz | 4.2<f≤6 GHz
Stage 2: DUT measurement
A2-18 | Misalignment DUT & pointing error for TRP | 0.3 | 0.3 | 0.3 | Rectangular | 1.7321 | 1 | 0.17321 | 0.17321 | 0.17321 | 0.03 | 0.03 | 0.03
C3-5 | Additional (COEX) emssions - Condcuted Uncertainty (minus missmatch) | 1.02 | 1.28 | 1.53 | Gaussian | 1 | 1 | 1.02 | 1.28 | 1.53 | 1.0404 | 1.6384 | 2.3409
A2-2a | Standing wave between DUT and test range antenna | 0.21 | 0.21 | 0.21 | U-shaped | 1.4142 | 1 | 0.14849 | 0.14849 | 0.14849 | 0.02205 | 0.02205 | 0.02205
A2-3 | RF leakage (SGH connector terminated & test range antenna connector cable terminated) | 0.0012 | 0.0012 | 0.0012 | Gaussian | 1 | 1 | 0.0012 | 0.0012 | 0.0012 | 1.44e-06 | 1.44e-06 | 1.44e-06
A2-4a | QZ ripple DUT | 0.0928 | 0.0928 | 0.0928 | Gaussian | 1 | 1 | 0.0928 | 0.0928 | 0.0928 | 0.0086118 | 0.0086118 | 0.0086118
A2-17 | Measurement system dynamic range uncertainty | 0.51 | 0.51 | 0.51 | Gaussian | 1 | 1 | 0.51 | 0.51 | 0.51 | 0.2601 | 0.2601 | 0.2601
A2-12 | Frequency flatness | 0.25 | 0.25 | 0.25 | Gaussian | 1 | 1 | 0.25 | 0.25 | 0.25 | 0.0625 | 0.0625 | 0.0625
Stage 1: Calibration measurement | [10 empty] | 0 | 0 | 0
C1-3 | Uncertainty of the network analyzer | 0.13 | 0.2 | 0.2 | Gaussian | 1 | 1 | 0.13 | 0.2 | 0.2 | 0.0169 | 0.04 | 0.04
A2-5 | Mismatch of receiver chain | 0.127 | 0.325 | 0.325 | U-shaped | 1.4142 | 1 | 0.089803 | 0.22981 | 0.22981 | 0.0080645 | 0.052812 | 0.052812
A2-6 | Insertion loss variation of receiver chain | 0.18 | 0.18 | 0.18 | Rectangular | 1.7321 | 1 | 0.10392 | 0.10392 | 0.10392 | 0.0108 | 0.0108 | 0.0108
A2-3 | RF leakage, (SGH connector terminated & test range antenna connector cable terminated) | 0.0012 | 0.0012 | 0.0012 | Gaussian | 1 | 1 | 0.0012 | 0.0012 | 0.0012 | 1.44e-06 | 1.44e-06 | 1.44e-06
A2-7 | Influence of the calibration antenna feed cable: | 0.022 | 0.022 | 0.022 | U-shaped | 1.4142 | 1 | 0.015556 | 0.015556 | 0.015556 | 0.000242 | 0.000242 | 0.000242
C1-4 | Uncertainty of the absolute gain of the reference antenna | 0.50229 | 0.43301 | 0.43301 | Rectangular | 1.7321 | 1 | 0.29 | 0.25 | 0.25 | 0.0841 | 0.0625 | 0.0625
A2-8 | Misalignment positioning system | 0 | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 | 0 | 0 | 0 | 0
A2-1b | Misalignment of calibration antenna and test range antenna | 0.5 | 0.5 | 0.5 | Exp. normal | 2 | 1 | 0.25 | 0.25 | 0.25 | 0.0625 | 0.0625 | 0.0625
A2-9 | Rotary Joints | 0.048 | 0.048 | 0.048 | U-shaped | 1.4142 | 1 | 0.033941 | 0.033941 | 0.033941 | 0.001152 | 0.001152 | 0.001152
A2-2b | Standing wave between calibration antenna and test range antenna | 0.09 | 0.09 | 0.09 | U-shaped | 1.4142 | 1 | 0.06364 | 0.06364 | 0.06364 | 0.00405 | 0.00405 | 0.00405
A2-4b | QZ ripple calibration antenna | 0.009 | 0.009 | 0.009 | Gaussian | 1 | 1 | 0.009 | 0.009 | 0.009 | 8.1e-05 | 8.1e-05 | 8.1e-05
A2-11 | Switching uncertainty | 0.26 | 0.26 | 0.26 | Rectangular | 1.7321 | 1 | 0.15011 | 0.15011 | 0.15011 | 0.022533 | 0.022533 | 0.022533
Combined standard uncertainty (1σ) [dB] | [7 empty] | 1.2783 | 1.5094 | 1.7265 | 1.2783 | 1.5094 | 1.7265
Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | [7 empty] | 2.5055 | 2.9585 | 3.384 | 2.5055 | 2.9585 | 3.384
TRP Summation error | [7 empty] | 0.75 | 0.75 | 0.75
Total MU | [7 empty] | 2.6153 | 3.052 | 3.4661 | same as 37.143
check 28/1/20
same as 37.843 - but teh pointing error is same as EIRP not TRP so inconsitant wioth latest update
Makes no difference to final value so use TRP value..