## Cover Sheet
3GPP RAN4 AH-1801 San Diego, US, 22 - 26 Jan, 2018
Tdoc number: | R4-1800761
Agenda item: | 4.7.2
Source: | CATR
Title: | MU Budget for EIS

## EIS
|  |  | Proposed MU values for EIS (D=5cm) |  |  |  | Notes
| Description of uncertainty contribution | Details in | Distribution of the probability | Divisor | Uncertainty value | Standard uncertainty (σ) [dB]
# | Stage 2-DUT measurement
1 | Positioning misalignment | B.1.1.4.1 | Rectangular | 1.732 | 0.5 | 0.28868
2 | Measure distance uncertainty | B.1.1.4.2 | Rectangular | 1.732 | 1 | 0.57737
3 | Quality of quiet zone* | B.1.1.4.3 | Actual | 1 | 1 | 1
4 | Mismatch | B.1.1.4.4 | U-shaped | 1.414 | 2.7432 | 1.94
5 | gNB emulator uncertainties | B.1.1.4.16 | Normal | 2 | 2 | 1 | different with EIRP
6 | Absolute antenna gain uncertainty of the measurement antenna | B.1.1.4.5 | Normal | 2 | 0 | 0
7 | Phase curvature | B.1.1.4.7 | U-shaped | 1.414 | 0 | 0
8 | Influence of the XPD | B.1.1.4.10 | U-shaped | 1.414 | 0.67872 | 0.48
9 | Amplifier uncertainties | B.1.1.4.8 | Normal | 2 | 2 | 1
10 | Random uncertainty | B.1.1.4.9 | Rectangular | 1.732 | 0.39836 | 0.23
Stage 1-Calibration measurement
11 | Mismatch | B.1.1.4.4 | U-shaped | 1.414 | 0 | 0
12 | Reference antenna positioning misalignment | B.1.1.4.11 | Rectangular | 1.732 | 0.29444 | 0.17
13 | Quality of quiet zone | B.1.1.4.3 | Actual | 1 | 1 | 1
14 | Amplifier uncertainties | B.1.1.4.8 | Normal | 2 | 0 | 0
15 | Uncertainty of the Network Analyzer | B.1.1.4.12 | Normal | 2 | 0.4 | 0.2
16 | Phase curvature | B.1.1.4.7 | U-shaped | 1.414 | 0 | 0
17 | Uncertainty of an absolute gain of the calibration antenna | B.1.1.4.14 | Normal | 2 | 1.6 | 0.8
18 | Positioning and pointing misalignment between the reference antenna and the receiving antenna | B.1.1.4.15 | Rectangular | 1.732 | 0.3464 | 0.2
Total Exp MU (1.96*sigma) [dB] | [5 empty] | 5.949