## Slide 1: 3GPP TSG-RAN WG4 Meeting #AH1801 R4-1800890 San Diego, US, 22nd – 26th Jan 2018
Source: Keysight Technologies, (MVG Industries)
Title: Extension of quiet zone characterization method to include phase characterization
Agenda item: 4.7.3.1
Document for: Approval

## Slide 2: Objective of this Contribution
This document addresses the method to characterize phase in the quiet zone. The outcome from the characterization would give the as phase ripple/curvature
R4-1711278 presents a TP (TR38810) describing the amplitude ripple characterization procedure for FR2.
  The outcome of the amplitude ripple test gives
    surface standard deviation that effectively predicts the variation of the TRP/TRS measurements
    spherical surface integrals of EIRP/EIS.
The objective of this contribution is to introduce a ‘Characterization of the phase ripple/ curvature in the Quiet Zone’ test procedure.
The reference antenna for the setup is FFS based on certain boundary conditions (R4-1711825)
Once agreement is reached, the TP (R4-1711278) will be appended.

## Slide 3: Contribution to MU due to phase and amplitude variation
Due to the use of phased array in the FR2 both amplitude and phase characteristics of the quiet zone is important to be characterized.
High amplitude or phase variation will cause non-coherent summation of the wave fronts from the individual elements.
The variation of MU due to these variations in the DUT performance depends on
  size of the phase array, topology of the array, element type used and the offset distance of the array from center of the quiet zone.

## Slide 4: QZ Phase ripple and taper test procedure
The phase characterization is a relative measurement made from the center of the test zone.
  Place the reference antenna at the center of the QZ and point it straight to the measurement antenna.
  Connect both the antennas to a VNA
  Calibrate this position with “through” measurement (S21) to set this as 0˚
    Calibration needs to be performed for both polarizations separately
  Move the reference antenna to all the points within the sphere defining the size of the quiet zone (D=15cm agreed)
    Measurement in both polarizations need to be performed.
  Step size required for calibration is λ/4 (2.6mm at 28GHz and 1.7mm at 43.5GHz)
    For reducing the number of steps, step size of n*3λ/4 can be considered, n to be FFS
    With larger step size the ripple cannot be captured but the phase taper can still be captured
Observation: The phase taper across the full volume of the quiet zone will impact the measurement results. Hence the through calibration is to be done only once at the center of the quiet zone
Proposal 1: Make phase characterization using vector measurement to highlight phase ripple and taper in the quiet zone
Proposal 2: Make phase measurements with a step size of n*λ/4 in the entire volume of the quiet zone, “n” is FFS
Proposal 3: Characterization is to be done for both polarizations individually
Proposal 4: Characterization frequency to be used is same as for amplitude characterization define in [R4-1711278]

## Slide 5: Proposals
Observation: The phase taper across the full volume of the quiet zone will impact the measurement results. Hence the through calibration is to be done only once at the center of the quiet zone
Proposal 1: Make phase characterization using vector measurement to highlight phase ripple and taper in the quiet zone
Proposal 2: Make phase measurements with a step size of n*λ/4 in the entire volume of the quiet zone, “n” is FFS
Proposal 3: Characterization is to be done for both polarizations individually
Proposal 4: Characterization frequency to be used is same as for amplitude characterization define in [R4-1711278]

## Slide 6: References
R4-1711278, “TP (TR38.810) on Quality of the Quiet Zone Characterization for FR2” Rohde & Schwarz, 3GPP TSG RAN WG4 #84bis, Dubrovnik, Croatia, October 2017
R4-1710077, WF on NR MU and test tolerance, CATR, Intel Corporation, Keysight Technologies, Rohde & Schwarz, Anritsu, 3GPP TSG RAN WG4 NR#3, Nagoya, Japan, September 2017
R4-1707221, Meaning of Quiet Zone in MU discussions, Anritsu, 3GPP TSG-RAN4 WG#84, Berlin, Germany, August 2017
R4-1711825, MVG Industries, 3GPP TSG RAN WG4 #84bis, Dubrovnik, Croatia, October 2017