## Slide 1: 3GPP TSG-RAN WG4 Meeting #AH1801 R4-1801263 San Diego, US, 22nd – 26th Jan 2018
Source: Keysight Technologies, MVG Industries
Title: Extension of quiet zone characterization method to include phase characterization
Agenda item: 4.7.3.1
Document for: Approval

## Slide 2: Objective of this Contribution
This document addresses the method to characterize phase in the quiet zone. The outcome from the characterization would give the as phase ripple/curvature
R4-1711278 presents a TP (TR38810) describing the amplitude ripple characterization procedure for FR2.
  The outcome of the amplitude ripple test gives
    surface standard deviation that effectively predicts the variation of the TRP/TRS measurements
    spherical surface integrals of EIRP/EIS.
The objective of this contribution is to introduce a ‘Characterization of the phase ripple/ curvature in the Quiet Zone’ test procedure.
The reference antenna for the setup is FFS based on certain boundary conditions (R4-1711825)
Once agreement is reached, the TP (R4-1711278) will be appended.

## Slide 3: Contribution to MU due to phase and amplitude variation
Due to the use of phased array in the FR2 both amplitude and phase characteristics of the quiet zone is important to be characterized.
High amplitude or phase variation will cause non-coherent summation of the wave fronts from the individual elements.
The variation of MU due to these variations in the DUT performance depends on
  size of the phase array, topology of the array, element type used and the offset distance of the array from center of the quiet zone.

## Slide 4: QZ Phase ripple and taper test procedure
The phase characterization is a relative measurement made from the center of the test zone.
  Place the reference antenna at the center of the QZ and point it towards the measurement antenna.
  Connect both the reference antenna and Measurement antennas to a VNA
  Calibrate this position with “through” measurement (S21) to set this as 0˚
    Calibration needs to be performed for both polarizations separately
  Move the reference antenna to all the points within the volume defining the size of the quiet zone (D=15cm agreed)
    Measurement in both polarizations need to be performed.
    To characterize the phase variation (taper and ripple) in the quiet zone, divide the measured phase by exp(j2pf/c * xn) to remove the plane wave propagation phase as the reference antenna is positioned within the quiet zone at positions Pn.
    Reduction of the test points/planes in the QZ volume is FFS
  Make phase measurements at specified test point TBD of the QZ using single pitch and
  The Reference antenna should be moved “n” cm more than the size of the expected QZ, and in the post processing the results of the dimension of QZ to be highlighted. “n” is FFS
Observation: The phase taper across the full volume of the quiet zone will impact the measurement results. Hence the through calibration is to be done only once at the center of the quiet zone. “kz” term to be compensated to remove the impact of natural phase variation, along the P2 to P3 axis. “k” is the freespace wave number and “z” is relative offset from through calibration point
Proposal 1: Make phase characterization using vector measurement to highlight phase ripple and taper in the quiet zone
Proposal 2: Make phase measurements at specified test point TBD of the QZ using single pitch and roll
Proposal 3: Characterization is to be done for both polarizations individually
Proposal 4: Characterization frequency to be used is same as for amplitude characterization define in [R4-1711278]

## Slide 5: Proposals
Observation: The phase taper across the full volume of the quiet zone will impact the measurement results. Hence the through calibration is to be done only once at the center of the quiet zone. “kz” term to be compensated to remove the impact of natural phase variation, along the P2 to P3 axis. “k” is the freespace wave number and “z” is relative offset from through calibration point
Proposal 1: Make phase characterization using vector measurement to highlight phase ripple and taper in the quiet zone
Proposal 2: Make phase measurements at specified test point TBD of the QZ using single pitch and roll
Proposal 3: Characterization is to be done for both polarizations individually
Proposal 4: Characterization frequency to be used is same as for amplitude characterization define in [R4-1711278]

## Slide 6: References
R4-1711278, “TP (TR38.810) on Quality of the Quiet Zone Characterization for FR2” Rohde & Schwarz, 3GPP TSG RAN WG4 #84bis, Dubrovnik, Croatia, October 2017
R4-1710077, WF on NR MU and test tolerance, CATR, Intel Corporation, Keysight Technologies, Rohde & Schwarz, Anritsu, 3GPP TSG RAN WG4 NR#3, Nagoya, Japan, September 2017
R4-1707221, Meaning of Quiet Zone in MU discussions, Anritsu, 3GPP TSG-RAN4 WG#84, Berlin, Germany, August 2017
R4-1711825, MVG Industries, 3GPP TSG RAN WG4 #84bis, Dubrovnik, Croatia, October 2017