3GPP TSG-RAN5 Meeting #82		R5-191628
Athens, Greece, 25 Feb - 1 Mar 2019

Agenda item:	5.3.8.17

Source:	Keysight Technologies

Title:	MU Elements: documentation of justification and interdependence

Document for:	Information

# Introduction

In the RAN5 AH#4 meeting, it was suggested to compile an overview of how the uncertainty values of the various MU elements have been justified and documented and how MU elements are interdependent. This contribution is meant to provide this overview.

# MU Table for MOP (EIRP, TRP, Spherical Coverage) with Overview

The table below corresponds to the MU Table for the maximum output power test cases based on EIRP, Spherical Coverage, and TRP for the IFF test methodology and for the 30cm DUT size/quiet zone size, i.e., Table B.3.2-3 in [1][2] . The first two columns provide additional overview whether the respective MU elements are dependent on the 8x2 UE antenna array assumption [3] or dependent on test equipment or component data sheets. The third component is a concise list of RAN4 or RAN5 contributions or references to write-up in [1] providing justification for the MU values.

Table 1: MU Table for MOP Test Case based on EIRP, Spherical Coverage, and TRP for IFF (DUT size/QZ size = 30cm)

| Related to 8x2 Antenna Assumption | Related to Equipment/ Component Datasheets | Contributions with documentation | UID | Uncertainty source | Uncertainty value [dB] | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  | Stage 2: DUT measurement |  |  |  |  |  |
| yes |  | R5-190389, R5-190566 | 1 | Positioning misalignment (NOTE 7) | 0 | Normal | 2 | 0 |
|  |  | N/A (inherent feature of IFF) | 2 | Measure distance uncertainty | 0 | Rectangular | 1.73 | 0 |
|  |  | R5-185956 (30cm) R5-185955 (15cm) | 3 | Quality of Quiet Zone (NOTE 1) | 0.5 | Actual | 1 | 0.5 |
|  | yes | R5-190345, R4-1714378 | 4 | Mismatch (NOTE 2, NOTE 7) | 1.3 | Actual | 1 | 1.3 |
|  |  | R4-1815960 | 5 | Standing wave between the DUT and measurement antenna | 0 | U-shaped | 1.41 | 0 |
|  | yes | R4-1713677, R5-190804 | 6 | Uncertainty of the RF power measurement equipment (NOTE 3) (NOTE 7) | 2.16 | Normal | 2 | 1.08 |
|  |  | inherent feature of IFF; see B.2.2.7 of TR38.903 | 7 | Phase curvature | 0 | U-shaped | 1.41 | 0 |
|  | yes | TBD | 8 | Amplifier uncertainties (NOTE 7) | [2.00] | Normal | 2 | 1 |
|  |  | B.2.1.9 of TR38.903 | 9 | Random uncertainty | 0.5 | Normal | 2 | 0.25 |
|  | yes | R5-190864 | 10 | Influence of the XPD | 0.01 | U-shaped | 1.41 | 0.00 |
|  |  | B.2.1.11 of TR38.903 (same cable assumption) | 11 | Insertion Loss Variation | 0 | Rectangular | 1.73 | 0 |
|  |  | B.2.1.12 of TR38.903, R5-187226 | 12 | RF leakage (from measurement antenna to the receiver/transmitter) | 0 | Actual | 1 | 0 |
| yes |  | R4-1808410 (LS to RAN5), R4-1813865 | 13 | Influence of TRP measurement grid (NOTE 4) | 0.25 | Actual | 1 | 0.25 |
| yes |  | R4-1816261 | 14 | Influence of beam peak search grid (NOTE 5) | 0 | Actual | 1 | 0 |
|  |  | R5-190387 | 15 | DUT repositioning (TRP, spherical coverage) | 0 | Rectangular | 1.73 | 0 |
| yes |  | R5-190387 | 15 | DUT repositioning (EIRP) | 0.08 | Rectangular | 1.73 | 0.05 |
|  | yes | R5-190863 | 16 | Multiple measurement antenna uncertainty (NOTE 8) | 0.15 | [Actual] | 1 | 0.15 |
| yes |  | R4-1814837 | 17 | Influence of spherical coverage grid (NOTE 10) | 0.12 | Actual | 1 | 0.12 |
|  |  |  | Stage 1: Calibration measurement |  |  |  |  |  |
|  | yes | R4-1714378 (combined Stage 1 and Stage 2) | 17 | Mismatch | 0 | U-shaped | 1.41 | 0 |
|  |  | combined Stage 1 and Stage | 18 | Amplifier Uncertainties (NOTE 7) | [0.00] | Normal | 2 | 0 |
|  | yes | R5-190389 | 19 | Misalignment of positioning System | 0 | Normal | 2 | 0 |
|  |  | R4-1713677, R5-191616 | 20 | Uncertainty of the Network Analyzer (NOTE 7) | [0.40] | Normal | 2 | 0.2 |
|  | yes | R5-185954 | 21 | Uncertainty of the absolute gain of the calibration antenna | 0.6 | Normal | 2 | 0.3 |
|  |  | R5-190805 | 22 | Positioning and pointing misalignment between the reference antenna and the measurement antenna (NOTE 7) | 0.01 | Normal | 2 | 0 |
|  |  | R4-1709839 | 23 | Phase centre offset of calibration antenna | 0 | Rectangular | 1.73 | 0 |
|  |  | R5-185956 (30cm) R5-185955 (15cm) | 24 | Quality of quiet zone for calibration process (NOTE 1) | 0.2 | Actual | 1 | 0.2 |
|  |  | R4-1815960 | 25 | Standing wave between reference calibration antenna and measurement antenna (NOTE 7) | 0 | U-shaped | 1.41 | 0 |
|  | yes | R5-190394 | 26 | Influence of the calibration antenna feed cable (NOTE 7) | 0.14 | Normal | 2 | 0.07 |
|  |  | B.2.1.11 of TR38.903 (same cable assumption) | 27 | Insertion Loss Variation | 0 | Rectangular | 1.73 | 0 |
|  |  |  |  | Systematic uncertainties (NOTE 6) |  |  |  | Value |
| yes |  | R4-1813865 | 28 | Mean error for constant step size grid (NOTE 4) |  |  |  | 0 |
|  |  | R5-190059, R5-190601 |  | Influence of noise (Minimum peak EIRP @23.45 GHz) (NOTE 8) |  |  |  | 0.02 |
|  |  |  |  | Influence of noise (Minimum peak EIRP @32.125 GHz) (NOTE 8) |  |  |  | 0.05 |
|  |  |  |  | Influence of noise (Minimum peak EIRP @40.8 GHz) (NOTE 8) |  |  |  | 0.42 |
| yes |  | R4-1816261 |  | Systematic error related to beam peak search (NOTE 5) |  |  |  | 0.5 |
|  |  |  | Total measurement uncertainty |  |  |  |  |  |
|  |  |  | EIRP Expanded uncertainty (@23.45 GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 4.62 |
|  |  |  | EIRP Expanded uncertainty (@32.125 GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 4.65 |
|  |  |  | EIRP Expanded uncertainty (@40.8 GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 5.02 |
|  |  |  | TRP Expanded uncertainty (@23.45 GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 4.14 |
|  |  |  | TRP Expanded uncertainty (@32.125 GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 4.17 |
|  |  |  | TRP Expanded uncertainty (@40.8 GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 4.54 |
|  |  |  | Spherical coverage (@23.45 GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 4.12 |
|  |  |  | Spherical coverage (@32.125 GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 4.15 |
|  |  |  | Spherical coverage (@40.8 GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 4.52 |
|  |  |  | NOTE 1: The quality of quiet zone is the same for EIRP and TRP. |  |  |  |  |  |
|  |  |  | NOTE 2: The analysis was done only for the case of operating at max output power, in-band, non-CA. |  |  |  |  |  |
|  |  |  | NOTE 3: The assessment assues maximum DUT output power. |  |  |  |  |  |
|  |  |  | NOTE 4: This contributor shall only be considered for TRP measurements. |  |  |  |  |  |
|  |  |  | NOTE 5: This contributor shall only be considered for EIRP measurements. |  |  |  |  |  |
|  |  |  | NOTE 6: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. |  |  |  |  |  |
|  |  |  | NOTE 7: Values extracted from TR 38.810 v2.6.1 in square brackets pending for further analysis. |  |  |  |  |  |
|  |  |  | NOTE 8: These new terms are yet to be approved to use in the Uncertainity Assessment and needs further technical analysis. |  |  |  |  |  |
|  |  |  | NOTE 9: Value based on procedure defined in Annex D.2 of TR 38.810 for Quiet Zone size of 30 cm. |  |  |  |  |  |
|  |  |  | NOTE 10: This contributor shall only be considered for spherical EIRP measurements. |  |  |  |  |  |

# References

- TR 38.903, Derivation of test tolerances and measurement uncertainty for User Equipment (UE) conformance test cases, V15.1.0 (2018-12)

- R5-191248, CR to TR38.903 to update MU tables and element descriptions, Keysight Technologies, 3GPP TSG-RAN5 Meeting #82, February 2019

- TR 38.810, Study on test methods for New Radio, V16.1.0 (2018-12)