3GPP TSG-RAN5 Meeting #86-e		R5-200327
Electronic Meeting, 17th – 28th February 2020

Agenda item:	5.3.2.17

Source:	Keysight Technologies

Title:	On the MU Element “Uncertainty of an absolute gain of the calibration antenna” for Spurious Emissions

Document for:	Discussion and Endorsement

# Background

The MU element “Uncertainty of an absolute gain of the calibration antenna” defined in [1]

B.2.1.15 Uncertainty of an absolute gain of the calibration antenna The calibration antenna only appears in Stage 2. Therefore, the gain uncertainty has to be taken into account. This uncertainty will come from a calibration report with traceability to a National Metrology Institute with measurement uncertainty budgets generated following the guidelines outlined in internationally accepted standards.

Has been defined with an uncertainty value of 0.6dB for in-band/OOB test cases in [1] based on [2]

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 17 | Uncertainty of an absolute gain of the calibration antenna | 0.60 | Normal | 2.00 | 0.30 |

No MU value has been defined for this MU element yet for spurious emissions test cases.

# MU Element Proposal “Uncertainty of an absolute gain of the calibration antenna” for Spurious Emissions

Since the MU value with uncertainty of 0.6dB has been agreed for in-band/OOB test cases within the 23.45-40.8GHz frequency range, it is proposed to re-use this MU value for the 23.45 and 40.8GHz frequencies of the spurious emissions MU table.

Proposal 1: For the spurious emissions MU tables that cover the (23.45GHz to 40.8GHz) frequency range, re-use the 0.6dB value for the “Uncertainty of an absolute gain of the calibration antenna” MU element from the in-band/OOB test cases

The MU element for frequencies of 6 and 12.75GHz can be defined using the same calibration certificate presented in [2]. Excerpts from the calibration certificate are shown below

Since the uncertainties below 23.45GHz are matching the uncertainties for 23.45-40.8GHz, it is proposed to re-use the 0.6dB value for the “Uncertainty of an absolute gain of the calibration antenna” MU element for the spurious emissions MU tables that cover the 6GHz and 12.75GHz frequencies.

Proposal 2: For the spurious emissions MU tables that cover the (6GHz to 12.75GHz) frequency range, re-use the 0.6dB value for the “Uncertainty of an absolute gain of the calibration antenna” MU element from the in-band/OOB test cases

# Conclusion

The following observations and proposals were made in this contribution

Proposal 1: For the spurious emissions MU tables that cover the (23.45GHz to 40.8GHz) frequency range, re-use the 0.6dB value for the “Uncertainty of an absolute gain of the calibration antenna” MU element from the in-band/OOB test cases

Proposal 2: For the spurious emissions MU tables that cover the (6GHz to 12.75GHz) frequency range, re-use the 0.6dB value for the “Uncertainty of an absolute gain of the calibration antenna” MU element from the in-band/OOB test cases

# References

- TR 38.903, Derivation of test tolerances and measurement uncertainty for User Equipment (UE) conformance test cases, V16.2.0 (2019-12)

- R5-185954, On the MU Element “Uncertainty of an absolute gain of the calibration antenna”, Keysight Technologies, 3GPPRAN5#3-5G-NR Adhoc, October 2018