3GPP TSG-RAN5 Meeting #86-e                                                    R5-201091

Electronic Meeting, 17th – 28th February 2020

Title:		On FR2 Transmission OFF power MU

Source:	Keysight Technologies

Agenda Item:   5.3.2.17

Document for: Endorsement

- Introduction

FR2 Measurement Uncertainties for IFF are still TBD in TR38.903 ([1]).

This document makes a proposal on how to define FR2 measurement uncertainties for FR2 TX OFF power measurement for IFF testing methodology.

- Pending MU factors to be defined

Basically, there are 3 MU factors pending:

- Uncertainty of the RF power measurement equipment: As power values to be measured in TX OFF power test case are lower than for maximum output power test case, RF power measurement uncertainty is higher. With our internal study specific to conformance test system, our view is 2.5dB is the estimated uncertainty for FR2a.

- Mismatch: Same values as in Maximum output power test case can be used for either FR2a and FR2b as similar path is expected to be used

- Noise impact: As described in [2], square brackets can be removed for noise impact FR2a range while FR2b is proposed not to be tested, hence TBD values for FR2b can be set as N/A.

- Additional alignments with Maximum output power uncertainty table

Some alignments with maximum output power test MU table are required (notes, notes applicability).

- Conclusion

This document makes a proposal on how to define remaining FR2 measurement uncertainties for FR2 TX OFF power measurement (IFF testing methodology) for FR2a frequency range. The following proposal is done:

-      Proposal 1: Apply 1.3 dB to Mismatch for FR2a assuming 30.4dB relaxation for 400MHz BW

-     Proposal 2: Apply 2.5 dB to Uncertainty of the RF power measurement equipment for FR2a assuming 30.4dB relaxation for 400MHz BW,

-       Proposal 3: Apply 1.0 dB to influence of noise for FR2a assuming 30.4dB relaxation for 400MHz BW.

-      Proposal 4: Define as N/A to influence of noise for FR2b.

- 	Proposal 5: Correct NOTE 1 as “The analysis was done only for the case of operating at TX OFF power, in-band, non-CA.”

- Proposal 6: Void NOTE 4 in table for FR2 Transmission OFF power MU and remove any references to it.

- 	Proposal 7: Void NOTE 6 in table for FR2 Transmission OFF power MU and remove any references to it.

- 	Proposal 8: Define NOTE 9 as “Applies to the system which has a structure of mechanical feed antenna positioning.”

Implementation of Proposals 1 to 8 would look like as shown in the table below:

Table B.8.2-2: Uncertainty assessment for TRP measurement (f=23.45GHz, 32.125GHz, 40.8GHz, Quiet Zone size ≤ 30 cm)

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment | 0.00 | Normal | 2.00 | 0.00 |
| 2 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 3 | Quality of Quiet Zone (NOTE 8) | 0.6 | Actual | 1.00 | 0.6 |
| 4 | Mismatch (NOTE 1) | 1.30 | Actual | 1.00 | 1.30 |
| 5 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | Uncertainty of the RF power measurement equipment (NOTE 2) | 2.5 | Normal | 2.00 | 1.25 |
| 7 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 8 | Amplifier uncertainties | 2.10 | Normal | 2.00 | 1.05 |
| 9 | Random uncertainty | 0.50 | Normal | 2.00 | 0.25 |
| 10 | Influence of the XPD | 0.01 | U-shaped | 1.41 | 0.00 |
| 11 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 13 | Influence of TRP measurement grid (NOTE 3) | 0.25 | Actual | 1 | 0.25 |
| 14 | Influence of beam peak search grid | 0.00 | Actual | 1 | 0.00 |
| 15 | Multiple measurement antenna uncertainty (NOTE 9) | 0.15 | Actual | 1 | 0.15 |
| 16 | DUT repositioning | 0.00 | Rectangular | 1.73 | 0.00 |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 17 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 18 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 19 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 20 | Uncertainty of the Network Analyzer | 0.73 | Normal | 2.00 | 0.37 |
| 21 | Uncertainty of the absolute gain of the calibration antenna | 0.60 | Normal | 2.00 | 0.30 |
| 22 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.01 | Rectangular | 1.73 | 0.00 |
| 23 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 24 | Quality of quiet zone for calibration process (NOTE 8) | 0.4 | Actual | 1.00 | 0.4 |
| 25 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 26 | Influence of the calibration antenna feed cable | 0.14 | Normal | 2.00 | 0.07 |
| 27 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
|  | Systematic uncertainties (NOTE 5) |  |  |  | Value |
| 28 | Systematic error due to TRP calculation/quadrature (NOTE 3) |  |  |  | 0.0 |
| 29 | Influence of noise (23.45GHz <= f <= 32.125GHz) |  |  |  | 1.0 |
| 30 | Influence of noise (32.125GHz < f <= 40.8GHz) |  |  |  | N/A |
| Total measurement uncertainty |  |  |  |  | Value |
| TRP Expanded uncertainty (23.45GHz <= f <= 32.125GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | TBD |
| TRP Expanded uncertainty (32.125GHz < f <= 40.8GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | N/A |
| NOTE 1: The analysis was done only for the case of operating at TX OFF power, in-band, non-CA. NOTE 2: The assessment assumes DUT Off power. NOTE 3: This contributor shall only be considered for TRP measurements. NOTE 4: Void NOTE 5: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 6: Void. NOTE 7: Void NOTE 8: Value based on procedure defined in Annex D.2 of TR 38.810 for Quiet Zone size less or equal to 30 cm. NOTE 9: Applies to the system which has a structure of mechanical feed antenna positioning. |  |  |  |  |  |

NOTE: MU assessment in Table B.8.2-2 for FR2a is based on the relaxation of 30.4dB for 400MHz BW.

- References

# [1] TR 38.903 v16.2.0, “Derivation of test tolerances and measurement uncertainty for User Equipment (UE) conformance test cases

# [2] R5-199497, “Proposal of noise impact for FR2 Transmitter Test Case”, Anritsu, Keysight, ROHDE & SCHWARZ, TSG WG RAN5 Meeting #85, Reno