3GPP TSG RAN Meeting #87-e	 	R5-202454

Electronic Meeting, 18th  – 29th May 2020

Source:		Rohde & Schwarz

Title:	On the MU for ALCR

Agenda Item:		5.3.2.17

Document for:	Discussion and Endorsement

# Introduction

The MU for ACLR has been discussed in RAN5 during the last meetings [1-11], however, not all  MU contributors’ values could be concluded.

In this discussion paper, we share our view on some of the remaining MU contributors for ACLR.

# Discussion

ACLR is based on the ratio of the power in the adjacent channel to that of the wanted channel. In the case of single carrier, the bandwidth to be covered by the measurement is 800 MHz for a channel bandwidth of 400 MHz. Due to the combination of two power measurements in the ACLR figure, the MU could be 2× the MU of a single power measurement.

## Uncertainty of the RF power measurement

The main MU contributions forming the uncertainty of the RF power measurement MU contributor are linearity and frequency response since the noise is treated as separate contributor in the MU tables.

While the frequency response can be compensated to some extent, limitations in the frequency correction exist and the residual frequency response cannot be neglected.

Observation 1: Limitations in the frequency correction exist and the residual frequency response cannot be neglected.

Depending on the characteristics of the frequency response, it may be reduced when averaging over a considerable bandwidth as it is the case for a full allocation. However, since ACLR is tested also with 1RB_Edge allocation, the frequency response is not reduced by averaging the channel power over a considerable bandwidth.

Observation 2: Since ACLR is tested also with 1RB_Edge allocation, the frequency response of the test system is not reduced by averaging the channel power over a considerable bandwidth.

As the frequency response can be compensated to some extent, 2× the MU of a single power measurement may overestimate the ACLR MU. As a middle ground, we propose to keep the value for uncertainty of the RF power measurement equipment in TR 38.903, which is commonly applied, e.g., for MOP, SEM and spurious emission test cases

Proposal 1: Keep the value of 2.16 dB for the Uncertainty of the RF power measurement equipment in TR 38.903 and remove the square brackets.

## Gain of the calibration antenna

When an antenna has a constant gain versus frequency dependence, the MU can safely be assumed to be fully correlated. In that case the uncertainty contribution in the calibration stage for a relative measurement would be zero. Real antennas, however, are characterized by a frequency dependent gain and the correlation of the MU of the gain for two frequencies can have a complicated dependence on the frequency separation. Therefore, certified labs performing antenna calibrations typically state only an uncertainty for the absolute gain in the calibration report. Regarding the total MU the middle ground between full correlation (MUrel = 0 dB) and statistical independence (MUrel=2×MUabs), i.e. MUrel = MUabs is a reasonable assumption since the type of correlation is unknown. Therefore, we propose to keep the value of 0.6 dB for the MU contributor Uncertainty of the absolute gain of the calibration antenna.

Proposal 2: Keep the value of 0.6 dB for the MU contributor Uncertainty of the absolute gain of the calibration antenna and remove the square brackets.

# Conclusion

In this contribution, the following observations and proposals were made:

Observation 1: Limitations in the frequency correction exist and the residual frequency response cannot be neglected.

Observation 2: Since ACLR is tested also with 1RB_Edge allocation, the frequency response of the test system is not reduced by averaging the channel power over a considerable bandwidth.

Proposal 1: Keep the value of 2.16 dB for the Uncertainty of the RF power measurement equipment in TR 38.903 and remove the square brackets.

Proposal 2: Keep the value of 0.6 dB for the MU contributor Uncertainty of the absolute gain of the calibration antenna and remove the square brackets.

# References

- R5-188189, On the FR2 MU for occupied BW and ACLR, Anritsu, 3GPP TSG RAN WG5 Meeting #81, November 2018

- R5-190868, Discussion on ACLR MU, Rohde & Schwarz, 3GPP TSG RAN WG5 #4-5G-NR Adhoc, January 2019

- R5-191625, MU for QoQZ for ACLR, Keysight Technologies UK Ltd, 3GPP TSG RAN WG5 Meeting #82, February 2019

- R5-192659, On the QoQZ impact for ACLR, Anritsu, 3GPP TSG RAN WG5 Meeting #82, February 2019

- R5-192661, On ACLR measurement uncertainty, Keysight Technologies UK Ltd, 3GPP TSG RAN WG5 Meeting #82, February 2019

- R5-194624, On mismatch uncertainty for ACLR, Rohde & Schwarz, 3GPP TSG RAN WG5 Meeting #83, May 2019

- R5-195133, On the minimum TRP assumption for FR2 ACLR MU, Anritsu, 3GPP TSG RAN WG5 Meeting #83, May 2019

- R5-196572, Impact of noise for FR2 ACLR, Anritsu, 3GPP TSG RAN WG5 Meeting #84, August 2019

- R5-197736, On ACLR test metric, NTT DOCOMO, 3GPP TSG RAN WG5 Meeting #85, November 2019

- R5-198451, On the FR2 ACLR test metric, Anritsu, 3GPP TSG RAN WG5 Meeting #85, November 2019

- R5-198562, On FR2 ACLR noise impact, Keysight Technologies UK Ltd, 3GPP TSG RAN WG5 Meeting #85, November 2019

- R5-200738, On FR2 ACLR MUs, Anritsu, 3GPP TSG RAN WG5 Meeting RAN5#86-e, February 2020