3GPP TSG RAN Meeting #87-e	 	R5-202458

Electronic Meeting, 18th  – 29th May 2020

Source:		Rohde & Schwarz

Title:	On the MU of FR2 OBW

Agenda Item:		5.3.2.17

Document for:	Discussion and Endorsement

# Introduction

The measurement uncertainty (MU) of occupied bandwidth (OBW) has been discussed during the last meetings [1] to [12]. However, no final conclusion has been reached due to the dependence of OBW on the UE ACLR. At RAN5#85, measured data has been provided to be used as basis for the OBW MU analysis [9]. This contribution is an updated version of the paper submitted to RAN5#86-e [10] with respect to interpolation of the measurement data and the calculation of OBW.

# Discussion

The spectra of each polarization and of the total power are plotted for 100 MHz CHBW and 4x100 MHz in Figures 1 and 2, respectively. During RAN5#85, it has been confirmed that the 4x100 MHz CA dataset can be applied to derive the MU for OBW for 400 MHz single carrier.

Figure 1: Spectrum for 100 MHz CHBW from [9]

Figure 2: Spectrum for 400 MHz CHBW from [9]

Table 1 summarizes some parameters of the dataset. Interpolation of the provided data has been applied in order to improve the resolution of the OBW results. The OBW has been calculated applying a evaluation window whose center is aligned on the center of the channel.

Table 1: Results of the two data sets

|  | 100 MHz CHBW dataset | 4x100 MHz CA dataset |
| --- | --- | --- |
| center frequency (GHz) | 27.9983 | 27.9955 |
| OBW (span=2.0⋅CHBW) | 94.240 MHz | 391.84 MHz |
| OBW (span=1.5⋅CHBW) | 94.224 MHz | 391.36 MHz |
| CP (dBm) | 20.38 | 18.306 |

For the channel power CP, it is assumed that the RAW data has been measured with a resolution bandwidth of 1 MHz. If the RBW has been different, there will be a constant shift in the absolute power level which should not be relevant for the OBW MU derivation.

In the following, we refer the SNR to the channel power. The SNR value is a total SNR accounting for both polarizations. In the diagrams the influence of noise on the OBW is visible. For both CHBWs a span of twice the CHBW has been assumed:

For a span of 1.5 times the channel bandwidth, the following results are obtained:

Observation 1: By applying a span to CHBW ratio of 1.5 instead of 2.0 the impact of the SNR on the OBW can be considerably reduced.

Proposal 1: Apply a span of 1.5 CHBW for 400 MHz and use the same span for the other channel bandwidth in order to optimize the impact of noise.

Occupied bandwidth is tested with CP-OFDM waveform which has a larger MPR than DTF-s-OFDM waveform. In order to reduce the impact of noise, DFT-s-OFDM should be selected.

Observation 2: Occupied bandwidth is tested with CP-OFDM waveform which has a larger MPR than DTF-s-OFDM waveform.

Proposal 2: Select DFT-s-OFDM QPSK as waveform and modulation since it has the lowest MPR and allows to reduce the impact of noise on the OBW MU.

We note that the impact of test system flatness of the OBW is FFS.

# Assumptions

| ID | Description | Assumption |
| --- | --- | --- |
| #1 | Frequency ranges under consideration | All Rel-15 FR2 bands for in-band measurements 6 GHz – 80 GHz for spurious emissions measurements |
| #2 | Size of QZ for IFF | 30 cm |
| #3 | UE power class | PC3 |
| #5 | Temperature range of the test equipment | 20°C – 35°C |
| #6 | Channel bandwidth | 400 MHz |

# Conclusion

In summary, the SNR impact on the MU for ACLR has been analyzed.

Observation 1: By applying a span to CHBW ratio of 1.5 instead of 2.0 the impact of the SNR on the OBW can be considerably reduced.

Proposal 1: Apply a span of 1.5 CHBW for 400 MHz and use the same span for the other channel bandwidth in order to optimize the impact of noise.

Observation 2: Occupied bandwidth is tested with CP-OFDM waveform which has a larger MPR than DTF-s-OFDM waveform.

Proposal 2: Select DFT-s-OFDM QPSK as waveform and modulation since it has the lowest MPR and allows to reduce the impact of noise on the OBW MU.

# References

- R5-188189, On the FR2 MU for occupied BW and ACLR, Anritsu, 3GPP TSG RAN WG 5 Meeting #81, November 2018

- R5-190390, On QoQZ MU Impact on OBW, Keysight, 3GPP TSG RAN WG 5 Meeting #4-5G-NR Adhoc, January 2019

- R5-192660, On the FR2 OBW MU, Anritsu, 3GPP TSG RAN WG 5 Meeting #82, February 2019

- R5-191617, On OBW measurement uncertainty, Keysight, 3GPP TSG RAN WG 5 Meeting #82, February 2019

- R5-194353, On the noise impact to FR2 OBW, Anritsu, 3GPP TSG RAN WG 5 Meeting #83, May 2019

- R5-192984, On the SNR for FR2 TRx test cases, Rohde & Schwarz, 3GPP TSG RAN WG 5 Meeting #5-5G-NR Adhoc, April 2019

- R5-193187, On noise impact, Keysight, 3GPP TSG RAN WG 5 Meeting #5-5G-NR Adhoc, April 2019

- R5-197626, Impact of noise for FR2 OBW, Anritsu, 3GPP TSG RAN WG 5 Meeting #83, May 2019

- R5-198175, Discussion on AP#84.22 Tx Spectrum Shape for FR2 OBW MU, Qualcomm, 3GPP TSG RAN WG 5 Meeting #85, November 2019

- R5-200869, On the MU of FR2 OBW, Rohde & Schwarz, 3GPP TSG RAN WG 5 Meeting #86-e, February 2019

- R5-200739, On FR2 OBW MU, Anritsu, 3GPP TSG RAN WG 5 Meeting #86-e, February 2019

- R5-200874, On FR2 OBW MU, Keysight Technologies, 3GPP TSG RAN WG 5 Meeting #86-e, February 2019