3GPP TSG-RAN5 Meeting #87-e		R5-202846
Electronic Meeting, 18 May – 29 May 2020

Agenda item:	5.3.2.17

Source:	Keysight Technologies

Title:	On NR FR2 Demodulation MUs

Document for:	Discussion and Endorsement

# Introduction

This contribution is investigating select MU elements for NR FR2 demodulation test cases and whether the elements apply to demodulation test cases just like they did for UE RF REFSENS test case.

# Discussion

In the last meeting, the baseline MU Table for Mode 1 (conditions with external noise) and for Mode 2 (noise free conditions) demodulation test cases were agreed in [1]. Many MU elements in those tables correspond to MU elements from the UE RF REFSENS MU table in Clause B.19 of [2]. However, the applicability of the UE RF MU elements to demodulation Mode 1 and Mode 2 test case has not been established and is the basis of this contribution.

In [3], we provided documentation and supporting information for various MU elements. Similar information has been added to the MU tables in [1] below, i.e., in Table 1 for Mode 1 and in Table 2 for Mode 2.

It is important to point out that the REFSENS test case requires the RX beam steered in the RX beam peak direction and select MU elements quantify the uncertainty of not measuring exactly at the “true” beam peak. On the other hand, demodulation test cases do not have to be performed at the RX beam peak as the EIS in the test direction “just” has to meet the REFSENS requirement. As such, some MU elements should not be applicable the same way they were for the REFSENS test case.

The MU elements related to positioning and/or misalignments:

- Positioning misalignment

- DUT repositioning

- Misalignment of positioning system

- Positioning and pointing misalignment between the reference antenna and the measurement antenna

were derived based on statistical analyses how small deviations in positioning misalignment affect EIS measurements at the desired beam peak position when compared to the actual position. If the verification that the UE meets the REFSENS requirement in the desired test direction is performed immediately before the execution of the demodulation test case or if the REFSENS verification was performed without any device re-positioning between the initial verification and the execution of the demodulation test case, these MU elements should not be applicable (uncertainty value of 0dB) as absolute power variations due to positioning/misalignments from the RX beam peak are irrelevant. Whether these MU elements apply to Mode 1 test cases needs to be further studied.

Observation 1: MU elements related to positioning and/or misalignments (Positioning misalignment, DUT repositioning, Misalignment of positioning system, Positioning and pointing misalignment between the reference antenna and the measurement antenna) for Mode 2 test cases can be considered negligible if the REFSENS condition was verified without any re-positioning between the verification and the start of the demodulation test case.

However, if the UE was re-positioned between the initial verification and the execution of the test case, these MU element descriptions and uncertainty values are applicable just as outlined for the REFSENS test case (worst case assumption).

Observation 2: MU element descriptions and uncertainties related to positioning and/or misalignments (Positioning misalignment, DUT repositioning, Misalignment of positioning system, Positioning and pointing misalignment between the reference antenna and the measurement antenna) for Mode 2 test cases should be re-used from the REFSENS test case if the UE has been re-positioned between the REFSENS verification and the start of the demodulation test case.

As outlined above, requiring a REFSENS verification before each demodulation test case would allow a small reduction in MU at the expense of test time. In order to optimize overall test time for demodulation test cases, it is proposed to re-use MU element descriptions and uncertainties the from the REFSENS test case

Proposal 1: Regarding MU elements related to positioning and/or misalignments (Positioning misalignment, DUT repositioning, Misalignment of positioning system, Positioning and pointing misalignment between the reference antenna and the measurement antenna) for Mode 2 test cases: re-use from the REFSENS test case.

The ‘Systematic error related to beam peak search’ of 0.5dB was based on a statistical analysis of the offset of beam peak search measurements from the beam peak that contains 95% of the distribution (alternatively, the value at which the CDF is 5%) and was a function of the beam peak search measurement grid [4]. This MU element should only be applicable to Mode 2 demodulation test cases when the RX beam peak search, documented in K.1.2 of [5], was used and little to no margin to the REFSENS requirement was observed. When sufficient margin, e.g., greater than the [0.5]dB systematic error, to the REFSENS requirement is observed at the test direction, this MU element should not be applicable and considered 0dB.

Given the lack of empirical data, either more time to further study this MU element is requested or the worst case assumption, i.e., re-use the ‘Systematic error related to beam peak search’ of 0.5dB from the REFSENS test case, should be assumed. Even if the worst case assumption of 0.5dB systematic error is assumed for now, this uncertainty can be further minimized in a Study Item tailored to improve overall FR2 MU [8]. Feedback from the OEMs is requested.

Proposal 2: Feedback from the OEMs is requested: allow more time to further study the ‘Systematic error related to beam peak search’ for Mode 2 demodulation test cases or assume the worst case, i.e., re-use the ‘Systematic error related to beam peak search’ of 0.5dB from the REFSENS test case.

Many MU elements should be applicable the same way for Mode 2 demodulation test cases as they were for the REFSENS test case as it was agreed that the UE RF system is applicable to demodulation test cases, e.g., it was decided to apply the multiple measurement antenna uncertainty to demodulation test cases even though this concept was introduced for spurious emissions testing. Due to the agreed re-use of the UE RF test methodology and the similarity of Mode 2 test cases with REFSENS test case, it is proposed to re-use most descriptions and uncertainty values of MU elements. Specifically, the descriptions and uncertainty values of MU elements: Measure distance uncertainty, Standing wave between the DUT and measurement antenna, Phase curvature, Random uncertainty, Influence of the XPD, Multiple measurement antenna uncertainty, Uncertainty of the Network Analyzer, Uncertainty of the absolute gain of the calibration antenna, Standing wave between reference calibration antenna and measurement antenna, Influence of the calibration antenna feed cable, RF leakage (from measurement antenna to the receiver/transmitter) for Mode 2 test cases should be re-used from REFSENS test case. However, since Mode 2 demodulation test cases assume different power levels, the following MU elements need to be re-evaluated: Mismatch, gNB uncertainty on absolute level, Amplifier uncertainties, Insertion Loss Variation.

Proposal 3: Re-use the descriptions and uncertainty values of MU elements for Mode 2 test cases from REFSENS test case: Measure distance uncertainty, Standing wave between the DUT and measurement antenna, Phase curvature, Random uncertainty, Influence of the XPD, Multiple measurement antenna uncertainty, Uncertainty of the Network Analyzer, Uncertainty of the absolute gain of the calibration antenna, Standing wave between reference calibration antenna and measurement antenna, Influence of the calibration antenna feed cable, RF leakage (from measurement antenna to the receiver/transmitter)

Proposal 4: Re-evaluate the uncertainty values of MU elements for Mode 2 test cases: Mismatch, gNB uncertainty on absolute level, Amplifier uncertainties, Insertion Loss Variation,

For Mode 1 test cases, many system/equipment uncertainties should not be re-used from REFSENS without further investigations as fixed SNR vs absolute power levels could yield different uncertainties.

Proposal 5: Further evaluate the descriptions and uncertainty values of the following MU elements for Mode 1 test cases: Measure distance uncertainty, Mismatch, Standing wave between the DUT and measurement antenna, Phase curvature, Amplifier uncertainties, Influence of the XPD, Insertion Loss Variation, RF leakage (from measurement antenna to the receiver/transmitter), Multiple measurement antenna uncertainty, Uncertainty of the absolute gain of the calibration antenna, Phase centre offset of calibration antenna, Standing wave between reference calibration antenna and measurement antenna, Influence of the calibration antenna feed cable, Insertion Loss Variation, Positioning misalignment, DUT repositioning, Misalignment of positioning system, Positioning and pointing misalignment between the reference antenna and the measurement antenna

The Quality of Quiet Zone validation approach for Mode 1 test cases also needs to be re-evaluated. Feedback from OEMs and chipset manufacturers is requested whether the assumptions made in [7] are still applicable to demodulation testing, i.e., cross-polarized transmission and no spatial MIMO:

- Proposal 1. Adopt “pure baseband” testing methodology for demodulation performance testing in FR2. - Assumes max rank 2 with cross polarized transmission - Rank 2 spatial MIMO would not be included

Proposal 6: OEMs and chipset vendors to confirm that demodulation testing only assumes cross-polarized transmission and no spatial MIMO.

The current idea is that the Quality of Quiet zone MU for Mode 1 test cases could be considered negligible if the transmission is based on cross polarization since relative signal ripple within the quiet zone has no effect on SNR if the antennas are co-located, i.e., transmission based on cross polarization.

Observation 3: Depending on the feedback in Proposal 6, the Quality of QZ MU for Mode 1 test cases could be considered zero.

Table 1: Uncertainty Contributions for Mode 1 Demodulation Test Cases

| Related to 8x2 Antenna Assumption | Related to Equipment/ Component Datasheets | UID | Uncertainty source | Proposal |
| --- | --- | --- | --- | --- |
|  |  | Signal-to-noise ratio uncertainty |  |  |
|  |  | Stage 2: DUT measurement |  |  |
| yes |  | 1 | Positioning misalignment | Re-Use from REFSENS |
|  |  | 2 | Measure distance uncertainty | Needs to be re-evaluated |
|  |  | 3 | Quality of Quiet Zone |  |
|  | yes | 4 | Mismatch | Needs to be re-evaluated |
|  |  | 5 | Standing wave between the DUT and measurement antenna | Needs to be re-evaluated |
|  | yes | 6 | gNB emulator SNR uncertainty | From datasheets |
|  |  | 7 | Phase curvature | Needs to be re-evaluated |
|  | yes | 8 | Amplifier uncertainties | Needs to be re-evaluated |
|  |  | 9 | Random uncertainty | Re-Use from REFSENS |
|  |  | 10 | Influence of the XPD | Needs to be re-evaluated |
|  |  | 11 | Insertion Loss Variation | Needs to be re-evaluated |
|  |  | 12 | RF leakage (from measurement antenna to the receiver/transmitter) | Needs to be re-evaluated |
|  | yes | 13 | Multiple measurement antenna uncertainty | Needs to be re-evaluated |
| yes |  | 14 | DUT repositioning | Re-Use from REFSENS |
|  |  | Stage 1: Calibration measurement |  |  |
|  | Combined with #4 | 15 | Mismatch | Needs to be re-evaluated |
|  | Combined with #8 | 16 | Amplifier Uncertainties | Needs to be re-evaluated |
| yes | yes | 17 | Misalignment of positioning System | Re-Use from REFSENS |
|  | yes | 18 | Uncertainty of the Network Analyzer | Re-Use from REFSENS |
|  | yes | 19 | Uncertainty of the absolute gain of the calibration antenna | Needs to be re-evaluated |
| yes |  | 20 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | Re-Use from REFSENS |
|  |  | 21 | Phase centre offset of calibration antenna | Needs to be re-evaluated |
|  |  | 22 | Quality of quiet zone for calibration process |  |
|  |  | 23 | Standing wave between reference calibration antenna and measurement antenna | Needs to be re-evaluated |
|  | yes | 24 | Influence of the calibration antenna feed cable | Needs to be re-evaluated |
|  |  | 25 | Insertion Loss Variation | Needs to be re-evaluated |
|  |  | Systematic uncertainties |  |  |
|  |  | 26 | Impact on non-ideal isolation between branches for the wireless cable mode | 0.45dB (for Rank 1) 0.6dB (for Rank 2) |

Table 2: Uncertainty Contributions for Mode 2 Demodulation Test Cases

| Related to 8x2 Antenna Assumption | Related to Equipment/ Component Datasheets | UID | Uncertainty source | Proposal |
| --- | --- | --- | --- | --- |
|  |  | Stage 2: DUT measurement |  |  |
| yes |  | 1 | Positioning misalignment | Re-Use from REFSENS |
|  |  | 2 | Measure distance uncertainty | Re-use from REFSENS |
|  |  | 3 | Quality of Quiet Zone | Re-use from REFSENS |
|  | yes | 4 | Mismatch | Needs to be re-evaluated |
|  |  | 5 | Standing wave between the DUT and measurement antenna | Re-use from REFSENS |
|  | yes | 6 | gNB uncertainty on absolute level | Needs to be re-evaluated |
|  |  | 7 | Phase curvature | Re-use from REFSENS |
|  | yes | 8 | Amplifier uncertainties | Needs to be re-evaluated |
|  |  | 9 | Random uncertainty | Re-use from REFSENS |
|  |  | 10 | Influence of the XPD | Re-use from REFSENS |
|  |  | 11 | Insertion Loss Variation | Needs to be re-evaluated |
|  |  | 12 | RF leakage (from measurement antenna to the receiver/transmitter) | Needs to be re-evaluated |
|  | yes | 13 | Multiple measurement antenna uncertainty | Re-use from REFSENS |
| yes |  | 14 | DUT repositioning | Re-Use from REFSENS |
|  |  | Stage 1: Calibration measurement |  |  |
|  | Combined with #4 | 15 | Mismatch | Needs to be re-evaluated |
|  | Combined with #8 | 16 | Amplifier Uncertainties | Needs to be re-evaluated |
| yes | yes | 17 | Misalignment of positioning System | Re-Use from REFSENS |
|  | yes | 18 | Uncertainty of the Network Analyzer | Re-use from REFSENS |
|  | yes | 19 | Uncertainty of the absolute gain of the calibration antenna | Re-use from REFSENS |
| yes |  | 20 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | Re-Use from REFSENS |
|  |  | 21 | Phase centre offset of calibration antenna | Re-use from REFSENS |
|  |  | 22 | Quality of quiet zone for calibration process | Re-use from REFSENS |
|  |  | 23 | Standing wave between reference calibration antenna and measurement antenna | Re-use from REFSENS |
|  | yes | 24 | Influence of the calibration antenna feed cable | Re-use from REFSENS |
|  |  | 25 | Insertion Loss Variation | Re-use from REFSENS |
|  |  | Systematic uncertainties |  |  |
| yes |  | 26 | Systematic error related to beam peak search | Feedback from OEMs |
|  |  | 27 | Impact on non-ideal isolation between branches for the wireless cable mode | 0.45dB (for Rank 1) 0.6dB (for Rank 2) |

# Conclusion

The following observations and proposals were made in this contribution

Observation 1: MU elements related to positioning and/or misalignments (Positioning misalignment, DUT repositioning, Misalignment of positioning system, Positioning and pointing misalignment between the reference antenna and the measurement antenna) for Mode 2 test cases can be considered negligible if the REFSENS condition was verified without any re-positioning between the verification and the start of the demodulation test case.

Observation 2: MU element descriptions and uncertainties related to positioning and/or misalignments (Positioning misalignment, DUT repositioning, Misalignment of positioning system, Positioning and pointing misalignment between the reference antenna and the measurement antenna) for Mode 2 test cases should be re-used from the REFSENS test case if the UE has been re-positioned between the REFSENS verification and the start of the demodulation test case.

Proposal 1: Regarding MU elements related to positioning and/or misalignments (Positioning misalignment, DUT repositioning, Misalignment of positioning system, Positioning and pointing misalignment between the reference antenna and the measurement antenna) for Mode 2 test cases: re-use from the REFSENS test case.

Proposal 2: Feedback from the OEMs is requested: allow more time to further study the ‘Systematic error related to beam peak search’ for Mode 2 demodulation test cases or assume the worst case, i.e., re-use the ‘Systematic error related to beam peak search’ of 0.5dB from the REFSENS test case.

Proposal 3: Re-use the descriptions and uncertainty values of MU elements for Mode 2 test cases from REFSENS test case: Measure distance uncertainty, Standing wave between the DUT and measurement antenna, Phase curvature, Random uncertainty, Influence of the XPD, Multiple measurement antenna uncertainty, Uncertainty of the Network Analyzer, Uncertainty of the absolute gain of the calibration antenna, Standing wave between reference calibration antenna and measurement antenna, Influence of the calibration antenna feed cable, RF leakage (from measurement antenna to the receiver/transmitter)

Proposal 4: Re-evaluate the uncertainty values of MU elements for Mode 2 test cases: Mismatch, gNB uncertainty on absolute level, Amplifier uncertainties, Insertion Loss Variation

Proposal 5: Further evaluate the descriptions and uncertainty values of the following MU elements for Mode 1 test cases: Measure distance uncertainty, Mismatch, Standing wave between the DUT and measurement antenna, Phase curvature, Amplifier uncertainties, Influence of the XPD, Insertion Loss Variation, RF leakage (from measurement antenna to the receiver/transmitter), Multiple measurement antenna uncertainty, Uncertainty of the absolute gain of the calibration antenna, Phase centre offset of calibration antenna, Standing wave between reference calibration antenna and measurement antenna, Influence of the calibration antenna feed cable, Insertion Loss Variation, Positioning misalignment, DUT repositioning, Misalignment of positioning system, Positioning and pointing misalignment between the reference antenna and the measurement antenna

Proposal 6: OEMs and chipset vendors to confirm that demodulation testing only assumes cross-polarized transmission and no spatial MIMO.

Observation 3: Depending on the feedback in Proposal 6, the Quality of QZ MU for Mode 1 test cases could be considered zero.

# References

- R5-201156, NR FR2 Demodulation MU Table, Keysight Technologies, 3GPP TSG-RAN5 Meeting #86-e, February 2020

- TR 38.903, Derivation of test tolerances and measurement uncertainty for User Equipment (UE) conformance test cases, V16.3.0 (2020-03)

- R5-191628, MU Elements: documentation of justification and interdependence, Keysight Technologies, 3GPP TSG-RAN5 Meeting #82, February 2019

- R4-1816261, On Beam Peak Search Measurement Grids, Keysight Technologies, 3GPP TSG-RAN WG4 Meeting #89, November 2018

- TS 38.101-2, User Equipment (UE) radio transmission and reception; Part 2: Range 2 Standalone, V16.3.0 (2020-03)

- TS 38.521-4: User Equipment (UE) conformance specification; Radio transmission and reception; Part 4: Performance requirements, V16.3.0 (2020-03)

- R4-1801288, WF on performance testing for FR2, Intel, Qualcomm, 3GPP TSG-RAN WG4 NR Adhoc 1801, January 2018

- R5-192820, WF on Maximum Test System Uncertainty Evolution for FR2, Keysight Technologies, Rohde &Schwarz, Anritsu, Samsung, MVG, ETS-Lindgren, PCTEST Engineering Lab, Ericsson, 3GPP TSG-RAN5 Meeting #82, February 2019