3GPP TSG-RAN5 Meeting #90-e	R5-211192
Electronic Meeting, 22 February – 5 March 2021

Agenda item:	5.3.2.17

Source:	Keysight Technologies

Title:	On n259 QoQZ and XPD MU

Document for:	Discussion and Endorsement

# Introduction

This contribution is addressing QoQZ and XPD MU for the n259 band.

# Discussion

In the RAN4#97-e WF on n262 [1], it was noted to complete the work on Band n259 (40-43.5GHz)

Observation 1 : In order to introduce 5G (Band n259; 40GHz band) service in Japan, it is required to complete RAN5 work by Feb-2021 at the latest.

The MU evaluation requires a maximum test frequency of 44.3GHz [2] for FR2c.

In this contribution, we present the results of an EIRP&TRP QoQZ MU campaign with 238 measurement points for a total of 5 test frequencies: 23.45, 32.125, 40.8, 44.3, and 49GHz. The results at 44.3GHz are addressing the n259 MUs while the results at 49GHz will eventually address the n262 (47.2-48.2GHz) MU. For this evaluation, a single, wide-band probe antenna was used to cover the entire frequency range.

The quiet zone size evaluated in this contribution is a 30cm sphere and the procedure is following the procedure and requirements outlined in [3].

The EIRP/TRP QoQZ standard deviations/MUs from this QoQZ campaign are tabulated in Table 1.

Table 1: EIRP and TRP std. deviations from the full QoQZ scan

| Frequency [GHz] | Standard Deviation [dB] |  |  |  |
| --- | --- | --- | --- | --- |
|  | EIRP (P1 only) | EIRP (P1-P7) | TRP (P1 only) | TRP (P1-P7) |
| 23.45 | 0.36 | 0.46 | 0.24 | 0.34 |
| 32.125 | 0.23 | 0.47 | 0.15 | 0.39 |
| 40.8 | 0.32 | 0.34 | 0.22 | 0.27 |
| 44.3 | 0.25 | 0.44 | 0.18 | 0.35 |
| 49.0 | 0.29 | 0.46 | 0.13 | 0.34 |

Clearly, the EIRP MUs are a generally a bit higher than the TRP MUs but these values are all within the 0.6dB example value defined in RAN5 for QZ up to 30 cm for Stage 2 (P1-P7) and within the 0.4dB example value for Stage 1 (P1 only). As such, no increase in QoQZ MU is necessary for FR2c.

From those QoQZ measurement results, the XPD MUs were calculated [4] and tabulated in Table 2

Table 2: XPD MUs from QoQZ scan

| Frequency [GHz] | XPD MU [dB] |
| --- | --- |
| 23.45 | 0.01 |
| 32.125 | 0.00 |
| 40.8 | 0.01 |
| 44.3 | 0.02 |
| 49.0 | 0.02 |

The XPD MU has increased slightly but given the small MU, it is proposed to leverage the existing sample MU of 0.01dB.

Proposal 1: For FR2c, keep the 0.6dB example QoQZ MU for Stage 2 (P1-P7) and the 0.4dB example QoQZ MU for Stage 1 (P1 only)

Proposal 2: For FR2c, keep the 0.01dB example XPD MU

# Conclusion

The following observations and proposals were made in this contribution

Proposal 1: For FR2c, keep the 0.6dB example QoQZ MU for Stage 2 (P1-P7) and the 0.4dB example QoQZ MU for Stage 1 (P1 only)

Proposal 2: For FR2c, keep the 0.01dB example XPD MU

# References

- R5-204851, Update of WF on Band n259, NTT DOCOMO, 3GPP TSG-RAN5 Meeting #88-e, August 2020

- R5-197733, Frequency range of Band n259, NTT DOCOMO, 3GPP TSG-RAN5 Meeting #85, November 2019

- TS 38.521-2, User Equipment (UE) conformance specification; Radio transmission and reception; Part 2: Range 2 Standalone, V16.60 (2020-12)

- TR 38.903, Derivation of test tolerances and measurement uncertainty for User Equipment (UE) conformance test cases, V16.6.0 (2020-12)