3GPP TSG RAN WG5 Meeting #94-e 	R5-221623

Electronic Meeting, 21st February – 4th March 2022

Source:				MVG Industries, Rohde & Schwarz

Title:					TP to TR 38.834 for Annex B MU tables

Agenda Item:		5.3.39.1

Document for:	Endorsement

# Introduction

The work item led by RAN4 for NR FR1 TRP-TRS [1] was approved during March plenary meeting, further updated during June plenary, which includes a secondary responsibility for RAN5 to develop the Measurement Uncertainty (MU) assessment. During RAN5#92-e, the work plan and initial agreements on this MU work were agreed in [2].

The aim of this contribution is to propose the Measurement Uncertainty tables including a preliminary MU estimation for Annex B of TR 38.834 [3].

The following changes have been made in relation to TR 37.544 v16.2.0 [5]

- Frequency Response term has been added

- Measurement Distance Term updated to reflect changes in R5-221624 [6].

- Sensitivity output level step resolution changed to 0.5 dB as defined in TR 38.834.

- Quality of the quiet zone aligned with new definition in [6]

It is understood that RAN5 will finalize the MU estimation following the performance requirements definition in a RAN5 test specification. This will allow MTSU to be defined and any revised MU estimation shall be based on detailed and documented analyses in RAN5.

Proposal 1: Endorse the text proposal below.

# Conclusion

The following proposal is made in this contribution:

Proposal 1: Endorse the text proposal below.

# References

- RP‑211158, “Revised WID: Introduction of UE TRP (Total Radiated Power) and TRS (Total Radiated Sensitivity) requirements and test methodologies for FR1 (NR SA and EN-DC)”, vivo, 3GPP TSG RAN Meeting #92-e, June 2021.

- R5-2145771, “MU workplan for NR FR1 TRP-TRS”, Rohde & Schwarz, 3GPP TSG RAN5 Meeting #92-e, August 2021.

- 3GPP TR 38.834 v0.3.0 (January 2022)

- R5-218480, “TP to TR 38.834 on skeleton for Annex B”, Rohde & Schwarz, 3GPP TSG RAN5 Meeting #93-e, November 2021.

- 3GPP TR 37.544 v16.2.0 (September 2021)

- R5-221624, “TP to TR 38.834 on MU General section and descriptions”, Rohde & Schwarz, MVG Industries, 3GPP TSG RAN5 Meeting #94-e, February 2022.

# Appendix – Text Proposal to TR 38.834 v0.3.0

< Unchanged Text Deleted >

# B.3	Total Radiated Power (TRP)

< Beginning of Changes >

The uncertainty contributions related to TRP are listed in Table B.3-1.  A preliminary example uncertainty budget is presented in Table B.3-2.

Table B.3-1 Uncertainty contributions in TRP measurement for anechoic chamber method

| UID | Description of uncertainty contribution | Details in clause |
| --- | --- | --- |
| Stage 2: DUT measurement (Figure TBD) |  |  |
| 1 | Mismatch of receiver chain | B.2.1 |
| 2 | Insertion loss of receiver chain | B.2.2 |
| 3 | Influence of the measurement antenna cable | B.2.3 |
| 4 | Measurement Receiver: uncertainty of the absolute level | B.2.4 |
| 5 | Measurement distance | B.2.7 |
| 6 | Quality of quiet zone | B.2.8 |
| 7 | DUT Tx-power drift | B.2.9 |
| 8 | Uncertainty related to the use of phantoms | B.2.11 |
| 9 | Coarse sampling grid | B.2.12 |
| 10 | Random uncertainty | B.2.13 |
| 11 | Frequency Response | B.2.14 |
| Stage 1: Calibration measurement, network analyzer method (Figure 7.3-1) |  |  |
| 12 | Uncertainty of network analyzer | B.2.15 |
| 13 | Mismatch of receiver chain | B.2.1 |
| 14 | Insertion loss of receiver chain | B.2.2 |
| 15 | Mismatch in the connection of calibration antenna | B.2.1 |
| 16 | Influence of the calibration antenna feed cable | B.2.3 |
| 17 | Influence of the measurement antenna cable | B.2.3 |
| 18 | Uncertainty of the absolute gain/ radiation efficiency of the calibration antenna | B.2.16 |
| 19 | Measurement distance | B.2.7 |
| 20 | Quality of the Quiet Zone | B.2.8 |

Table B.3-2 Preliminary example of uncertainty budget for TRP hand only (browsing mode) measurement for anechoic chamber method for NR FR1 bands

| UID | Uncertainty Source | Comment | Uncertainty Value [dB] | Prob Distr | Div | ci | Standard Uncertainty [dB] |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |
| 1 | Mismatch of receiver chain | Гpower meter <0.05 Гmeasurement antenna <0.16 | 0.07 | U-shaped | 1.41 | 1 | 0.05 |
| 2 | Insertion loss of receiver chain | Systematic with Stage 1 (=> cancels) | 0 | Rectangular | 1.73 | 1 | 0.00 |
| 3 | Influence of the measurement antenna cable | Systematic with Stage 1 (=> cancels) | 0 | Rectangular | 1.73 | 1 | 0.00 |
| 4 | Measurement Receiver: uncertainty of the absolute level | Power Meter | 0.06 | Rectangular | 1.73 | 1 | 0.03 |
| 5 | Measurement distance | d=1.6m, Δd=0.05m | 0.27 | Rectangular | 1.73 | 1 | 0.16 |
| 6 | Quality of quiet zone | Surface standard deviation of power measurements in ripple test | 0.5 | Actual | 1 | 1 | 0.50 |
| 7 | DUT Tx-power drift | Drift | 0.2 | Rectangular | 1.73 | 1 | 0.12 |
| 8 | Uncertainty related to the use of phantoms | U [dB] = 0.20 U [dB] = 0.15 | 0.32 | Rectangular | 1.73 | 1 | 0.18 |
| 9 | Coarse sampling grid | Negligible 15° sampling grid | 0 | Actual | 1 | 1 | 0.00 |
| 10 | Random Uncertainty | Monoblock, clamshell and PDA design used for testing | 0.81 | Rectangular | 1.73 | 1 | 0.47 |
| 11 | Frequency Response | Average path loss corrected | 0 | Rectangular | 1.73 | 1 | 0.00 |
|  | Stage 1: Calibration measurement, network analyzer method |  |  |  |  |  |  |
| 12 | Uncertainty of network analyzer | Manufacturer's uncertainty | 0.5 | Rectangular | 1.73 | 1 | 0.29 |
| 13 | Mismatch of receiver chain | Taken into account in VNA uncertainty term | 0 | U-shaped | 1.41 | 1 | 0.00 |
| 14 | Insertion loss of receiver chain | Systematic with Stage 2 (=> cancels) | 0 | Rectangular | 1.73 | 1 | 0.00 |
| 15 | Mismatch in the connection of calibration antenna | Taken in to account in VNA setup uncertainty | 0 | U-shaped | 1.41 | 1 | 0.00 |
| 16 | Influence of the calibration antenna feed cable | Gain calibration with a dipole | 0.3 | Rectangular | 1.73 | 1 | 0.17 |
| 17 | Influence of the measurement antenna cable | Systematic with Stage 2 (=> cancels) | 0 | Rectangular | 1.73 | 1 | 0.00 |
| 18 | Uncertainty of the absolute gain/ radiation efficiency of the calibration antenna | Calibration certificate | 0.5 | Rectangular | 1.73 | 1 | 0.29 |
| 19 | Measurement distance | Dipole: aligned with phase center | 0 | Rectangular | 1.73 | 1 | 0.00 |
| 20 | Quality of the Quiet Zone | Peak-to-null ripple | 0.5 | Rectangular | 1.73 | 1 | 0.29 |
| Combined standard uncertainty |  |  |  |  |  |  | 0.91 |
| Expanded uncertainty (Confidence interval of 95 %) |  |  |  |  |  |  | 1.78 |

# B.4	Total Radiated Sensitivity (TRS)

The uncertainty contributions related to TRS are listed in Table B.4-1. A preliminary example uncertainty budget is presented in Table B.4-2.

Table B.4-1 Uncertainty contributions in TRS measurement for anechoic chamber method

| UID | Description of uncertainty contribution | Details in clause |
| --- | --- | --- |
| Stage 2: DUT measurement (Figure TBD) |  |  |
| 1 | Mismatch of transmitter chain | B.2.1 |
| 2 | Insertion loss of transmitter chain | B.2.2 |
| 3 | Influence of the measurement antenna cable | B.2.3 |
| 4 | Communication Tester: uncertainty of the absolute output level | B.2.5 |
| 5 | Sensitivity measurement: output level step resolution | B.2.6 |
| 6 | Measurement distance | B.2.7 |
| 7 | Quality of quiet zone | B.2.8 |
| 8 | DUT sensitivity drift | B.2.10 |
| 9 | Uncertainty related to the use of phantoms | B.2.11 |
| 10 | Coarse sampling grid | B.2.12 |
| 11 | Random uncertainty | B.2.13 |
| 12 | Frequency Response | B.2.14 |
| Stage 1: Calibration measurement, network analyzer method (Figure 7.3-1) |  |  |
| 13 | Uncertainty of network analyzer | B.2.15 |
| 14 | Mismatch of transmitter chain | B.2.1 |
| 15 | Insertion loss of transmitter chain | B.2.2 |
| 16 | Mismatch in the connection of calibration antenna | B.2.1 |
| 17 | Influence of the calibration antenna feed cable | B.2.3 |
| 18 | Influence of the measurement antenna cable | B.2.3 |
| 19 | Uncertainty of the absolute gain/radiation efficiency of the calibration antenna | B.2.16 |
| 20 | Measurement distance | B.2.7 |
| 21 | Quality of quiet zone | B.2.8 |

Table B.4-2: Preliminary example of uncertainty budget for TRS hand only (browsing mode) measurement for anechoic chamber method for NR FR1 bands

| UID | Uncertainty Source | Comment | Uncertainty Value [dB] | Prob Distr | Div | ci | Standard Uncertainty [dB] |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |
| 1 | Mismatch of transmitter chain | ГCommTester <0.13 Г antenna connection <0.03 | 0.07 | U-shaped | 1.41 | 1 | 0.05 |  |
| 2 | Insertion loss of transmitter chain | Systematic with Stage 1 (=> cancels) | 0 | Rectangular | 1.73 | 1 | 0.00 |  |
| 3 | Influence of the measurement antenna cable | Systematic with Stage 1 (=> cancels) | 0 | Rectangular | 1.73 | 1 | 0.00 |  |
| 4 | Communication Tester: uncertainty of the absolute output level | Manufacturer's data sheet | 1 | Rectangular | 1.73 | 1 | 0.58 |  |
| 5 | Sensitivity measurement: output level step resolution | Step of 0.5 dB | 0.25 | Rectangular | 1.73 | 1 | 0.14 |  |
| 6 | Measurement distance | d=1.6m, Δd=0.05m | 0.27 | Rectangular | 1.73 | 1 | 0.16 |  |
| 7 | Quality of quiet zone | Surface standard deviation of power measurements in ripple test | 0.5 | Actual | 1 | 1 | 0.5 |  |
| 8 | DUT sensitivity drift | Drift measurement | 0.2 | Rectangular | 1.73 | 1 | 0.12 |  |
| 9 | Uncertainty related to the use of phantoms | U [dB] = 0.20 U [dB] = 0.15 | 0.32 | Rectangular | 1.73 | 1 | 0.18 |  |
| 10 | Coarse sampling grid | 30° sampling grid | 0.15 | Actual | 1 | 1 | 0.15 |  |
| 11 | Random uncertainty | Monoblock, clamshell and PDA used for testing | 0.91 | Rectangular | 1.73 | 1 | 0.53 |  |
| 12 | Frequency Response | Included in the output level step resolution | 0 | Rectangular | 1.73 | 1 | 0.00 |  |
| Stage 1: Calibration measurement, network analyzer method |  |  |  |  |  |  |  |  |
| 13 | Uncertainty of network analyzer | Manufacturer's uncertainty calculator, covers NA setup | 0.5 | Rectangular | 1.73 | 1 | 0.29 |  |
| 14 | Mismatch of transmitter chain | Taken in to account in VNA setup uncertainty | 0 | U-shaped | 1.41 | 1 | 0.00 |  |
| 15 | Insertion loss of transmitter chain | Systematic with Stage 2 (=> cancels) | 0 | Rectangular | 1.73 | 1 | 0.00 |  |
| 16 | Mismatch in the connection of calibration antenna | Taken in to account in VNA setup uncertainty | 0 | U-shaped | 1.41 | 1 | 0.00 |  |
| 17 | Influence of the calibration antenna feed cable | Gain calibration with dipole | 0.3 | Rectangular | 1.73 | 1 | 0.17 |  |
| 18 | Influence of the measurement antenna cable | Systematic with Stage 2 (=> cancels) | 0 | Rectangular | 1.73 | 1 | 0.00 |  |
| 19 | Uncertainty of the absolute gain/ radiation efficiency of the calibration antenna | Calibration certificate | 0.5 | Rectangular | 1.73 | 1 | 0.29 |  |
| 20 | Measurement distance | Dipole: aligned with phase center | 0 | Rectangular | 1.73 | 1 | 0.00 |  |
| 21 | Quality of quiet zone | Peak-to-null ripple | 0.5 | Rectangular | 1.73 | 1 | 0.29 |  |
| Combined standard uncertainty |  |  |  |  |  |  | 1.12 |
| Expanded uncertainty (Confidence interval of 95 %) |  |  |  |  |  |  | 2.20 |

< End of Changes >