3GPP TSG RAN Meeting #96-e	 	R5-225643

Electronic Meeting, 15th – 26th August 2022

Source:		Rohde & Schwarz

Title:	On the MU for FR2 PC1 TRx test cases

Agenda Item:		5.4.17

Document for:	Discussion and Endorsement

# Introduction

At previous meetings, estimation for the SNR for PC1 MOP TCs have been discussed [1, 2, 3]. In this discussion paper, we present the results of our analysis of for influence of noise and relaxation based on the common assumptions.

# Discussion

In this paper we focus our analysis on the requirements for MOP and Refsens in band n257.

## Relaxation

The requirement for MOP from TS 38.101-2 [4] is as follows:

Table 6.2.1.1-2: UE maximum output power limits for power class 1

| Operating band | Max TRP (dBm) | Max EIRP (dBm) |
| --- | --- | --- |
| n257 | 35 | 55 |
| n258 | 35 | 55 |
| n260 | 35 | 55 |
| n261 | 35 | 55 |
| n262 | 35 | 55 |

We have analyzed the SNR for MOP for an UE carrier rms level of 55 dBm (as shown above) equivalent to 68 dBm PEP.

Given that this maximum power is higher than the 43 dBm max from Power Class 3, no relaxation of the requirement is needed, since this higher power can be compensated for by utilizing higher input attenuation in the measurement equipment. Thus no relaxation is needed on the requirement.

Proposal 1: No relaxation is needed for the MOP TC in PC1.

For the Refsens TC, the requirements from TS 38.101-2 [4] are shown in the following table:

Table 7.3.2.1-1: Reference sensitivity for power class 1

| Operating band | REFSENS (dBm) / Channel bandwidth |  |  |  |
| --- | --- | --- | --- | --- |
|  | 50 MHz | 100 MHz | 200 MHz | 400 MHz |
| n257 | -97.5 | -94.5 | -91.5 | -88.5 |
| n258 | -97.5 | -94.5 | -91.5 | -88.5 |
| n260 | -94.5 | -91.5 | -88.5 | -85.5 |
| n261 | -97.5 | -94.5 | -91.5 | -88.5 |
| n262 | -92.5 | -89.5 | -86.5 | -83.5 |
| NOTE 1: The transmitter shall be set to PUMAX as defined in clause 6.2.4 |  |  |  |  |

Compared to the PC3 Refsens values, the PC1 values are 9.2 dB lower. These values can be achieved by either lowering the output power from the system simulator or by increasing the attenuation in the test system. This however has no impact on the testability of the requirement and no relaxation is required.

Proposal 2: No relaxation is needed for the Refsens TC in PC1.

## Measurement Uncertainty

In the following, we summarize the results of our MU assessment for MOP EIRP and TRP and Reference sensitivity EIS. We propose to apply these values for the MU for PC1, FR2a, CHBW ≤ 400 MHz, IFF, max device size ≤ 30 cm, and NTC.

Table 1: Uncertainty assessment for EIRP and TRP measurement (f=23.45GHz, 32.125GHz, 40.8GHz, Quiet Zone size ≤ 30 cm) for PC1 UEs and normal temperature condition

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment | 0.02 | Normal | 2.00 | 0.01 |
| 2 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 3 | Quality of Quiet Zone (NOTE 1) | 0.6 | Actual | 1.00 | 0.6 |
| 4 | Mismatch | 1.30 | Actual | 1.00 | 1.30 |
| 5 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | Uncertainty of the RF power measurement equipment (NOTE 3) | 2.16 | Normal | 2.00 | 1.08 |
| 7 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 8 | Amplifier uncertainties | 2.10 | Normal | 2.00 | 1.05 |
| 9 | Random uncertainty | 0.50 | Normal | 2.00 | 0.25 |
| 10 | Influence of the XPD | 0.01 | U-shaped | 1.41 | 0.00 |
| 11 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 13 | Influence of TRP measurement grid (NOTE 4) | 0.25 | Actual | 1 | 0.25 |
| 14 | Influence of beam peak search grid (NOTE 5) | 0.00 | Actual | 1 | 0.00 |

| 15 | Multiple measurement antenna uncertainty (NOTE 9) | 0.15 | Actual | 1 | 0.15 |
| --- | --- | --- | --- | --- | --- |
| 16 | DUT repositioning | 0.00 (NOTE 4) 0.35 (NOTE 5) | Rectangular | 1.73 | 0.00 (NOTE 4) 0.20 (NOTE 5) |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 17 | Mismatch | 0.00 | U-shaped. | 1.41 | 0.00 |
| 18 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 19 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 20 | Uncertainty of the Network Analyzer | 1.50 | Normal | 2.00 | 0.75 |
| 21 | Uncertainty of the absolute gain of the calibration antenna | 0.60 | Normal | 2.00 | 0.30 |
| 22 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.01 | Rectangular | 1.73 | 0.00 |
| 23 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 24 | Quality of quiet zone for calibration process (NOTE 1) | 0.4 | Actual | 1.00 | 0.4 |
| 25 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 26 | Influence of the calibration antenna feed cable | 0.14 | Normal | 2.00 | 0.07 |
| 27 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
|  | Systematic uncertainties (NOTE 6) |  |  |  | Value |
| 28 | Systematic error due to TRP calculation/quadrature (NOTE 4) |  |  |  | 0.00 |
| 29 | Influence of noise (23.45GHz <= f <= 32.125GHz) |  |  |  | 0.13 |
| 29 | Influence of noise (32.125GHz < f <= 40.8GHz) |  |  |  | FFS |
| 30 | Systematic error related to beam peak search (NOTE 5) |  |  |  | 0.7 |
| Total measurement uncertainty |  |  |  |  | Value |
| EIRP Expanded uncertainty (23.45GHz <= f <= 32.125GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 5.33 |
| EIRP Expanded uncertainty (32.125GHz < f <= 40.8GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | FFS |
| TRP Expanded uncertainty (23.45GHz <= f <= 32.125GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 4.64 |
| TRP Expanded uncertainty (32.125GHz < f <= 40.8GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | FFS |
| NOTE 1: The quality of quiet zone is the same for EIRP and TRP. Value based on procedure defined in clause D.2 of TR 38.810 for Quiet Zone size less or equal to 30 cm. NOTE 2: The analysis was done only for the case of operating at max output power, in-band, non-CA. NOTE 3: The assessment assumes maximum DUT output power. NOTE 4: This contributor shall only be considered for TRP measurements. NOTE 5: This contributor shall only be considered for EIRP measurements. NOTE 6: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 7: Void. NOTE 8: Void NOTE 9: Applies to the system which has a structure of mechanical feed antenna positioning. |  |  |  |  |  |

Proposal 3: Apply the values in Table 1 for the MU of MOP TC for PC1, FR2a, CHBW ≤ 400 MHz, IFF, max device size ≤ 30 cm, and NTC.

Table 2: Uncertainty assessment for EIS measurement (f=23.45GHz, 32.125GHz, 40.8GHz, Quiet Zone size ≤ 30 cm) for PC1 UEs and normal temperature condition

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment | 0.02 | Normal | 2.00 | 0.01 |
| 2 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 3 | Quality of Quiet Zone (NOTE 7) | 0.6 | Actual | 1.00 | 0.6 |
| 4 | Mismatch | 1.30 | Actual | 1.00 | 1.30 |
| 5 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | gNB uncertainty on absolute level | 2.9 | Normal | 2.00 | 1.45 |
| 7 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 8 | Amplifier uncertainties | 2.1 | Normal | 2.00 | 1.05 |
| 9 | Random uncertainty | 0.50 | Normal | 2.00 | 0.25 |
| 10 | Influence of the XPD | 0.01 | U-shaped | 1.41 | 0.00 |
| 11 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 13 | Multiple measurement antenna uncertainty (NOTE 6) | 0.15 | Actual | 1.00 | 0.15 |
| 14 | DUT repositioning | 0.00 (NOTE 4) 0.35 (NOTE 5) | Rectangular | 1.73 | 0.00 (NOTE 4) 0.20 (NOTE 5) |
| 15 | Influence of spherical coverage grid (NOTE 4) | 0.13 | Actual | 1 | 0.13 |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 16 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 17 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 18 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 19 | Uncertainty of the Network Analyzer | 1.50 | Normal | 2.00 | 0.75 |
| 20 | Uncertainty of the absolute gain of the calibration antenna | 0.60 | Normal | 2.00 | 0.30 |
| 21 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.01 | Rectangular | 1.73 | 0.00 |
| 22 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 23 | Quality of quiet zone for calibration process (NOTE 7) | 0.4 | Actual | 1.00 | 0.4 |
| 24 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 25 | Influence of the calibration antenna feed cable | 0.14 | Normal | 2.00 | 0.07 |
| 26 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
|  | Systematic uncertainties (NOTE 3) |  |  |  | Value |
| 27 | Systematic error related to beam peak search (NOTE 5) |  |  |  | 0.7 |
| 28 | Systematic error related to EIS spherical coverage (NOTE 4) |  |  |  | DL power step size, 0.2 |

| Total measurement uncertainty | Value |
| --- | --- |
| EIS Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | 5.58 |
| EIS Spherical coverage Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] | FFS |
| NOTE 1: The analysis was done only for the case of operating at max output power, in-band, non-CA. NOTE 2: Void. NOTE 3: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 4: This contributor shall only be considered for spherical EIS measurements. NOTE 5: This contributor shall only be considered for EIS measurements. NOTE 6: Applies to the system which has a structure of mechanical feed antenna positioning. NOTE 7: Value based on procedure defined in clause D.2 of TR 38.810 for Quiet Zone size less or equal to 30 cm. |  |

Proposal 4: Apply the values in Table 2 for the MU of Reference Sensitivity TC for PC1, FR2a, CHBW ≤ 400 MHz, IFF, max device size ≤ 30 cm, and NTC.

# Assumptions

| ID | Description | Assumption |
| --- | --- | --- |
| #1 | Frequency ranges under consideration | FR2a |
| #2 | Size of QZ for IFF | 30 cm |
| #3 | UE power class | PC1 |
| #4 | Max EIRP | EIRP MAX= +55 dBm (rms) |
| #5 | Wanted signal PAPR | 13 dB |
| #6 | Temperature range of the test equipment | 20°C – 35°C |
| #7 | Channel bandwidth | 400 MHz |

# Conclusion

In summary no relaxation is needed for the MOP and Refsens TCs in band n257 for PC1.

Proposal 1: No relaxation is needed for the MOP TC in PC1.

Proposal 2: No relaxation is needed for the Refsens TC in PC1.

Proposal 3: Apply the values in Table 1 for the MU of MOP TC for PC1, FR2a, CHBW ≤ 400 MHz, IFF, max device size ≤ 30 cm, and NTC.

Proposal 4: Apply the values in Table 2 for the MU of Reference Sensitivity TC for PC1, FR2a, CHBW ≤ 400 MHz, IFF, max device size ≤ 30 cm, and NTC.

# References

- R5-211275, Measurement uncertainties for PC1 devices, Keysight Technologies, GPP TSG RAN WG5 Meeting #90-e, February 2021

- R5-213314, On the SNR of FR2 TRx test cases for PC1, Rohde & Schwarz, 3GPP TSG RAN WG 5 Meeting #91-e, May 2021

- R5-215481, On the SNR of FR2 TRx test cases for PC1, Rohde & Schwarz, 3GPP TSG RAN WG 5 Meeting #92-e, August 2021

- TS 38.101-2, v17.6.0, 3GPP TSG RAN WG4, June 2022