3GPP TSG-RAN5 Meeting #96-e	R5-225644
Electronic Meeting, 15th August – 26th August 2022

Agenda item:	5.4.17

Source:	Keysight Technologies

Title:	On 40cm QoQZ and XPD MUs

Document for:	Discussion and Endorsement

# Introduction

This contribution is reviewing previously presented validation for 40cm QZ [1][2][3].

# Discussion

In RAN5#95-e, discussions were held whether a single, i.e., for ETC & NTC combined, uncertainty value for 40cm QZ should be defined vs two, i.e., separately for ETC and NTC. Arguments and pros/cons from discussions held in RAN5#95-e are summarized in Table 1.

Table 1: 40cm QoQZ MTSU Approach

| Impact | Single MTSU | Two Separate MTSUs |
| --- | --- | --- |
| Argumentation in RAN5#95-e | “In 30 cm QZ, NTC MU and ETC MU were separately defined due to historical reasons that there was no ETC test system at the first time. However, we should define the common MU value for both of NTC and ETC from now on. Adding and removing ETC enclosure for each test significantly increases test time, so I proposed MTSU calculations based on only ETC” | “we previously tried to optimize the MU for each respective test case given the feedback from industry on high MUs” |
| MTSU | Single MTSU applies to ETC and NTC test cases which simplifies documentation in 38.903 | Defining ETC and NTC MTSUs is aligned with previous approach and could lead to confusion (single MTSU for 40cm, two for 30cm QZ) and would allow test systems dedicated to NTC unnecessary extra margin. |
| TT | Single TT applies to ETC and NTC test cases | Defining separate TTs for ETC and NTC is aligned with previous approach and would give UEs unnecessarily additional margin for NTC test environment |

We still believe that the current approach to define MTSU/TT separately for ETC and NTC test cases for the 40cm QZ is the most appropriate approach to define the most optimized MTSU and to avoid defining unnecessary increases in MTSU/TT for NTC cases.

Proposal 1: In order to define the most optimized MTSU and to avoid defining unnecessary increases in MTSU/TT for NTC cases, define MTSU/TT separately for ETC and NTC test cases for the 40cm QZ.

In [1][2][3], 40cm QoQZ validation measurements were presented based on various assumptions which are outlined in Table 2 with the differences highlighted in yellow. The main difference between these three sets of measurements is the test environment, i.e., two sets of results were performed assuming NTC without any enclosure surrounding the QZ/AUT while the other set of results had an enclosure surrounding the QZ/AUT for ETC type testing.

Table 2: 40cm QoQZ Validation Measurement Assumption

| Assumptions | Keysight [1] | Anritsu [2] | R&S [3] |
| --- | --- | --- | --- |
| Re-Positioning Approach considered | yes | yes | yes |
| Number of test points | 238 | 238 | 238 |
| Reference AUT Coordinates | Position; x; y; z. P1; 0; 0; 0. P2; 20cm; 0; 0. P3; -20cm; 0; 0. P4; 0; 20cm; 0. P5; 0; -20cm; 0. P6; 0; 0; 20cm. P7; 0; 0; -20cm. | Position; x; y; z. P1; 0; 0; 0. P2; 20cm; 0; 0. P3; -20cm; 0; 0. P4; 0; 20cm; 0. P5; 0; -20cm; 0. P6; 0; 0; 16.5cm. P7; 0; 0; -16.5cm. | Position; x; y; z. P1; 0; 0; 0. P2; 20cm; 0; 0. P3; -20cm; 0; 0. P4; 0; 20cm; 0. P5; 0; -20cm; 0. P6; 0; 0; 20cm. P7; 0; 0; -15cm. |
| Test Frequencies [GHz] | In-Band/OOB: 23.45, 32.125, 40.8, 44.3, 49.0 | In-Band/OOB: 23.45, 32.125, 40.8, 44.3 Spurious: 6, 12.75, 66, 80, 87 | In-Band/OOB: 23.45, 32.125, 40.8, 44.3 |
| Test Environment | NTC (no enclosure/bubble) | ETC (with enclosure/bubble) | NTC (no enclosure/bubble) |

Table 3: QoQZ Measurements/Uncertainties

| QoQZ MU | Keysight [1] | Anritsu [2] | R&S [3] |
| --- | --- | --- | --- |
| 40cm QoQZ MU [dB] (In-Band/OOB) | f [GHz]; NTC. EIRP&TRP P1 Only; EIRP P1-P7; TRP P1-P7. 23.45-40.8; ≤0.37; ≤0.50; ≤0.47. 44.3; 0.20; 0.45; 0.30. 49; 0.23; 0.56; 0.32. | f [GHz]; ETC. EIRP&TRP P1 Only; EIRP P1-P7; TRP P1-P7. 23.45-40.8; ≤0.37; ≤0.88; ≤0.56. 44.3; 0.41; 0.85; 0.85. | f [GHz]; NTC. EIRP&TRP P1 Only; EIRP P1-P7; TRP P1-P7. 23.45-40.8; ≤0.25; ≤0.75; ≤0.68. 44.3; 0.21; 0.98; 0.70. 49; N/A; N/A; N/A. |
| Previously Agreed NTC 30cm QoQZ Uncertainty Values [4][5] | f [GHz]; EIRP&TRP P1 Only; EIRP P1-P7; TRP P1-P7. 23.45-40.8; 0.4; 0.6; 0.6. 44.3; 0.5; 0.7; 0.6. |
| Previously Agreed ETC 30cm QoQZ Uncertainty Values [4][6] | f [GHz]; EIRP&TRP P1 Only; EIRP P1-P7; TRP P1-P7. 23.45-44.3; 0.6; 0.9; 0.9. |

The measurements presented from two system vendors [1][2] clearly indicate that the extension from the 30cm to 40cm QZ did not result in any need to increase the uncertainty values currently defined for 30cm as the presented measurements are within those limits. The result from another system vendor [3] show increases in QoQZ MU for 40cm QZ when compared to 30cm QZ, see Table 3.

The XPD MUs for 40cm QZ were presented by a two vendors [1][3] with the measurement results summarized in Table 4.

Table 4: Measured XPD Uncertainty Values for 40cm QZ

| f [GHz] | XPD MU [1] | XPD MU [3] |
| --- | --- | --- |
| 23.45-40.8 | ≤0.03 | ≤0.03 |
| 44.3 | 0.07 | 0.03 |
| 49 | 0.08 | N/A |

Instead of trying to agree different QoQZ and XPD uncertainty values among different vendors, a different approach is proposed below that considers the bigger picture of Maximum Test System Uncertainty (MTSU) and Test Tolerance (TT) instead. Clearly, it would be best for industry to align the MTSU and TT for the 30 cm and 40 cm QZs since existing systems are used/extended for these 2 QZs, i.e., ideally define just a single MTSU and a single TT regardless of QZ (20cm vs 30cm vs 40cm). In Table 4, the MTSUs and TTs (NTC only, highest frequency limited to 40.8 GHz) are presented for the existing 30 cm QZ which assume the following MUs: QoQZ MU (Stage 2, P1-P7) of 0.6 dB and XPD MU of 0.01 dB [4]. They are compared with a theoretical MTSU and TT for the 40cm QZ based on the worst case QoQZ MU (Stage 2, P1-P7) of 0.75 dB, see Table 3, and the worst case XPD MU of 0.07 dB, see Table 4. Clearly, the MTSU and TT differences between the 30 cm and 40 cm QZs are insignificant with less than 0.1dB. These very small differences justify defining the MTSU for the 40 QZ to match the MTSU for the 30 cm QZ. It should be noted that this approach is not applicable to the next larger QZ of 55cm.

Table 4: Comparison of MTSUs and TTs for 30cm QZ and 40cm QZ (based on presented QoQZ and XPD results)

|  |  | MTSU [dB] |  |  | TT [dB] |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Test Case | Frequency | 30cm QZ with QoQZ MU (Stage 2) of 0.6dB and XPD MU of 0.01dB | 40cm QZ with QoQZ MU (Stage 2) of 0.75dB and XPD MU of 0.07dB | Diff. [dB] | 30cm QZ | 40cm QZ | Diff. [dB] |
| MOP-EIRP | 23.45GHz <= f <= 32.125GHz | 4.89 | 4.98 | 0.09 | 2.87 | 2.93 | 0.05 |
|  | 32.125GHz < f <= 40.8GHz | 5.09 | 5.18 | 0.09 | 2.87 | 2.93 | 0.05 |
| MOP-TRP | 23.45GHz <= f <= 32.125GHz | 4.42 | 4.51 | 0.09 | 2.65 | 2.71 | 0.05 |
|  | 32.125GHz < f <= 40.8GHz | 4.62 | 4.71 | 0.09 | 2.77 | 2.83 | 0.05 |
| REFSENS-EIS | 23.45GHz <= f <= 40.8GHz | 5.19 | 5.28 | 0.08 | 2.34 | 2.37 | 0.04 |

Proposal 2: Define the MTSU (and correspondingly the TT) for the 40 cm QZ to be the same as for the 30 cm QZ.

# Conclusion

The following observations and conclusions were made in this contribution.

Proposal 1: In order to define the most optimized MTSU and to avoid defining unnecessary increases in MTSU/TT for NTC cases, define MTSU/TT separately for ETC and NTC test cases for the 40cm QZ.

Proposal 2: Define the MTSU (and correspondingly the TT) for the 40 cm QZ to be the same as for the 30 cm QZ.

# References

- R5-223614, On QoQZ for 40cm QZ, Keysight Technologies, 3GPP TSG-RAN5 Meeting #95-e, May 2022

- R5-223613, MU discussion on 40 cm Quiet Zone, Anritsu, 3GPP TSG-RAN5 Meeting #95-e, May 2022

- R5-225206r1, Discussion on QoQZ and XPD MUs for 40cm, Rohde & Schwarz, 3GPP TSG RAN WG5 Meeting #96-e, August 2022

- TR 38.903, Derivation of test tolerances and measurement uncertainty for User Equipment (UE) conformance test cases, V16.12.0 (2022-06)

- R5-211656, On the MU for n259, Rohde & Schwarz, 3GPP TSG RAN Meeting #90-e, March 2021

- R5-211944, On the QoQZ standard deviation for ETC testing, Rohde & Schwarz, 3GPP TSG RAN Meeting #90-e, March 2021