3GPP TSG-RAN5 Meeting #96-e	R5-224612r1

# Online,  , 15th Aug 2022 - 26th Aug 2022

| CR-Form-v12.2 |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CHANGE REQUEST |  |  |  |  |  |  |  |  |
|  | 38.903 | CR | 0359 | rev | - | Current version: | 16.12.0 |  |
| For HELP on using this form: comprehensive instructions can be found at http://www.3gpp.org/Change-Requests. |  |  |  |  |  |  |  |  |

| Proposed change affects: | UICC apps |  | ME |  | Radio Access Network |  | Core Network |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |

| Title: | Definition of PC1 MU |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Source to WG: | Anritsu |  |  |  |  |  |  |  |  |  |
| Source to TSG: | R5 |  |  |  |  |  |  |  |  |  |
| Work item code: | TEI15_Test, 5GS_NR_LTE-UEConTest |  |  |  |  |  | Date: |  |  | 2022-08-08 |
| Category: | F |  |  |  |  |  | Release: |  |  | Rel-16 |
|  | Use one of the following categories: F (correction) A (mirror corresponding to a change in an earlier release) B (addition of feature), C (functional modification of feature) D (editorial modification) Detailed explanations of the above categories can be found in 3GPP TR 21.900. |  |  |  |  |  |  |  | Use one of the following releases: Rel-8 (Release 8) Rel-9 (Release 9) Rel-10 (Release 10) Rel-11 (Release 11) … Rel-16 (Release 16) Rel-17 (Release 17) Rel-18 (Release 18) Rel-19 (Release 19) |  |
| Reason for change: |  | Total MU is not defined for FR2 PC1 test cases. |  |  |  |  |  |  |  |  |
| Summary of change: |  | MU thresholds for the following test cases were defined for PC1, FR2a: - MOP (EIRP, TRP) - OFF power - Frequency error - ACLR - Tx Spurious - REFSENS - ACS - Rx Spurious Following uncertainty assessment sheets were updated or added for PC1, FR2a: - Table B.3.2-6 - Table B.8.2-6 - Table B.17.2-3 - Table B.18.2-12 to Table B.18.2-15 - Table B.19.2-3 - Table B.21.2-3 - Table B.25.2-12 to Table B.25.2-15 Relaxation values for PC1 were defined in Table B.18.2-17. [ ] for frequency range was removed from Table B.25.2-12 and Table B.25.2-13. |  |  |  |  |  |  |  |  |
| Consequences if not approved: |  | FR2 PC1 cannot be tested. |  |  |  |  |  |  |  |  |
| Clauses affected: |  | B.3, B.8, B.10, B.17, B.18, B.19, B.21, B.25 |  |  |  |  |  |  |  |  |
|  |  | Y | N |  |  |  |  |  |  |  |
| Other specs |  |  | x | Other core specifications |  |  |  | TS/TR ... CR ... |  |  |
| affected: |  | x |  | Test specifications |  |  |  | TS 38.521-2 CR R5-224611 |  |  |
| (show related CRs) |  |  | x | O&M Specifications |  |  |  | TS/TR ... CR ... |  |  |
| Other comments: |  | Depending on the outcome of MU discussion R5-224610 |  |  |  |  |  |  |  |  |
| This CR's revision history: |  | r1: B.3 and B.19 were removed to avoid overlapping with R5-224310/4311. |  |  |  |  |  |  |  |  |

<<Unchanged sections skipped>>

B.8	Transmit OFF power

Following tables summarize the MU threshold for TRP measurements for Transmit OFF power. The origin MU values for different test setups can be found in following clauses.

Table B.8-1: MU threshold for TRP measurement for Transmit OFF power

| Power Class | Frequency | MBW | Power | Threshold MU value (NOTE1) |
| --- | --- | --- | --- | --- |
| PC3 | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = Off Power | 5.67 |
|  | 32.125GHz < f <= 40.8GHz |  |  | N/A |
| PC1 | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = Off Power | 5.67 |
|  | 32.125GHz < f <= 40.8GHz |  |  | N/A |
| NOTE 1: Total TRP Expanded MU for IFF for Quiet Zone size ≤ 30cm in Table B.8.2-2 for PC3 UEs and Table B.8.2-6 for PC1 UEs. |  |  |  |  |

Table B.8-2: Void

<<Unchanged sections skipped>>

B.8.2	Uncertainty budget format and assessment for IFF

The uncertainty contributions that may impact the overall MU value are listed in Table B.8.2-1.

Table B.8.2-1: Uncertainty contributions for TRP measurement

| UID | Description of uncertainty contribution | Details in annex |
| --- | --- | --- |
| Stage 2: DUT measurement |  |  |
| 1 | Positioning misalignment | # B.2.2.1 |
| 2 | Measure distance uncertainty | B.2.2.2 |
| 3 | Quality of Quiet Zone | B.2.2.3 |
| 4 | Mismatch | B.2.2.4 |
| 5 | Standing wave between the DUT and measurement antenna | B.2.2.5 |
| 6 | Uncertainty of the RF power measurement equipment | B.2.2.6 |
| 7 | Phase curvature | B.2.2.7 |
| 8 | Amplifier uncertainties | B.2.2.8 |
| 9 | Random uncertainty | B.2.2.9 |
| 10 | Influence of the XPD | B.2.2.10 |
| 11 | Insertion Loss Variation | B.2.2.11 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | B.2.2.12 |
| 13 | Influence of TRP measurement grid | B.2.2.22 |
| 14 | Influence of beam peak search grid | B.2.2.23 |
| 15 | Multiple measurement antenna uncertainty | B.2.2.25 |
| 16 | DUT repositioning | B.2.2.26 |
| Stage 1: Calibration measurement |  |  |
| 17 | Mismatch | B.2.2.4 |
| 18 | Amplifier Uncertainties | B.2.2.8 |
| 19 | Misalignment of positioning System | B.2.2.13 |
| 20 | Uncertainty of the Network Analyzer | B.2.2.14 |
| 21 | Uncertainty of the absolute gain of the calibration antenna | B.2.2.15 |
| 22 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | B.2.2.16 |
| 23 | Phase centre offset of calibration antenna | B.2.2.18 |
| 24 | Quality of quiet zone for calibration process | B.2.2.19 |
| 25 | Standing wave between reference calibration antenna and measurement antenna | B.2.2.20 |
| 26 | Influence of the calibration antenna feed cable | B.2.2.21 |
| 27 | Insertion Loss Variation | B.2.2.11 |
| Systematic uncertainties |  |  |
| 28 | Systematic error due to TRP calculation/quadrature | B.2.2.24 |
| 29 | Influence of noise | B.2.2.27 |

The uncertainty assessment tables are organized as follows:

-	For the purpose of uncertainty assessment, the radiating antenna aperture of the DUT is denoted as D

-	The uncertainty assessment has been derived for the case of Quiet Zone size ≤ 30 cm, f = {23.45GHz, 32.125GHz, 40.8GHz}, P = Off power.

-	The uncertainty assessment for TRP is provided in Table B.8.2-2 for PC3 UEs and Table B.8.2-6 for PC1 UEs.

Table B.8.2-2: Uncertainty assessment for TRP measurement (f=23.45GHz, 32.125GHz, 40.8GHz, Quiet Zone size ≤ 30 cm) for PC3 UEs

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment | 0.00 | Normal | 2.00 | 0.00 |
| 2 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 3 | Quality of Quiet Zone (NOTE 8) | 0.6 | Actual | 1.00 | 0.6 |
| 4 | Mismatch (NOTE 1) | 1.30 | Actual | 1.00 | 1.30 |
| 5 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | Uncertainty of the RF power measurement equipment (NOTE 2) | 2.50 | Normal | 2.00 | 1.25 |
| 7 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 8 | Amplifier uncertainties | 2.10 | Normal | 2.00 | 1.05 |
| 9 | Random uncertainty | 0.50 | Normal | 2.00 | 0.25 |
| 10 | Influence of the XPD | 0.01 | U-shaped | 1.41 | 0.00 |
| 11 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 13 | Influence of TRP measurement grid (NOTE 3) | 0.25 | Actual | 1 | 0.25 |
| 14 | Influence of beam peak search grid | 0.00 | Actual | 1 | 0.00 |
| 15 | Multiple measurement antenna uncertainty (NOTE 9) | 0.15 | Actual | 1 | 0.15 |
| 16 | DUT repositioning | 0.00 | Rectangular | 1.73 | 0.00 |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 17 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 18 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 19 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 20 | Uncertainty of the Network Analyzer | 1.5 | Normal | 2.00 | 0.75 |
| 21 | Uncertainty of the absolute gain of the calibration antenna | 0.60 | Normal | 2.00 | 0.30 |
| 22 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.01 | Rectangular | 1.73 | 0.00 |
| 23 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 24 | Quality of quiet zone for calibration process (NOTE 8) | 0.4 | Actual | 1.00 | 0.4 |
| 25 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 26 | Influence of the calibration antenna feed cable | 0.14 | Normal | 2.00 | 0.07 |
| 27 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
|  | Systematic uncertainties (NOTE 5) |  |  |  | Value |
| 28 | Systematic error due to TRP calculation/quadrature (NOTE 3) |  |  |  | 0.0 |
| 29 | Influence of noise (23.45GHz <= f <= 32.125GHz) |  |  |  | 1.0 |
| 30 | Influence of noise (32.125GHz < f <= 40.8GHz) |  |  |  | N/A |
| Total measurement uncertainty |  |  |  |  | Value |
| TRP Expanded uncertainty (23.45GHz <= f <= 32.125GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 5.67 |
| TRP Expanded uncertainty (32.125GHz < f <= 40.8GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | N/A |
| NOTE 1: The analysis was done only for the case of operating at TX OFF power, in-band, non-CA. NOTE 2: The assessment assumes DUT Off power. NOTE 3: This contributor shall only be considered for TRP measurements. NOTE 4: Void NOTE 5: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 6: Void. NOTE 7: Void NOTE 8: Value based on procedure defined in Annex D.2 of TR 38.810 for Quiet Zone size less or equal to 30 cm. NOTE 9: Applies to the system which has a structure of mechanical feed antenna positioning. |  |  |  |  |  |

NOTE: MU assessment in Table B.8.2-2 for FR2a is based on the relaxation of 30.4dB for 400MHz BW.

Table B.8.2-3: Void

Table B.8.2-4: Void

Table B.8.2-5: Void

Table B.8.2-6: Uncertainty assessment for TRP measurement (f=23.45GHz, 32.125GHz, 40.8GHz, Quiet Zone size ≤ 30 cm) for PC1 UEs

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment | 0.02 | Normal | 2.00 | 0.01 |
| 2 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 3 | Quality of Quiet Zone (NOTE 8) | 0.6 | Actual | 1.00 | 0.6 |
| 4 | Mismatch (NOTE 1) | 1.30 | Actual | 1.00 | 1.30 |
| 5 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | Uncertainty of the RF power measurement equipment (NOTE 2) | 2.50 | Normal | 2.00 | 1.25 |
| 7 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 8 | Amplifier uncertainties | 2.10 | Normal | 2.00 | 1.05 |
| 9 | Random uncertainty | 0.50 | Normal | 2.00 | 0.25 |
| 10 | Influence of the XPD | 0.01 | U-shaped | 1.41 | 0.00 |
| 11 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 13 | Influence of TRP measurement grid (NOTE 3) | 0.25 | Actual | 1 | 0.25 |
| 14 | Influence of beam peak search grid | 0.00 | Actual | 1 | 0.00 |
| 15 | Multiple measurement antenna uncertainty (NOTE 9) | 0.15 | Actual | 1 | 0.15 |
| 16 | DUT repositioning | 0.00 | Rectangular | 1.73 | 0.00 |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 17 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 18 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 19 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 20 | Uncertainty of the Network Analyzer | 1.5 | Normal | 2.00 | 0.75 |
| 21 | Uncertainty of the absolute gain of the calibration antenna | 0.60 | Normal | 2.00 | 0.30 |
| 22 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.01 | Rectangular | 1.73 | 0.00 |
| 23 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 24 | Quality of quiet zone for calibration process (NOTE 8) | 0.4 | Actual | 1.00 | 0.4 |
| 25 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 26 | Influence of the calibration antenna feed cable | 0.14 | Normal | 2.00 | 0.07 |
| 27 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
|  | Systematic uncertainties (NOTE 5) |  |  |  | Value |
| 28 | Systematic error due to TRP calculation/quadrature (NOTE 3) |  |  |  | 0.0 |
| 29 | Influence of noise (23.45GHz <= f <= 32.125GHz) |  |  |  | 1.0 |
| 30 | Influence of noise (32.125GHz < f <= 40.8GHz) |  |  |  | N/A |
| Total measurement uncertainty |  |  |  |  | Value |
| TRP Expanded uncertainty (23.45GHz <= f <= 32.125GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 5.67 |
| TRP Expanded uncertainty (32.125GHz < f <= 40.8GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | N/A |
| NOTE 1: The analysis was done only for the case of operating at TX OFF power, in-band, non-CA. NOTE 2: The assessment assumes DUT Off power. NOTE 3: This contributor shall only be considered for TRP measurements. NOTE 4: Void NOTE 5: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 6: Void. NOTE 7: Void NOTE 8: Value based on procedure defined in Annex D.2 of TR 38.810 for Quiet Zone size less or equal to 30 cm. NOTE 9: Applies to the system which has a structure of mechanical feed antenna positioning. |  |  |  |  |  |

NOTE: MU assessment in Table B.8.2-2 and Table B.8.2-6 for FR2a is based on the relaxation of 30.4dB for 400MHz BW.

<<Unchanged sections skipped>>

B.10	Frequency error

Following tables summarize the MU threshold for EIRP measurements for Frequency error. The origin MU values for different test setups can be found in following subclauses.

Table B.10-1: MU threshold for beam peak measurement for Frequency error

| Power Class | Frequency | MBW | Power | Threshold MU value for NTC and ETC (NOTE1) |
| --- | --- | --- | --- | --- |
| PC3 | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = Max Output Power | +/- 0.01 ppm |
|  | 32.125GHz < f <= 40.8GHz |  |  | +/- 0.01 ppm |
| PC1 | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = Max Output Power | +/- 0.01 ppm |
|  | 32.125GHz < f <= 40.8GHz |  |  | FFS |
| NOTE 1: Total Expanded MU for IFF for Quiet Zone size ≤ 30cm in section B.10.2 |  |  |  |  |

B.10.1	Uncertainty budget format and assessment for DFF

+/- 0.01 ppm

-	The uncertainty assessment has been derived for the case of D = [5 cm], f = {23.45 GHz, 32.125 GHz, 40.8 GHz}, P = [Maximum output power]. This uncertainty has no dependency with extreme temperature conditions.

B.10.2	Uncertainty budget format and assessment for IFF

+/- 0.01 ppm

-	The uncertainty assessment has been derived for the case of Quiet zone size ≤ [30 cm], f = {23.45GHz, 32.125GHz, 40.8GHz}, P = [Maximum output power].

<<Unchanged sections skipped>>

B.17	Adjacent Channel Leakage Ratio

Editor’s Note: MU value analysis for PC2 and 4 are not complete.

Following tables summarize the MU threshold for EIRP measurements for Adjacent Channel Leakage Ratio. The origin MU values for different test setups can be found in following subclauses.

Table B.17-1: MU threshold for TRP measurement for Spectrum emission mask

| Power Class | Frequency | MBW | Power | Threshold MU value (NOTE 1) |
| --- | --- | --- | --- | --- |
| PC3 | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = Max Output Power | TBD |
|  | 32.125GHz < f <= 40.8GHz |  |  | TBD |
| PC1 | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = Max Output Power | TBD |
|  | 32.125GHz < f <= 40.8GHz |  |  | TBD |
| NOTE 1: Total Expanded MU for IFF for a Quiet Zone size ≤ 30 cm in Table B.17.2-2 for PC3 UEs and B.17.2-3 for PC1 UEs. |  |  |  |  |

Table B.17-1B: MU threshold for EIRP measurement for Adjacent Channel Leakage Ratio

| Power Class | Frequency | CBW | Power | Threshold MU value for NTC and ETC (NOTE 1) |
| --- | --- | --- | --- | --- |
| PC3 | 23.45GHz <= f <= 32.125GHz | 50MHz | P = Max Output Power | 5.63 |
|  |  | 100MHz |  | 6.09 |
|  |  | 200MHz |  | 6.09 (NOTE5) |
|  |  | 400MHz |  | 6.09 (NOTE2) |
|  | 32.125GHz < f <= 40.8GHz | 50MHz | P = Max Output Power | 6.09 (NOTE7) |
|  |  | 100MHz |  | 6.09 (NOTE6) |
|  |  | 200MHz |  | 6.09 (NOTE3) |
|  |  | 400MHz |  | 6.09 (NOTE4) |
| PC1 | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = Max Output Power | 5.59 |
|  | 32.125GHz < f <= 40.8GHz | BW <= 400MHz | P = Max Output Power | TBD |
| NOTE 1: Total Expanded MU for IFF for a Quiet Zone size ≤ 30 cm in Table B.17.2-2 for PC3 UEs and Table B.17.2-3 for PC1 UEs. NOTE 2: This value is based on the relaxation of (MPR – 3.0) dB for MPR > 3.0dB. NOTE 3: Not applicable for MPR > 3.5dB NOTE 4: Not applicable for MPR > 2.0dB NOTE 5: This value is based on the relaxation of (MPR – 5.0) dB for MPR > 5.0dB. NOTE 6: Not applicable for MPR > 5.0dB NOTE 7: Not applicable for MPR >7. 5 dB |  |  |  |  |

<<Unchanged sections skipped>>

B.17.2	Uncertainty budget format and assessment for IFF

The uncertainty contributions that may impact the overall MU value are listed in Table B.17.2-1.

Table B.17.2-1: Uncertainty contributions for EIRP measurement

| UID | Description of uncertainty contribution | Details in clause |
| --- | --- | --- |
| Stage 2: DUT measurement |  |  |
| 1 | Quality of Quiet Zone | B.2.2.3 |
| 2 | Mismatch | B.2.2.4 |
| 3 | Standing wave between the DUT and measurement antenna | B.2.2.5 |
| 4 | Uncertainty of the RF power measurement equipment | B.2.2.6 |
| 5 | Phase curvature | B.2.2.7 |
| 6 | Amplifier uncertainties | B.2.2.8 |
| 7 | Random uncertainty | B.2.2.9 |
| 8 | Influence of the XPD | B.2.2.10 |
| 9 | RF leakage (from measurement antenna to the receiver/transmitter) | B.2.2.12 |
| 10 | Multiple measurement antenna uncertainty | B.2.2.25 |
| Stage 1: Calibration measurement |  |  |
| 11 | Mismatch | B.2.2.4 |
| 12 | Amplifier Uncertainties | B.2.2.8 |
| 13 | Misalignment of positioning System | B.2.2.13 |
| 14 | Uncertainty of the Network Analyzer | B.2.2.14 |
| 15 | Uncertainty of the absolute gain of the calibration antenna | B.2.2.15 |
| 16 | Phase centre offset of calibration antenna | B.2.2.18 |
| 17 | Quality of quiet zone for calibration process | B.2.2.19 |
| 18 | Standing wave between reference calibration antenna and measurement antenna | B.2.2.20 |
| 19 | Influence of the calibration antenna feed cable | B.2.2.21 |
| 20 | Insertion Loss Variation | B.2.2.11 |
| Systematic uncertainties |  |  |
| 21 | Influence of noise | B.2.2.27 |

The uncertainty assessment tables are organized as follows:

-	For the purpose of uncertainty assessment, the radiating antenna aperture of the DUT is denoted as D

-	The uncertainty assessment has been derived for the case of Quiet Zone size ≤ 30 cm, f = {23.45GHz, 32.125GHz, 40.8GHz}, P = Maximum output power - MPR – MBR(Multi-band relaxation).

-	The uncertainty assessment for EIRP is provided in Table B.17.2-2 for PC3 UEs and Table B.17.2-3 for PC1 UEs.

Table B.17.2-2: Uncertainty assessment for EIRP measurement (f=23.45GHz, 32.125GHz, 40.8GHz, Quiet Zone size ≤ 30 cm) for PC3 UEs and normal and extreme temperature condition

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Quality of Quiet Zone (NOTE 10) | 0.52 | Actual | 1.00 | 0.52 |
| 2 | Mismatch (NOTE 2) | 1.84 | Actual | 1.00 | 1.84 |
| 3 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 4 | Uncertainty of the RF power measurement equipment (NOTE 3, 7) | 2.16 | Normal | 2.00 | 1.08 |
| 5 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | Amplifier uncertainties | 2.1 | Normal | 2.00 | 1.05 |
| 7 | Random uncertainty | 0.50 | Normal | 2.00 | 0.25 |
| 8 | Influence of the XPD | 0.00 | U-shaped | 1.41 | 0.00 |
| 9 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 10 | Multiple measurement antenna uncertainty (NOTE 9) | 0.0 | Actual | 1 | 0.0 |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 11 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 12 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 13 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 14 | Uncertainty of the Network Analyzer | 1.5 | Normal | 2.00 | 0.75 |
| 15 | Uncertainty of the absolute gain of the calibration antenna | 0.60 | Normal | 2.00 | 0.30 |
| 16 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 17 | Quality of quiet zone for calibration process (NOTE 10) | 0.32 | Actual | 1.00 | 0.32 |
| 18 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 19 | Influence of the calibration antenna feed cable | 0.00 | Normal | 2.00 | 0.00 |
| 20 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| EIRP Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 5.09 |
|  | Systematic uncertainties (NOTE 6) |  |  |  | Value |
| 21 | Influence of noise |  |  |  | Table B.17.2-4 |
| Total measurement uncertainty |  |  |  |  | Value |
| EIRP total measurement uncertainty [dB] |  |  |  |  | 5.09 + Influence of Noise |
| NOTE 1: Void NOTE 2: The analysis was done only for the case of operating at max output power – MPR – MBR(Multi-band relaxation)., in-band, non-CA. NOTE 3: The assessment assumes maximum DUT output power – MPR – MBR(Multi-band relaxation). NOTE 4: Void NOTE 5: Void NOTE 6: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 7: Void NOTE 8: Void NOTE 9: Void NOTE 10: Defined as fixed value MU contributor. |  |  |  |  |  |

Table B.17.2-3: Uncertainty assessment for EIRP measurement (f=23.45GHz, 32.125GHz, 40.8GHz, Quiet Zone size ≤ 30 cm) for PC1 UEs

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Quality of Quiet Zone (NOTE 10) | 0.52 | Actual | 1.00 | 0.52 |
| 2 | Mismatch (NOTE 2, NOTE 7) | 1.84 | Actual | 1.00 | 1.84 |
| 3 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 4 | Uncertainty of the RF power measurement equipment (NOTE 3, 7) | 2.16 | Normal | 2.00 | 1.08 |
| 5 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | Amplifier uncertainties | 2.1 | Normal | 2.00 | 1.05 |
| 7 | Random uncertainty | 0.50 | Normal | 2.00 | 0.25 |
| 8 | Influence of the XPD | 0.00 | U-shaped | 1.41 | 0.00 |
| 9 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 10 | Multiple measurement antenna uncertainty (NOTE 9) | 0.00 | Actual | 1 | 0.00 |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 11 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 12 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 13 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 14 | Uncertainty of the Network Analyzer | 1.5 | Normal | 2.00 | 0.75 |
| 15 | Uncertainty of the absolute gain of the calibration antenna | 0.60 | Normal | 2.00 | 0.30 |
| 16 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 17 | Quality of quiet zone for calibration process (NOTE 10) | 0.32 | Actual | 1.00 | 0.32 |
| 18 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 19 | Influence of the calibration antenna feed cable | 0.00 | Normal | 2.00 | 0.00 |
| 20 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| TRP Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 5.09 |
|  | Systematic uncertainties (NOTE 6) |  |  |  | Value |
| 21 | Influence of noise (23.45GHz <= f <= 32.125GHz) |  |  |  | 0.50 |
| 21 | Influence of noise (32.125GHz < f <= 40.8GHz) |  |  |  | FFS |
| Total measurement uncertainty |  |  |  |  | Value |
| TRP total measurement uncertainty [dB] (23.45GHz <= f <= 32.125GHz) |  |  |  |  | 5.59 |
| TRP total measurement uncertainty [dB] (32.125GHz < f <= 40.8GHz) |  |  |  |  | FFS |
| NOTE 1: Void NOTE 2: The analysis was done only for the case of operating at max output power, in-band, non-CA. NOTE 3: The assessment assumes maximum DUT output power. NOTE 4: Void NOTE 5: Void NOTE 6: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 7: Values extracted from TR 38.810 v2.6.1 in square brackets pending for further analysis. NOTE 8: Void. NOTE 9: Void NOTE 10: Defined as fixed value MU contributor. |  |  |  |  |  |

Table B.17.2-4: Influence of noise measurement (f=23.45GHz, 32.125GHz, 40.8GHz, Quiet Zone size ≤ 30 cm) for PC3 UEs

|  | FR2a | FR2b |
| --- | --- | --- |
| ChBW (50MHz) | 0.54 | 1.0 (NOTE 6) |
| ChBW (100MHz) | 1.0 | 1.0 (NOTE 5) |
| ChBW (200MHz) | 1.0 (NOTE 4) | 1.0 (NOTE 2) |
| ChBW (400MHz) | 1.0 (NOTE 1) | 1.0 (NOTE 3) |
| NOTE 1: This value is based on the relaxation of (MPR – 3.0) dB for MPR > 3.0dB. NOTE 2: Not applicable for MPR > 3.5dB NOTE 3: Not applicable for MPR > 2.0dB NOTE 4: This value is based on the relaxation of (MPR – 5.0) dB for MPR > 5.0dB. NOTE 5: Not applicable for MPR > 5.0dB NOTE 6: Not applicable for MPR >7. 5 dB |  |  |

B.18	Spurious emissions

Editor’s Note:

-	MU value analysis and offset value analysis for PC1, 2 and 4 are not complete.

-	MU value analysis for various test setups in clause B.18.x is not complete for above 80 GHz.

-	Offset value analysis is not complete as it is derived from MU value analysis for above 80 GHz.

Test procedure of general spurious emission comprises 2 stages: coarse TRP measurement and fine TRP measurement BW. Coarse TRP measurement is introduced to reduce the measurement time by applying sparser grids and/or wider measurement BW than fine TRP measurement while having offset dB more stringent test requirement in order not to cause additional misjudgement risk. For the frequency ranges for which coarse TRP measurement does not PASS, the measurement is continued with fine TRP measurement procedure.

Tables B.18-1,  B.18-1a, B.18-1b summarizes the MU threshold for fine TRP measurements for General spurious emissions, spurious emission band UE co-existence and additional spurious emission, respectively. The origin MU values for fine TRP measurement for different test setups can be found in following subclauses.

Table B.18-1: MU threshold for TRP measurement for general spurious emission

| Power Class | Frequency | In-band BW | In-band Power (NOTE2) | Threshold MU value [dB] (NOTE1) |
| --- | --- | --- | --- | --- |
| PC3 | 6 GHz <= f <=12.75 GHz | BW <= 400MHz | P = Max Output Power | 5.14 |
|  | 12.75 GHz <= f <= 23.45 GHz |  |  | 5.11 |
|  | 23.45 GHz <= f <= 40.8 GHz |  |  | 5.41 |
|  | 40.8 GHz <= f <= 66 GHz |  |  | 7.42 |
|  | 66 GHz <= f <= 80 GHz |  |  | 7.72 |
| PC1 | 6 GHz <= f <=12.75 GHz | BW <= 400MHz | P = Max Output Power | [5.24] |
|  | 12.75 GHz <= f <= 23.45 GHz |  |  | [6.81] |
|  | 23.45 GHz <= f <= 40.8 GHz |  |  | [7.11] |
|  | 40.8 GHz <= f <= 66 GHz |  |  | [9.09] |
|  | 66 GHz <= f <= 80 GHz |  |  | FFS |
| NOTE 1: Total EIRP Expanded MU for IFF for Quiet Zone size ≤ 30cm in Table B.18.2-3 to Table B.18.2-11 for PC3 UEs and in Table B.18.2.12 to Table B.18.2.16 for PC1 UEs. NOTE 2: Max output power level for device with corresponding power class. |  |  |  |  |

Table B.18-1a: MU threshold for TRP measurement for spurious emission band UE co-existence

| Power Class | Frequency | In-band BW | In-band Power (NOTE2) | Threshold MU value [dB] (NOTE1) |
| --- | --- | --- | --- | --- |
| PC3 | n257, n260, n261 | BW <= 400MHz | P = Max Output Power | 6.00 |
|  | 23.6 GHz < f <= 24.0 GHz |  |  | 6.00 |
|  | 36 GHz <= f <= 37 GHz |  |  | 6.00 |
|  | 57 GHz <= f <= 66 GHz |  |  | 8.01 |
| PC1 | n257, n261 | BW <= 400MHz | P = Max Output Power | FFS |
|  | n260 |  |  | FFS |
|  | 23.6 GHz < f <= 24.0 GHz |  |  | FFS |
|  | 36 GHz <= f <= 37 GHz |  |  | FFS |
|  | 57 GHz <= f <= 66 GHz |  |  | FFS |
| NOTE 1: Total EIRP Expanded MU for IFF for Quiet Zone size ≤ 30cm in Table B.18.2-3 to Table B.18.2-11 for PC3 UEs and in Table B.18.2.12 to Table B.18.2.16 for PC1 UEs. NOTE 2: Max output power level for device with corresponding power class. |  |  |  |  |

Table B.18-1b: MU threshold for TRP measurement for additional spurious emission

| Power Class | Frequency | In-band BW | In-band Power (NOTE2) | Threshold MU value [dB] (NOTE1) |
| --- | --- | --- | --- | --- |
| PC3 | 6 GHz <= f <=12.75 GHz NS_202 | BW <= 400MHz | P = Max Output Power | 5.14 |
|  | 12.75 GHz <= f <= 23.45 GHz NS_202 |  |  | 5.70 |
|  | 23.45 GHz <= f <= 40.8 GHz NS_202, NS_203 |  |  | 6.00 |
|  | 40.8 GHz <= f <= 2nd harmonic of the upper frequency edge of the UL operating band NS_202 |  |  | 8.01 |
| PC1 | 6 GHz <= f <=12.75 GHz NS_202 | BW <= 400MHz | P = Max Output Power | FFS |
|  | 12.75 GHz <= f <= 23.45 GHz NS_202 |  |  | FFS |
|  | 23.45 GHz <= f <= 40.8 GHz NS_202, NS_203 |  |  | FFS |
|  | 40.8 GHz <= f <= 2nd harmonic of the upper frequency edge of the UL operating band NS_202 |  |  | FFS |
| NOTE 1: Total EIRP Expanded MU for IFF for Quiet Zone size ≤ 30cm in Table B.18.2-3 to Table B.18.2-11 for PC3 UEs and in Table B.18.2.12 to Table B.18.2.16 for PC1 UEs. NOTE 2: Max output power level for device with corresponding power class. |  |  |  |  |

Table B.18-2 provides valid coarse TRP measurement grids and corresponding offset dB value that may be used for UE general spurious emission test case. The offset value is derived as 95%-tile TRP measurement uncertainty including the effect from uncertainty due to Coarse TRP measurement grid, excluding influence of noise.

Table B.18-2: Coarse TRP measurement grids and offset values for UE Tx spurious emission

| Power Class | Coarse TRP measurement grid | Frequency | Min Number of measurement points on the grid | Influence of coarse TRP measurement grid (dB) | Systematic error due to coarse TRP calculation/quadrature (dB) | Offset value (dB) |
| --- | --- | --- | --- | --- | --- | --- |
| PC3 | Constant density grid (charged particle based) | 6 GHz <= f <=12.75 GHz | 35 | 0.94 | 0.09 | 5.13 |
|  |  | 12.75 GHz <= f <= 23.45 GHz |  |  |  | 5.09 |
|  |  | 23.45 GHz <= f <= 40.8 GHz |  |  |  | 5.38 |
|  |  | 40.8 GHz <= f <= 66 GHz |  |  |  | 7.31 |
|  |  | 66 GHz <= f <= 80 GHz |  |  |  | 7.61 |
|  | Constant step size grid | 6 GHz <= f <=12.75 GHz | 62 | 0.97 | 0.2 | 5.26 |
|  |  | 12.75 GHz <= f <= 23.45 GHz |  |  |  | 5.23 |
|  |  | 23.45 GHz <= f <= 40.8 GHz |  |  |  | 5.52 |
|  |  | 40.8 GHz <= f <= 66 GHz |  |  |  | 7.43 |
|  |  | 66 GHz <= f <= 80 GHz |  |  |  | 7.73 |
| PC1 | Constant density grid (charged particle based) | 6 GHz <= f <=12.75 GHz | FFS | FFS | FFS | FFS |
|  |  | 12.75 GHz <= f <= 23.45 GHz |  |  |  | FFS |
|  |  | 23.45 GHz <= f <= 40.8 GHz |  |  |  | FFS |
|  |  | 40.8 GHz <= f <= 66 GHz |  |  |  | FFS |
|  |  | 66 GHz <= f <= 80 GHz |  |  |  | FFS |
|  | Constant step size grid | 6 GHz <= f <=12.75 GHz | FFS | FFS | FFS | FFS |
|  |  | 12.75 GHz <= f <= 23.45 GHz |  |  |  | FFS |
|  |  | 23.45 GHz <= f <= 40.8 GHz |  |  |  | FFS |
|  |  | 40.8 GHz <= f <= 66 GHz |  |  |  | FFS |
|  |  | 66 GHz <= f <= 80 GHz |  |  |  | FFS |
| NOTE 1: Based on Total EIRP Expanded MU for IFF for Quiet Zone size ≤ 30cm in Table B.18.2-3 to Table B.18.2-11 for PC3 UEs and in Table B.18.2.12 to Table B.18.2.16 for PC1 UEs, replacing “Influence of TRP measurement grid” and “Systematic error due to TRP calculation/quadrature” by the values for coarse TRP grid, and excluding “Influence of noise”. NOTE 2: Max output power level for device with corresponding power class. |  |  |  |  |  |  |

<<Unchanged sections skipped>>

B.18.2	Uncertainty budget format and assessment for IFF

The uncertainty contributions that may impact the overall MU value are listed in Table B.18.2-1.

Table B.18.2-1: Uncertainty contributions for TRP measurement

| UID | Description of uncertainty contribution | Details in clause |
| --- | --- | --- |
| Stage 2: DUT measurement |  |  |
| 1 | Positioning misalignment | # B.2.2.1 |
| 2 | Measure distance uncertainty | B.2.2.2 |
| 3 | Quality of Quiet Zone | B.2.2.3 |
| 4 | Mismatch | B.2.2.4 |
| 5 | Standing wave between the DUT and measurement antenna | B.2.2.5 |
| 6 | Uncertainty of the RF power measurement equipment | B.2.2.6 |
| 7 | Phase curvature | B.2.2.7 |
| 8 | Amplifier uncertainties | B.2.2.8 |
| 9 | Random uncertainty | B.2.2.9 |
| 10 | Influence of the XPD | B.2.2.10 |
| 11 | Insertion Loss Variation | B.2.2.11 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | B.2.2.12 |
| 13 | Influence of TRP measurement grid | B.2.2.22 |
| 14 | Influence of beam peak search grid | B.2.2.23 |
| 15 | Multiple measurement antenna uncertainty | B.2.2.25 |
| 16 | DUT repositioning | B.2.2.26 |
| 17 | Misalignment of DUT due to change of DUT orientation | B.2.2.31 |
| Stage 1: Calibration measurement |  |  |
| 18 | Mismatch | B.2.2.4 |
| 19 | Amplifier Uncertainties | B.2.2.8 |
| 20 | Misalignment of positioning System | B.2.2.13 |
| 21 | Uncertainty of the Network Analyzer | B.2.2.14 |
| 22 | Uncertainty of the absolute gain of the calibration antenna | B.2.2.15 |
| 23 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | B.2.2.16 |
| 24 | Phase centre offset of calibration antenna | B.2.2.18 |
| 25 | Quality of quiet zone for calibration process | B.2.2.19 |
| 26 | Standing wave between reference calibration antenna and measurement antenna | B.2.2.20 |
| 27 | Influence of the calibration antenna feed cable | B.2.2.21 |
| 28 | Insertion Loss Variation | B.2.2.11 |
| Systematic uncertainties |  |  |
| 29 | Systematic error due to TRP calculation/quadrature | B.2.2.24 |
| 30 | Influence of noise | B.2.2.27 |

The uncertainty assessment tables are organized as follows:

-	For the purpose of uncertainty assessment, the radiating antenna aperture of the DUT is denoted as D

-	The uncertainty assessment has been derived for the case of Quiet zone size ≤ 30 cm, f = {6 GHz to 80 GHz}, P = Maximum output power.

-	The uncertainty assessment for TRP is provided from Table B.18.2-2 to Table B.18.2-11 for PC3 UEs and from Table B.18.2.12 to Table B.18.2.16 for PC1 UEs.

Table B.18.2-2: Void

Table B.18.2-3: Uncertainty assessment for TRP measurement (f=6 GHz to 12.75GHz, Quiet Zone size ≤ 30 cm) for PC3 UEs

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment | 0.00 | Normal | 2.00 | 0.00 |
| 2 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 3 | Quality of Quiet Zone (NOTE 4) | 0.70 | Actual | 1.00 | 0.70 |
| 4 | Mismatch | 1.50 | Actual | 1.00 | 1.50 |
| 5 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | Uncertainty of the RF power measurement equipment | 2.00 | Normal | 2.00 | 1.00 |
| 7 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 8 | Amplifier uncertainties | 2.1 | Normal | 2.00 | 1.05 |
| 9 | Random uncertainty | 0.5 | Normal | 2.00 | 0.25 |
| 10 | Influence of the XPD | 0.09 | U-shaped | 1.41 | 0.064 |
| 11 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 13 | Influence of TRP measurement grid (NOTE 1) | 0.32 | Actual | 1 | 0.32 |
| 14 | Influence of beam peak search grid (NOTE 2) | N/A | Actual | 1 | N/A |
| 15 | Multiple measurement antenna uncertainty (NOTE 5) | 0.15 | Actual | 1 | 0.15 |
| 16 | DUT repositioning | 0.00 | Rectangular | 1.73 | 0.00 |
| 17 | Misalignment of DUT due to change of DUT orientation | 0.10 | Actual | 1 | 0.10 |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 18 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 19 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 20 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 21 | Uncertainty of the Network Analyzer | 0.90 | Normal | 2.00 | 0.45 |
| 22 | Uncertainty of the absolute gain of the calibration antenna | 0.60 | Normal | 2.00 | 0.30 |
| 23 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.05 | Rectangular | 1.73 | 0.03 |
| 24 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 25 | Quality of quiet zone for calibration process (NOTE 4) | 0.70 | Actual | 1.00 | 0.70 |
| 26 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 27 | Influence of the calibration antenna feed cable | 0.14 | Normal | 2.00 | 0.07 |
| 28 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
|  | Expanded uncertainty (1.96σ - confidence interval of 95 %) |  |  |  | Value |
|  | TRP Expanded uncertainty (6 GHz < f <= 12.75 GHz) [dB] (a) |  |  |  | 4.73 |
|  | Systematic uncertainties (NOTE 3) |  |  |  | Value |
| 29 | Systematic error due to TRP calculation/quadrature (NOTE 1) (b) |  |  |  | 0.0 |
| 30 | General spurious emissions Influence of noise (c1) (6 GHz < f <= 12.75 GHz) |  |  |  | 0.41 |
| 31 | Additional spurious emissions Influence of noise (c2) NS_202 (6 GHz < f <= 12.75 GHz) |  |  |  | 0.41 |
| 32 | Systematic error related to beam peak search (NOTE 2) |  |  |  | N/A |
| Total measurement uncertainty |  |  |  |  | Value |
| General spurious emissions Total measurement uncertainty (a)+(b)+(c1) [dB] NS_202 (6 GHz < f <= 12.75 GHz) |  |  |  |  | 5.14 |
| Additional spurious emissions Total measurement uncertainty (a)+(b)+(c2) [dB] NS_202 (6 GHz < f <= 12.75 GHz) |  |  |  |  | 5.14 |
| NOTE 1: This contributor shall only be considered for TRP measurements. NOTE 2: This contributor shall only be considered for EIRP measurements. NOTE 3: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 4: Value based on procedure defined in clause D.2 of TR 38.810 for Quiet Zone size of less or equal to 30 cm. NOTE 5: Applies to the system which has a structure of mechanical feed antenna positioning. |  |  |  |  |  |

Table B.18.2-4: Void

Table B.18.2-5: Uncertainty assessment for TRP measurement (f=12.75 GHz to 23.45 GHz, Quiet Zone size ≤ 30 cm) for PC3 UEs

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment | 0.00 | Normal | 2.00 | 0.00 |
| 2 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 3 | Quality of Quiet Zone (NOTE 4) | 0.60 | Actual | 1.00 | 0.60 |
| 4 | Mismatch | 1.50 | Actual | 1.00 | 1.50 |
| 5 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | Uncertainty of the RF power measurement equipment | 2.16 | Normal | 2.00 | 1.08 |
| 7 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 8 | Amplifier uncertainties | 2.1 | Normal | 2.00 | 1.05 |
| 9 | Random uncertainty | 0.5 | Normal | 2.00 | 0.25 |
| 10 | Influence of the XPD | 0.09 | U-shaped | 1.41 | 0.064 |
| 11 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 13 | Influence of TRP measurement grid (NOTE 1) | 0.32 | Actual | 1 | 0.32 |
| 14 | Influence of beam peak search grid (NOTE 2) | N/A | Actual | 1 | N/A |
| 15 | Multiple measurement antenna uncertainty (NOTE 5) | 0.15 | Actual | 1 | 0.15 |
| 16 | DUT repositioning | 0.00 | Rectangular | 1.73 | 0.00 |
| 17 | Misalignment of DUT due to change of DUT orientation | 0.10 | Actual | 1 | 0.10 |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 18 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 19 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 20 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 21 | Uncertainty of the Network Analyzer | 0.90 | Normal | 2.00 | 0.45 |
| 22 | Uncertainty of the absolute gain of the calibration antenna | 0.60 | Normal | 2.00 | 0.30 |
| 23 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.05 | Rectangular | 1.73 | 0.03 |
| 24 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 25 | Quality of quiet zone for calibration process (NOTE 4) | 0.60 | Actual | 1.00 | 0.60 |
| 26 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 27 | Influence of the calibration antenna feed cable | 0.14 | Normal | 2.00 | 0.07 |
| 28 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
|  | Expanded uncertainty (1.96σ - confidence interval of 95 %) |  |  |  | Value |
|  | TRP Expanded uncertainty (12.75 GHz < f <= 23.45 GHz) [dB] (a) |  |  |  | 4.70 |
|  | Systematic uncertainties (NOTE 3) |  |  |  | Value |
| 29 | Systematic error due to TRP calculation/quadrature (NOTE 1) (b) |  |  |  | 0.0 |
| 30 | General spurious emissions Influence of noise (c1) (12.75 GHz < f <= 23.45 GHz) |  |  |  | 0.41 |
| 31 | Additional spurious emissions Influence of noise (c2) NS_202 (12.75 GHz < f <= 23.45 GHz) |  |  |  | 1.0 |
| 32 | Systematic error related to beam peak search (NOTE 2) |  |  |  | N/A |
| Total measurement uncertainty |  |  |  |  | Value |
| General spurious emissions Total measurement uncertainty (a)+(b)+(c1) [dB] (12.75 GHz < f <= 23.45 GHz) |  |  |  |  | 5.11 |
| Additional spurious emissions Total measurement uncertainty (a)+(b)+(c2) [dB] NS_202 (12.75 GHz < f <= 23.45 GHz) |  |  |  |  | 5.70 |
| NOTE 1: This contributor shall only be considered for TRP measurements. NOTE 2: This contributor shall only be considered for EIRP measurements. NOTE 3: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 4: Value based on procedure defined in clause D.2 of TR 38.810 for Quiet Zone size of less or equal to 30 cm. NOTE 5: Applies to the system which has a structure of mechanical feed antenna positioning. |  |  |  |  |  |

Table B.18.2-6: Void

Table B.18.2-7: Uncertainty assessment for TRP measurement (f=23.45 GHz to 40.8 GHz, Quiet Zone size ≤ 30 cm) for PC3 UEs

| UID |  | Uncertainty source |  | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |
| 1 |  | Positioning misalignment |  | 0.00 | Normal | 2.00 | 0.00 |  |
| 2 |  | Measure distance uncertainty |  | 0.00 | Rectangular | 1.73 | 0.00 |  |
| 3 |  | Quality of Quiet Zone (NOTE 4) |  | 0.6 | Actual | 1.00 | 0.6 |  |
| 4 |  | Mismatch |  | 1.40 | Actual | 1.00 | 1.40 |  |
| 5 |  | Standing wave between the DUT and measurement antenna |  | 0.00 | U-shaped | 1.41 | 0.00 |  |
| 6 |  | Uncertainty of the RF power measurement equipment |  | 2.73 | Normal | 2.00 | 1.37 |  |
| 7 |  | Phase curvature |  | 0.00 | U-shaped | 1.41 | 0.00 |  |
| 8 |  | Amplifier uncertainties |  | 2.1 | Normal | 2.00 | 1.05 |  |
| 9 |  | Random uncertainty |  | 0.5 | Normal | 2.00 | 0.25 |  |
| 10 |  | Influence of the XPD |  | 0.01 | U-shaped | 1.41 | 0.00 |  |
| 11 |  | Insertion Loss Variation |  | 0.00 | Rectangular | 1.73 | 0.00 |  |
| 12 |  | RF leakage (from measurement antenna to the receiver/transmitter) |  | 0.00 | Actual | 1.00 | 0.00 |  |
| 13 |  | Influence of TRP measurement grid (NOTE 1) |  | 0.32 | Actual | 1 | 0.32 |  |
| 14 |  | Influence of beam peak search grid (NOTE 2) |  | N/A | Actual | 1 | N/A |  |
| 15 |  | Multiple measurement antenna uncertainty (NOTE 5) |  | 0.15 | Actual | 1 | 0.15 |  |
| 16 |  | DUT repositioning |  | 0.00 | Rectangular | 1.73 | 0.00 |  |
| 17 |  | Misalignment of DUT due to change of DUT orientation |  | 0.10 | Actual | 1 | 0.10 |  |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |  |
| 18 |  | Mismatch |  | 0.00 | U-shaped | 1.41 | 0.00 |  |
| 19 |  | Amplifier Uncertainties |  | 0.00 | Normal | 2.00 | 0.00 |  |
| 20 |  | Misalignment of positioning System |  | 0.00 | Normal | 2.00 | 0.00 |  |
| 21 |  | Uncertainty of the Network Analyzer |  | 1.5 | Normal | 2.00 | 0.75 |  |
| 22 |  | Uncertainty of the absolute gain of the calibration antenna |  | 0.6 | Normal | 2.00 | 0.3 |  |
| 23 |  | Positioning and pointing misalignment between the reference antenna and the measurement antenna |  | 0.05 | Rectangular | 1.73 | 0.03 |  |
| 24 |  | Phase centre offset of calibration antenna |  | 0.00 | Rectangular | 1.73 | 0.00 |  |
| 25 |  | Quality of quiet zone for calibration process (NOTE 4) |  | 0.6 | Actual | 1.00 | 0.6 |  |
| 26 |  | Standing wave between reference calibration antenna and measurement antenna |  | 0.00 | U-shaped | 1.41 | 0.00 |  |
| 27 |  | Influence of the calibration antenna feed cable |  | 0.14 | Normal | 2.00 | 0.07 |  |
| 28 |  | Insertion Loss Variation |  | 0.00 | Rectangular | 1.73 | 0.00 |  |
|  |  | Expanded uncertainty (1.96σ - confidence interval of 95 %) |  |  |  |  | Value |  |
|  |  | TRP Expanded uncertainty (23.45 GHz < f <= 40.8 GHz) [dB] (a) |  |  |  |  | 5.00 |  |
|  |  | Systematic uncertainties (NOTE 3) |  |  |  |  | Value |  |
| 29 |  | Systematic error due to TRP calculation/quadrature (NOTE 1) (b) |  |  |  |  | 0.0 |  |
| 30 |  | General spurious emissions Influence of noise (c1) (23.45 GHz < f <= 40.8 GHz) |  |  |  |  | 0.41 |  |
| 31 |  | Spurious emission band UE co-existence Influence of noise (c2) (f within NR Bands n257, n260 or n261) |  |  |  |  | 1.0 |  |
| 32 |  | Spurious emission band UE co-existence Influence of noise (c3) (36 GHz <= f <= 37 GHz) |  |  |  |  | 1.0 |  |
| 33 |  | Additional spurious emissions Influence of noise (c4) NS_202 (23.6 GHz < f <= 24.0 GHz) |  |  |  |  | 1.0 |  |
| 34 |  | Additional spurious emissions Influence of noise (c5) NS_202 (23.45 GHz < f <= 40.8 GHz) |  |  |  |  | 1.0 |  |
| 35 |  | Additional spurious emissions Influence of noise (c6) NS_203 (23.6 GHz < f <= 24.0 GHz) |  |  |  |  | 1.0 |  |
| 36 |  | Spurious emission band UE co-existence Influence of noise (c7) (23.6 GHz < f <= 24.0 GHz) |  |  |  |  | 1.0 |  |
| 37 |  | Systematic error related to beam peak search (NOTE 2) |  |  |  |  | N/A |  |
| Total measurement uncertainty |  |  |  |  |  |  | Value |  |
| General spurious emissions Total measurement uncertainty (a)+(b)+(c1) [dB] (23.45 GHz < f <= 40.8 GHz) |  |  |  |  |  |  | 5.41 |  |
| Spurious emission band UE co-existence Total measurement uncertainty (a)+(b)+(c2) [dB] (f within NR Bands n257, n260 or n261) |  |  |  |  |  |  | 6.00 |  |
| Spurious emission band UE co-existence Total measurement uncertainty (a)+(b)+(c3) [dB] (36 GHz <= f <= 37 GHz) |  |  |  |  |  |  | 6.00 |  |
| Additional spurious emissions Total measurement uncertainty (a)+(b)+(c4) [dB] NS_202 (23.6 GHz < f <= 24.0 GHz) |  |  |  |  |  |  | 6.00 |  |
| Additional spurious emissions Total measurement uncertainty (a)+(b)+(c5) [dB] NS_202 (23.45 GHz < f <= 40.8 GHz) |  |  |  |  |  |  | 6.00 |  |
| Additional spurious emissions Total measurement uncertainty (a)+(b)+(c6) [dB] NS_203 (23.6 GHz < f <= 24.0 GHz) |  |  |  |  |  |  | 6.00 |  |
| Additional spurious emissions Total measurement uncertainty (a)+(b)+(c7) [dB] (23.6 GHz < f <= 24.0 GHz) |  |  |  |  |  |  | 6.00 |  |
| NOTE 1: This contributor shall only be considered for TRP measurements. NOTE 2: This contributor shall only be considered for EIRP measurements. NOTE 3: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 4: Value based on procedure defined in clause D.2 of TR 38.810 for Quiet Zone size of less or equal to 30 cm. NOTE 5: Applies to the system which has a structure of mechanical feed antenna positioning. |  |  |  |  |  |  |  |  |

Table B.18.2-8: Void

Table B.18.2-9: Uncertainty assessment for TRP measurement (f= 40.8 GHz to 66 GHz, Quiet Zone size ≤ 30 cm) for PC3 UEs

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment | 0.0 | Normal | 2.00 | 0.0 |
| 2 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 3 | Quality of Quiet Zone (NOTE 4) | 0.6 | Actual | 1.00 | 0.6 |
| 4 | Mismatch | 2.30 | Actual | 1.00 | 2.30 |
| 5 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | Uncertainty of the RF power measurement equipment | 4.0 | Normal | 2.00 | 2.00 |
| 7 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 8 | Amplifier uncertainties | 2.1 | Normal | 2.00 | 1.05 |
| 9 | Random uncertainty | 0.5 | Normal | 2.00 | 0.25 |
| 10 | Influence of the XPD | 0.09 | U-shaped | 1.41 | 0.064 |
| 11 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 13 | Influence of TRP measurement grid (NOTE 1) | 0.32 | Actual | 1 | 0.32 |
| 14 | Influence of beam peak search grid (NOTE 2) | N/A | Actual | 1 | N/A |
| 15 | Multiple measurement antenna uncertainty (NOTE 5) | 0.15 | Actual | 1 | 0.15 |
| 16 | DUT repositioning | 0.00 | Rectangular | 1.73 | 0.00 |
| 17 | Misalignment of DUT due to change of DUT orientation | 0.10 | Actual | 1 | 0.10 |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 18 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 19 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 20 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 21 | Uncertainty of the Network Analyzer | 1.7 | Normal | 2.00 | 0.85 |
| 22 | Uncertainty of the absolute gain of the calibration antenna | 1.70 | Normal | 2.00 | 0.85 |
| 23 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.05 | Rectangular | 1.73 | 0.03 |
| 24 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 25 | Quality of quiet zone for calibration process (NOTE 4) | 0.6 | Actual | 1.00 | 0.6 |
| 26 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 27 | Influence of the calibration antenna feed cable | 0.28 | Normal | 2.00 | 0.14 |
| 28 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
|  | Expanded uncertainty (1.96σ - confidence interval of 95 %) |  |  |  | Value |
|  | TRP Expanded uncertainty ( 40.8 GHz < f <= 66 GHz) [dB] (a) |  |  |  | 7.01 |
|  | Systematic uncertainties (NOTE 3) |  |  |  | Value |
| 29 | Systematic error due to TRP calculation/quadrature (NOTE 1) (b) |  |  |  | 0.00 |
| 30 | General spurious emissions Influence of noise (c1) (40.8 GHz < f <= 66 GHz) |  |  |  | 0.41 |
| 31 | Spurious emission band UE co-existence Influence of noise (c2) (57 GHz <= f <= 66 GHz) |  |  |  | 1.0 |
| 32 | Additional spurious emissions Influence of noise (c3) NS_202 (40.8 GHz < f <= 2nd harmonic of the upper frequency edge of the UL operating band) |  |  |  | 1.0 |
| 33 | Systematic error related to beam peak search (NOTE 2) |  |  |  | N/A |
| Total measurement uncertainty |  |  |  |  | Value |
| General spurious emissions Total measurement uncertainty (a)+(b)+(c1) [dB] (40.8 GHz < f <= 66 GHz) |  |  |  |  | 7.42 |
| Spurious emission band UE co-existence Total measurement uncertainty (a)+(b)+(c2) [dB] (57 GHz <= f <= 66 GHz) |  |  |  |  | 8.01 |
| Additional spurious emissions Total measurement uncertainty (a)+(b)+(c3) [dB] NS_202 (40.8 GHz < f <= 2nd harmonic of the upper frequency edge of the UL operating band) |  |  |  |  | 8.01 |
| NOTE 1: This contributor shall only be considered for TRP measurements. NOTE 2: This contributor shall only be considered for EIRP measurements. NOTE 3: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 4: Value based on procedure defined in clause D.2 of TR 38.810 for Quiet Zone size of less or equal to 30 cm. NOTE 5: Applies to the system which has a structure of mechanical feed antenna positioning. NOTE 6: Void |  |  |  |  |  |

Table B.18.2-10: Void

Table B.18.2-11: Uncertainty assessment for TRP measurement (f= 66 GHz to 80 GHz, Quiet Zone size ≤ 30 cm) for PC3 UEs

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment | 0.00 | Normal | 2.00 | 0.00 |
| 2 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 3 | Quality of Quiet Zone (NOTE 4) | 0.6 | Actual | 1.00 | 0.6 |
| 4 | Mismatch | 2.30 | Actual | 1.00 | 2.30 |
| 5 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | Uncertainty of the RF power measurement equipment | 4.00 | Normal | 2.00 | 2.00 |
| 7 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 8 | Amplifier uncertainties | 3.0 | Normal | 2.00 | 1.50 |
| 9 | Random uncertainty | 0.5 | Normal | 2.00 | 0.25 |
| 10 | Influence of the XPD | 0.09 | U-shaped | 1.41 | 0.064 |
| 11 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 13 | Influence of TRP measurement grid (NOTE 1) | 0.32 | Actual | 1 | 0.32 |
| 14 | Influence of beam peak search grid (NOTE 2) | N/A | Actual | 1 | N/A |
| 15 | Multiple measurement antenna uncertainty (NOTE 5) | 0.15 | Actual | 1 | 0.15 |
| 16 | DUT repositioning | 0.00 | Rectangular | 1.73 | 0.00 |
| 17 | Misalignment of DUT due to change of DUT orientation | 0.10 | Actual | 1 | 0.10 |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 18 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 19 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 20 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 21 | Uncertainty of the Network Analyzer | 1.70 | Normal | 2.00 | 0.85 |
| 22 | Uncertainty of the absolute gain of the calibration antenna | 1.70 | Normal | 2.00 | 0.85 |
| 23 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.05 | Rectangular | 1.73 | 0.03 |
| 24 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 25 | Quality of quiet zone for calibration process (NOTE 4) | 0.60 | Actual | 1.00 | 0.60 |
| 26 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 27 | Influence of the calibration antenna feed cable | 0.28 | Normal | 2.00 | 0.14 |
| 28 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
|  | Expanded uncertainty (1.96σ - confidence interval of 95 %) |  |  |  | Value |
|  | TRP Expanded uncertainty (66 GHz < f <= 80 GHz) [dB] (a) |  |  |  | 7.31 |
|  | Systematic uncertainties (NOTE 3) |  |  |  | Value |
| 29 | Systematic error due to TRP calculation/quadrature (NOTE 1) (b) |  |  |  | 0.00 |
| 30 | General spurious emissions Influence of noise (c1) (66 GHz < f <= 80 GHz) |  |  |  | 0.41 |
| 31 | Systematic error related to beam peak search (NOTE 2) |  |  |  | N/A |
| Total measurement uncertainty |  |  |  |  | Value |
| General spurious emissions Total measurement uncertainty (a)+(b)+(c1) [dB] (66 GHz < f <= 80 GHz) |  |  |  |  | 7.72 |
| NOTE 1: This contributor shall only be considered for TRP measurements. NOTE 2: This contributor shall only be considered for EIRP measurements. NOTE 3: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 4: Value based on procedure defined in clause D.2 of TR 38.810 for Quiet Zone size of less or equal to 30 cm. NOTE 5: Applies to the system which has a structure of mechanical feed antenna positioning. |  |  |  |  |  |

Table B.18.2-12: Uncertainty assessment for TRP measurement (f=6 GHz to 12.75GHz, Quiet Zone size ≤ 30 cm) for PC1 UEs

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment | 0.02 | Normal | 2.00 | 0.01 |
| 2 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 3 | Quality of Quiet Zone (NOTE 4) | 0.70 | Actual | 1.00 | 0.70 |
| 4 | Mismatch | 1.50 | Actual | 1.00 | 1.50 |
| 5 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | Uncertainty of the RF power measurement equipment | 2.00 | Normal | 2.00 | 1.00 |
| 7 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 8 | Amplifier uncertainties | 2.1 | Normal | 2.00 | 1.05 |
| 9 | Random uncertainty | 0.5 | Normal | 2.00 | 0.25 |
| 10 | Influence of the XPD | 0.09 | U-shaped | 1.41 | 0.064 |
| 11 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 13 | Influence of TRP measurement grid (NOTE 1) | 0.32 | Actual | 1 | 0.32 |
| 14 | Influence of beam peak search grid (NOTE 2) | N/A | Actual | 1 | N/A |
| 15 | Multiple measurement antenna uncertainty (NOTE 5) | 0.15 | Actual | 1 | 0.15 |
| 16 | DUT repositioning | 0.00 | Rectangular | 1.73 | 0.00 |
| 17 | Misalignment of DUT due to change of DUT orientation | [0.50] | Actual | 1 | [0.50] |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 18 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 19 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 20 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 21 | Uncertainty of the Network Analyzer | 0.90 | Normal | 2.00 | 0.45 |
| 22 | Uncertainty of the absolute gain of the calibration antenna | 0.60 | Normal | 2.00 | 0.30 |
| 23 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.05 | Rectangular | 1.73 | 0.03 |
| 24 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 25 | Quality of quiet zone for calibration process (NOTE 4) | 0.70 | Actual | 1.00 | 0.70 |
| 26 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 27 | Influence of the calibration antenna feed cable | 0.14 | Normal | 2.00 | 0.07 |
| 28 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
|  | Expanded uncertainty (1.96σ - confidence interval of 95 %) |  |  |  | Value |
|  | TRP Expanded uncertainty (6 GHz < f <= 12.75 GHz) [dB] (a) |  |  |  | [4.83] |
|  | Systematic uncertainties (NOTE 3) |  |  |  | Value |
| 29 | Systematic error due to TRP calculation/quadrature (NOTE 1) (b) |  |  |  | 0.0 |
| 30 | General spurious emissions Influence of noise (c1) (6 GHz < f <= 12.75 GHz) |  |  |  | 0.41 |
| 31 | Additional spurious emissions Influence of noise (c2) NS_202 (6 GHz < f <= 12.75 GHz) |  |  |  | FFS |
| 32 | Systematic error related to beam peak search (NOTE 2) |  |  |  | N/A |
| Total measurement uncertainty |  |  |  |  | Value |
| General spurious emissions Total measurement uncertainty (a)+(b)+(c1) [dB] (6 GHz < f <= 12.75 GHz) |  |  |  |  | [5.24] |
| Additional spurious emissions Total measurement uncertainty (a)+(b)+(c2) [dB] NS_202 (6 GHz < f <= 12.75 GHz) |  |  |  |  | FFS |
| NOTE 1: This contributor shall only be considered for TRP measurements. NOTE 2: This contributor shall only be considered for EIRP measurements. NOTE 3: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 4: Value based on procedure defined in clause D.2 of TR 38.810 for Quiet Zone size of less or equal to 30 cm. NOTE 5: Applies to the system which has a structure of mechanical feed antenna positioning. |  |  |  |  |  |

Table B.18.2-13: Uncertainty assessment for TRP measurement (f=12.75 GHz to 23.45 GHz, Quiet Zone size ≤ 30 cm) for PC1 UEs

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment | 0.02 | Normal | 2.00 | 0.01 |
| 2 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 3 | Quality of Quiet Zone (NOTE 4) | 0.60 | Actual | 1.00 | 0.60 |
| 4 | Mismatch | 1.50 | Actual | 1.00 | 1.50 |
| 5 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | Uncertainty of the RF power measurement equipment | 2.16 | Normal | 2.00 | 1.08 |
| 7 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 8 | Amplifier uncertainties | 2.1 | Normal | 2.00 | 1.05 |
| 9 | Random uncertainty | 0.5 | Normal | 2.00 | 0.25 |
| 10 | Influence of the XPD | 0.09 | U-shaped | 1.41 | 0.064 |
| 11 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 13 | Influence of TRP measurement grid (NOTE 1) | 0.32 | Actual | 1 | 0.32 |
| 14 | Influence of beam peak search grid (NOTE 2) | N/A | Actual | 1 | N/A |
| 15 | Multiple measurement antenna uncertainty (NOTE 5) | 0.15 | Actual | 1 | 0.15 |
| 16 | DUT repositioning | 0.00 | Rectangular | 1.73 | 0.00 |
| 17 | Misalignment of DUT due to change of DUT orientation | [0.50] | Actual | 1 | [0.50] |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 18 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 19 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 20 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 21 | Uncertainty of the Network Analyzer | 0.90 | Normal | 2.00 | 0.45 |
| 22 | Uncertainty of the absolute gain of the calibration antenna | 0.60 | Normal | 2.00 | 0.30 |
| 23 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.05 | Rectangular | 1.73 | 0.03 |
| 24 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 25 | Quality of quiet zone for calibration process (NOTE 4) | 0.60 | Actual | 1.00 | 0.60 |
| 26 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 27 | Influence of the calibration antenna feed cable | 0.14 | Normal | 2.00 | 0.07 |
| 28 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
|  | Expanded uncertainty (1.96σ - confidence interval of 95 %) |  |  |  | Value |
|  | TRP Expanded uncertainty (12.75 GHz < f <= 23.45 GHz) [dB] (a) |  |  |  | [4.80] |
|  | Systematic uncertainties (NOTE 3) |  |  |  | Value |
| 29 | Systematic error due to TRP calculation/quadrature (NOTE 1) (b) |  |  |  | 0.0 |
| 30 | General spurious emissions Influence of noise (c1) (12.75 GHz < f <= 23.45 GHz) |  |  |  | 2.02 |
| 31 | Additional spurious emissions Influence of noise (c2) NS_202 (12.75 GHz < f <= 23.45 GHz) |  |  |  | FFS |
| 32 | Systematic error related to beam peak search (NOTE 2) |  |  |  | N/A |
| Total measurement uncertainty |  |  |  |  | Value |
| General spurious emissions Total measurement uncertainty (a)+(b)+(c) [dB] (12.75 GHz < f <= 23.45 GHz) |  |  |  |  | [6.81] |
| Additional spurious emissions Total measurement uncertainty (a)+(b)+(c2) [dB] NS_202 (12.75 GHz < f <= 23.45 GHz) |  |  |  |  | FFS |
| NOTE 1: This contributor shall only be considered for TRP measurements. NOTE 2: This contributor shall only be considered for EIRP measurements. NOTE 3: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 4: Value based on procedure defined in clause D.2 of TR 38.810 for Quiet Zone size of less or equal to 30 cm. NOTE 5: Applies to the system which has a structure of mechanical feed antenna positioning. |  |  |  |  |  |

Table B.18.2-14: Uncertainty assessment for TRP measurement (f=23.45 GHz to 40.8 GHz, Quiet Zone size ≤ 30 cm) for PC1 UEs

| UID |  | Uncertainty source |  | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |
| 1 |  | Positioning misalignment |  | 0.02 | Normal | 2.00 | 0.01 |  |
| 2 |  | Measure distance uncertainty |  | 0.00 | Rectangular | 1.73 | 0.00 |  |
| 3 |  | Quality of Quiet Zone (NOTE 4) |  | 0.6 | Actual | 1.00 | 0.6 |  |
| 4 |  | Mismatch |  | 1.40 | Actual | 1.00 | 1.40 |  |
| 5 |  | Standing wave between the DUT and measurement antenna |  | 0.00 | U-shaped | 1.41 | 0.00 |  |
| 6 |  | Uncertainty of the RF power measurement equipment |  | 2.73 | Normal | 2.00 | 1.37 |  |
| 7 |  | Phase curvature |  | 0.00 | U-shaped | 1.41 | 0.00 |  |
| 8 |  | Amplifier uncertainties |  | 2.1 | Normal | 2.00 | 1.05 |  |
| 9 |  | Random uncertainty |  | 0.5 | Normal | 2.00 | 0.25 |  |
| 10 |  | Influence of the XPD |  | 0.01 | U-shaped | 1.41 | 0.00 |  |
| 11 |  | Insertion Loss Variation |  | 0.00 | Rectangular | 1.73 | 0.00 |  |
| 12 |  | RF leakage (from measurement antenna to the receiver/transmitter) |  | 0.00 | Actual | 1.00 | 0.00 |  |
| 13 |  | Influence of TRP measurement grid (NOTE 1) |  | 0.32 | Actual | 1 | 0.32 |  |
| 14 |  | Influence of beam peak search grid (NOTE 2) |  | N/A | Actual | 1 | N/A |  |
| 15 |  | Multiple measurement antenna uncertainty (NOTE 5) |  | 0.15 | Actual | 1 | 0.15 |  |
| 16 |  | DUT repositioning |  | 0.00 | Rectangular | 1.73 | 0.00 |  |
| 17 |  | Misalignment of DUT due to change of DUT orientation |  | [0.50] | Actual | 1 | [0.50] |  |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |  |
| 18 |  | Mismatch |  | 0.00 | U-shaped | 1.41 | 0.00 |  |
| 19 |  | Amplifier Uncertainties |  | 0.00 | Normal | 2.00 | 0.00 |  |
| 20 |  | Misalignment of positioning System |  | 0.00 | Normal | 2.00 | 0.00 |  |
| 21 |  | Uncertainty of the Network Analyzer |  | 1.5 | Normal | 2.00 | 0.75 |  |
| 22 |  | Uncertainty of the absolute gain of the calibration antenna |  | 0.6 | Normal | 2.00 | 0.3 |  |
| 23 |  | Positioning and pointing misalignment between the reference antenna and the measurement antenna |  | 0.05 | Rectangular | 1.73 | 0.03 |  |
| 24 |  | Phase centre offset of calibration antenna |  | 0.00 | Rectangular | 1.73 | 0.00 |  |
| 25 |  | Quality of quiet zone for calibration process (NOTE 4) |  | 0.6 | Actual | 1.00 | 0.6 |  |
| 26 |  | Standing wave between reference calibration antenna and measurement antenna |  | 0.00 | U-shaped | 1.41 | 0.00 |  |
| 27 |  | Influence of the calibration antenna feed cable |  | 0.14 | Normal | 2.00 | 0.07 |  |
| 28 |  | Insertion Loss Variation |  | 0.00 | Rectangular | 1.73 | 0.00 |  |
|  |  | Expanded uncertainty (1.96σ - confidence interval of 95 %) |  |  |  |  | Value |  |
|  |  | TRP Expanded uncertainty (23.45 GHz < f <= 40.8 GHz) [dB] (a) |  |  |  |  | [5.09] |  |
|  |  | Systematic uncertainties (NOTE 3) |  |  |  |  | Value |  |
| 29 |  | Systematic error due to TRP calculation/quadrature (NOTE 1) (b) |  |  |  |  | 0.0 |  |
| 30 |  | General spurious emissions Influence of noise (c1) (23.45 GHz < f <= 40.8 GHz) |  |  |  |  | 2.02 |  |
| 31 |  | Spurious emission band UE co-existence Influence of noise (c2) (f within NR Bands n257, n260 or n261) |  |  |  |  | FFS |  |
| 32 |  | Spurious emission band UE co-existence Influence of noise (c3) (36 GHz <= f <= 37 GHz) |  |  |  |  | FFS |  |
| 33 |  | Additional spurious emissions Influence of noise (c4) NS_202 (23.6 GHz < f <= 24.0 GHz) |  |  |  |  | FFS |  |
| 34 |  | Additional spurious emissions Influence of noise (c5) NS_202 (23.45 GHz < f <= 40.8 GHz) |  |  |  |  | FFS |  |
| 35 |  | Additional spurious emissions Influence of noise (c6) NS_203 (23.6 GHz < f <= 24.0 GHz) |  |  |  |  | FFS |  |
| 36 |  | Spurious emission band UE co-existence Influence of noise (c7) (23.6 GHz < f <= 24.0 GHz) |  |  |  |  | FSS |  |
| 37 |  | Systematic error related to beam peak search (NOTE 2) |  |  |  |  | N/A |  |
| Total measurement uncertainty |  |  |  |  |  |  | Value |  |
| General spurious emissions Total measurement uncertainty (a)+(b)+(c1) [dB] (23.45 GHz < f <= 40.8 GHz) |  |  |  |  |  |  | [7.11] |  |
| Spurious emission band UE co-existence Total measurement uncertainty (a)+(b)+(c2) [dB] (f within NR Bands n257, n260 or n261) |  |  |  |  |  |  | FFS |  |
| Spurious emission band UE co-existence Total measurement uncertainty (a)+(b)+(c3) [dB] (36 GHz <= f <= 37 GHz) |  |  |  |  |  |  | FFS |  |
| Additional spurious emissions Total measurement uncertainty (a)+(b)+(c4) [dB] NS_202 (23.6 GHz < f <= 24.0 GHz) |  |  |  |  |  |  | FFS |  |
| Additional spurious emissions Total measurement uncertainty (a)+(b)+(c5) [dB] NS_202 (23.45 GHz < f <= 40.8 GHz) |  |  |  |  |  |  | FFS |  |
| Additional spurious emissions Total measurement uncertainty (a)+(b)+(c6) [dB] NS_203 (23.6 GHz < f <= 24.0 GHz) |  |  |  |  |  |  | FFS |  |
| Additional spurious emissions Total measurement uncertainty (a)+(b)+(c7) [dB] (23.6 GHz < f <= 24.0 GHz) |  |  |  |  |  |  | FFS |  |
| NOTE 1: This contributor shall only be considered for TRP measurements. NOTE 2: This contributor shall only be considered for EIRP measurements. NOTE 3: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 4: Value based on procedure defined in clause D.2 of TR 38.810 for Quiet Zone size of less or equal to 30 cm. NOTE 5: Applies to the system which has a structure of mechanical feed antenna positioning. |  |  |  |  |  |  |  |  |

Table B.18.2-15: Uncertainty assessment for TRP measurement (f= 40.8 GHz to 66 GHz, Quiet Zone size ≤ 30 cm) for PC1 UEs

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment | 0.02 | Normal | 2.00 | 0.01 |
| 2 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 3 | Quality of Quiet Zone (NOTE 4) | 0.6 | Actual | 1.00 | 0.6 |
| 4 | Mismatch | 2.30 | Actual | 1.00 | 2.30 |
| 5 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | Uncertainty of the RF power measurement equipment | 4.0 | Normal | 2.00 | 2.00 |
| 7 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 8 | Amplifier uncertainties | 2.1 | Normal | 2.00 | 1.05 |
| 9 | Random uncertainty | 0.5 | Normal | 2.00 | 0.25 |
| 10 | Influence of the XPD | 0.09 | U-shaped | 1.41 | 0.064 |
| 11 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 13 | Influence of TRP measurement grid (NOTE 1) | 0.32 | Actual | 1 | 0.32 |
| 14 | Influence of beam peak search grid (NOTE 2) | N/A | Actual | 1 | N/A |
| 15 | Multiple measurement antenna uncertainty (NOTE 5) | 0.15 | Actual | 1 | 0.15 |
| 16 | DUT repositioning | 0.00 | Rectangular | 1.73 | 0.00 |
| 17 | Misalignment of DUT due to change of DUT orientation | [0.50] | Actual | 1 | [0.50] |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 18 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 19 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 20 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 21 | Uncertainty of the Network Analyzer | 1.7 | Normal | 2.00 | 0.85 |
| 22 | Uncertainty of the absolute gain of the calibration antenna | 1.70 | Normal | 2.00 | 0.85 |
| 23 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.05 | Rectangular | 1.73 | 0.03 |
| 24 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 25 | Quality of quiet zone for calibration process (NOTE 4) | 0.6 | Actual | 1.00 | 0.6 |
| 26 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 27 | Influence of the calibration antenna feed cable | 0.28 | Normal | 2.00 | 0.14 |
| 28 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
|  | Expanded uncertainty (1.96σ - confidence interval of 95 %) |  |  |  | Value |
|  | TRP Expanded uncertainty ( 40.8 GHz < f <= 66 GHz) [dB] (a) |  |  |  | [7.08] |
|  | Systematic uncertainties (NOTE 3) |  |  |  | Value |
| 29 | Systematic error due to TRP calculation/quadrature (NOTE 1) (b) |  |  |  | 0.00 |
| 30 | General spurious emissions Influence of noise (c1) (40.8 GHz < f <= 66 GHz) |  |  |  | 2.02 |
| 31 | Spurious emission band UE co-existence Influence of noise (c2) (57 GHz <= f <= 66 GHz) |  |  |  | FFS |
| 32 | Additional spurious emissions Influence of noise (c3) NS_202 (40.8 GHz < f <= 2nd harmonic of the upper frequency edge of the UL operating band) |  |  |  | FFS |
| 33 | Systematic error related to beam peak search (NOTE 2) |  |  |  | N/A |
| Total measurement uncertainty |  |  |  |  | Value |
| General spurious emissions Total measurement uncertainty (a)+(b)+(c1) [dB] (40.8 GHz < f <= 66 GHz) |  |  |  |  | [9.09] |
| Spurious emission band UE co-existence Total measurement uncertainty (a)+(b)+(c2) [dB] (57 GHz <= f <= 66 GHz) |  |  |  |  | FFS |
| Additional spurious emissions Total measurement uncertainty (a)+(b)+(c3) [dB] NS_202 (40.8 GHz < f <= 2nd harmonic of the upper frequency edge of the UL operating band) |  |  |  |  | FFS |
| NOTE 1: This contributor shall only be considered for TRP measurements. NOTE 2: This contributor shall only be considered for EIRP measurements. NOTE 3: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 4: Value based on procedure defined in clause D.2 of TR 38.810 for Quiet Zone size of less or equal to 30 cm. NOTE 5: Applies to the system which has a structure of mechanical feed antenna positioning. NOTE 6: Void |  |  |  |  |  |

Table B.18.2-16: Uncertainty assessment for TRP measurement (f= 66 GHz to 80 GHz, Quiet Zone size ≤ 30 cm) for PC1 UEs

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment | FFS | Normal | 2.00 | FFS |
| 2 | Measure distance uncertainty | FFS | Rectangular | 1.73 | FFS |
| 3 | Quality of Quiet Zone (NOTE 4) | FFS | Actual | 1.00 | FFS |
| 4 | Mismatch | FFS | Actual | 1.00 | FFS |
| 5 | Standing wave between the DUT and measurement antenna | FFS | U-shaped | 1.41 | FFS |
| 6 | Uncertainty of the RF power measurement equipment | FFS | Normal | 2.00 | FFS |
| 7 | Phase curvature | FFS | U-shaped | 1.41 | FFS |
| 8 | Amplifier uncertainties | FFS | Normal | 2.00 | FFS |
| 9 | Random uncertainty | FFS | Normal | 2.00 | FFS |
| 10 | Influence of the XPD | FFS | U-shaped | 1.41 | FFS |
| 11 | Insertion Loss Variation | FFS | Rectangular | 1.73 | FFS |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | FFS | Actual | 1.00 | FFS |
| 13 | Influence of TRP measurement grid (NOTE 1) | FFS | Actual | 1 | FFS |
| 14 | Influence of beam peak search grid (NOTE 2) | N/A | Actual | 1 | N/A |
| 15 | Multiple measurement antenna uncertainty (NOTE 5) | FFS | Actual | 1 | FFS |
| 16 | DUT repositioning | FFS | Rectangular | 1.73 | FFS |
| 17 | Misalignment of DUT due to change of DUT orientation | FFS | Actual | 1 | FFS |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 18 | Mismatch | FFS | U-shaped | 1.41 | FFS |
| 19 | Amplifier Uncertainties | FFS | Normal | 2.00 | FFS |
| 20 | Misalignment of positioning System | FFS | Normal | 2.00 | FFS |
| 21 | Uncertainty of the Network Analyzer | FFS | Normal | 2.00 | FFS |
| 22 | Uncertainty of the absolute gain of the calibration antenna | FFS | Normal | 2.00 | FFS |
| 23 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | FFS | Rectangular | 1.73 | FFS |
| 24 | Phase centre offset of calibration antenna | FFS | Rectangular | 1.73 | FFS |
| 25 | Quality of quiet zone for calibration process (NOTE 4) | FFS | Actual | 1.00 | FFS |
| 26 | Standing wave between reference calibration antenna and measurement antenna | FFS | U-shaped | 1.41 | FFS |
| 27 | Influence of the calibration antenna feed cable | FFS | Normal | 2.00 | FFS |
| 28 | Insertion Loss Variation | FFS | Rectangular | 1.73 | FFS |
|  | Expanded uncertainty (1.96σ - confidence interval of 95 %) |  |  |  | Value |
|  | TRP Expanded uncertainty (66 GHz < f <= 80 GHz) [dB] (a) |  |  |  | FFS |
|  | Systematic uncertainties (NOTE 3) |  |  |  | Value |
| 29 | Systematic error due to TRP calculation/quadrature (NOTE 1) (b) |  |  |  | FFS |
| 30 | General spurious emissions Influence of noise (c1) (66 GHz < f <= 80 GHz) |  |  |  | FFS |
| 31 | Systematic error related to beam peak search (NOTE 2) |  |  |  | N/A |
| Total measurement uncertainty |  |  |  |  | Value |
| General spurious emissions Total measurement uncertainty (a)+(b)+(c1) [dB] (66 GHz < f <= 80 GHz) |  |  |  |  | FFS |
| NOTE 1: This contributor shall only be considered for TRP measurements. NOTE 2: This contributor shall only be considered for EIRP measurements. NOTE 3: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 4: Value based on procedure defined in clause D.2 of TR 38.810 for Quiet Zone size of less or equal to 30 cm. NOTE 5: Applies to the system which has a structure of mechanical feed antenna positioning. |  |  |  |  |  |

NOTE: MU assessment for additional spurious in Table B.18.2-3 to Table B.18.2-16 is based on the following relaxations:

Table B.18.2-17: Transmitter Spurious emissions relaxation considered in MU assessment (Quiet Zone size ≤ 30 cm)

| Power Class | Frequency | Relaxation |
| --- | --- | --- |
| PC1 | 6 GHz < f <= 12.75 GHz | 0 dB |
|  | 12.75 GHz < f <= 23.45 GHz | 0 dB |
|  | 23.45GHz <= f <= 40.8GHz | 0 dB |
|  | 40.8 GHz < f <= 66 GHz | 0 dB |
|  | 66 GHz < f <= 80 GHz | FFS |
| PC2 | 6 GHz < f <= 12.75 GHz | FFS |
|  | 12.75 GHz < f <= 23.45 GHz | FFS |
|  | 23.45GHz <= f <= 40.8GHz | FFS |
|  | 40.8 GHz < f <= 66 GHz | FFS |
|  | 66 GHz < f <= 80 GHz | FFS |
| PC3 | 6 GHz < f <= 12.75 GHz | 0 dB |
|  | 12.75 GHz < f <= 23.45 GHz | 0 dB |
|  | 23.45GHz <= f <= 40.8GHz | 0 dB |
|  | 40.8 GHz < f <= 66 GHz | 0 dB |
|  | 66 GHz < f <= 80 GHz | 0 dB |
| PC4 | 6 GHz < f <= 12.75 GHz | FFS |
|  | 12.75 GHz < f <= 23.45 GHz | FFS |
|  | 23.45GHz <= f <= 40.8GHz | FFS |
|  | 40.8 GHz < f <= 66 GHz | FFS |
|  | 66 GHz < f <= 80 GHz | FFS |

Table B.18.2-18: Spurious emissions band UE co-existence relaxation considered in MU assessment (Quiet Zone size ≤ 30 cm)

| Power Class | Frequency | Relaxation |
| --- | --- | --- |
| PC1 | 23.45GHz <= f <= 40.8GHz | FFS |
|  | 40.8 GHz < f <= 66 GHz | FFS |
| PC2 | 23.45GHz <= f <= 40.8GHz | FFS |
|  | 40.8 GHz < f <= 66 GHz | FFS |
| PC3 | 23.45GHz <= f <= 40.8GHz | 3.3 dB (for protected bands n257, n261) 5 dB (for protected band n260) 0.3 dB (for 23.6 GHz ≤ f ≤ 24.0 GHz) |
|  | 40.8 GHz < f <= 66 GHz | 6 dB (for 36.0 GHz ≤ f ≤ 37.0 GHz) 0 dB (for 57.0 GHz ≤ f ≤ 66.0 GHz) |
| PC4 | 23.45GHz <= f <= 40.8GHz | FFS |
|  | 40.8 GHz < f <= 66 GHz | FFS |

Table B.18.2-19: Additional Spurious emissions relaxation considered in MU assessment (Quiet Zone size ≤ 30 cm)

| Power Class | Frequency | Relaxation |
| --- | --- | --- |
| PC1 | 6 GHz < f <= 12.75 GHz | FFS |
|  | 12.75 GHz < f <= 23.45 GHz | FFS |
|  | 23.45GHz <= f <= 40.8GHz | FFS |
|  | 40.8 GHz < f <= 66 GHz | FFS |
|  | 66 GHz < f <= 80 GHz | FFS |
| PC2 | 6 GHz < f <= 12.75 GHz | FFS |
|  | 12.75 GHz < f <= 23.45 GHz | FFS |
|  | 23.45GHz <= f <= 40.8GHz | FFS |
|  | 40.8 GHz < f <= 66 GHz | FFS |
|  | 66 GHz < f <= 80 GHz | FFS |
| PC3 | 6 GHz < f <= 12.75 GHz | 0 dB (NS_202) |
|  | 12.75 GHz < f <= 23.45 GHz | 13 dB (NS_202) |
|  | 23.45GHz <= f <= 40.8GHz | 13 dB (whole frequency range for NS_202) 0.3 dB (for 23.6 GHz ≤ f ≤ 24.0 GHz for NS_202 & NS_203) |
|  | 40.8 GHz < f <= 66 GHz | 13 dB (NS_202) |
| PC4 | 6 GHz < f <= 12.75 GHz | FFS |
|  | 12.75 GHz < f <= 23.45 GHz | FFS |
|  | 23.45GHz <= f <= 40.8GHz | FFS |
|  | 40.8 GHz < f <= 66 GHz | FFS |
|  | 66 GHz < f <= 80 GHz | FFS |

<<Unchanged sections skipped>>

B.21	Adjacent Channel Selectivity

Following tables summarize the MU threshold for Adjacent Channel Selectivity measurement. The origin MU values for different test setups with varies parameters can be found in following subclauses.

Table B.21-1: MU threshold for Adjacent Channel Selectivity

| Power Class | Frequency | MBW | Power | Threshold MU value (NOTE 1) |
| --- | --- | --- | --- | --- |
| PC3 | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = Max Output Power | 7.84 |
|  | 32.125GHz < f <= 40.8GHz |  |  | 7.84 |
| PC1 | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = Max Output Power | 8.06 |
|  | 32.125GHz < f <= 40.8GHz |  |  | FFS |
| NOTE 1: Total Expanded MU for IFF for Quiet Zone size ≤ 30cm in Table B.21.2-2 for PC3 UEs and Table B.21.2-3 for PC1 UEs. |  |  |  |  |

B.21.1	Uncertainty budget format and assessment for DFF

FFS

B.21.2	Uncertainty budget format and assessment for IFF

The uncertainty contributions that may impact the overall MU value are listed in Table B.21.2-1.

Table B.21.2-1: Total Uncertainty contributions for Adjacent Channel Selectivity measurement

| UID | Description of uncertainty contribution | Details in clause |
| --- | --- | --- |
| Stage 2: DUT measurement (Wanted Signal contributions) |  |  |
| 1 | Positioning misalignment | B.2.2.1 |
| 2 | Measure distance uncertainty | B.2.2.2 |
| 3 | Quality of Quiet Zone | B.2.2.3 |
| 4 | Mismatch | B.2.2.4 |
| 5 | Standing wave between the DUT and measurement antenna | B.2.2.5 |
| 6 | gNB emulator uncertainty | B.2.2.17 |
| 7 | Phase curvature | B.2.2.7 |
| 8 | Amplifier uncertainties | B.2.2.8 |
| 9 | Random uncertainty | B.2.2.9 |
| 10 | Influence of the XPD | B.2.2.10 |
| 11 | Insertion Loss Variation | B.2.2.11 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | B.2.2.12 |
| 13 | Multiple measurement antenna uncertainty | B.2.2.25 |
| 14 | DUT repositioning | B.2.2.26 |
| Stage 2: DUT measurement (Modulated Interferer Signal specific contributions) |  |  |
| 15 | Positioning misalignment | B.2.2.1 |
| 16 | Measure distance uncertainty | B.2.2.2 |
| 17 | Quality of Quiet Zone | B.2.2.3 |
| 18 | Mismatch | B.2.2.4 |
| 19 | Standing wave between the DUT and measurement antenna | B.2.2.5 |
| 20 | Modulated Interferer uncertainty | B.2.2.33 |
| 21 | Phase curvature | B.2.2.7 |
| 22 | Amplifier uncertainties | B.2.2.8 |
| 23 | Random uncertainty | B.2.2.9 |
| 24 | Influence of the XPD | B.2.2.10 |
| 25 | Insertion Loss Variation | B.2.2.11 |
| 26 | RF leakage (from measurement antenna to the receiver/transmitter) | B.2.2.12 |
| 27 | Multiple measurement antenna uncertainty | B.2.2.25 |
| 28 | DUT repositioning | B.2.2.26 |
| 29 | Influence of offset antenna (Std.Dev) | B.2.2.35 |
| Stage 1: Calibration measurement (Wanted Signal contributions) |  |  |
| 30 | Mismatch | B.2.2.4 |
| 31 | Amplifier Uncertainties | B.2.2.8 |
| 32 | Misalignment of positioning System | B.2.2.13 |
| 33 | Uncertainty of the Network Analyzer | B.2.2.14 |
| 34 | Uncertainty of the absolute gain of the calibration antenna | B.2.2.15 |
| 35 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | B.2.2.16 |
| 36 | Phase centre offset of calibration antenna | B.2.2.18 |
| 37 | Quality of quiet zone for calibration process | B.2.2.19 |
| 38 | Standing wave between reference calibration antenna and measurement antenna | B.2.2.20 |
| 39 | Influence of the calibration antenna feed cable | B.2.2.21 |
| 40 | Insertion Loss Variation | B.2.2.11 |
| Stage 1: Calibration measurement (Modulated Interferer Signal contributions) |  |  |
| 41 | Mismatch | B.2.2.4 |
| 42 | Amplifier Uncertainties | B.2.2.8 |
| 43 | Misalignment of positioning System | B.2.2.13 |
| 44 | Uncertainty of the Network Analyzer | B.2.2.14 |
| 45 | Uncertainty of the absolute gain of the calibration antenna | B.2.2.15 |
| 46 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | B.2.2.16 |
| 47 | Phase centre offset of calibration antenna | B.2.2.18 |
| 48 | Quality of quiet zone for calibration process | B.2.2.19 |
| 48 | Standing wave between reference calibration antenna and measurement antenna | B.2.2.20 |
| 50 | Influence of the calibration antenna feed cable | B.2.2.21 |
| 51 | Insertion Loss Variation | B.2.2.11 |
| Systematic uncertainties |  |  |
| 52 | Systematic error related to beam peak search | B.2.2.28 |
| 53 | Additional impact of interferer ACLR | B.2.2.32 |
| 54 | Influence of offset antenna (mean error) | B.2.2.35 |

The uncertainty assessment tables are organized as follows:

-	For the purpose of uncertainty assessment, the radiating antenna aperture of the DUT is denoted as D

-	The uncertainty assessment has been derived for the case of Quiet Zone size ≤ [30 cm], f = {23.45GHz, 32.125GHz, 40.8GHz}, [P = maximum output power].

-	The uncertainty assessment for ACS is provided in Table B.21.2-2 for PC3 UEs and Table B.21.2-3 for PC1 UEs.

Table B.21.2-2: Uncertainty assessment for Adjacent Channel Selectivity measurement (f=23.45GHz, 32.125GHz, 40.8GHz, Quiet Zone size ≤ 30 cm) for PC3 UEs

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement (Wanted Signal contributions) |  |  |  |  |  |
| 1 | Positioning misalignment | 0.00 | Normal | 2.00 | 0.00 |
| 2 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 3 | Quality of Quiet Zone (NOTE 4) | 0.6 | Actual | 1.00 | 0.6 |
| 4 | Mismatch | 1.30 | Actual | 1.00 | 1.30 |
| 5 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | gNB uncertainty on absolute level | 2.9 | Normal | 2.00 | 1.45 |
| 7 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 8 | Amplifier uncertainties | 2.1 | Normal | 2.00 | 1.05 |
| 9 | Random uncertainty | 0.50 | Normal | 2.00 | 0.25 |
| 10 | Influence of the XPD | 0.01 | U-shaped | 1.41 | 0.00 |
| 11 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 13 | Multiple measurement antenna uncertainty (NOTE 3) | 0.15 | Actual | 1.00 | 0.15 |
| 14 | DUT repositioning | 0.08 | Rectangular | 1.73 | 0.05 |
| Stage 2: DUT measurement (Modulated Interferer Signal specific contributions) |  |  |  |  |  |
| 15 | Positioning misalignment | 0.00 | Normal | 2.00 | 0.00 |
| 16 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 17 | Quality of Quiet Zone (NOTE 4) | 0.6 | Actual | 1.00 | 0.6 |
| 18 | Mismatch | 1.30 | Actual | 1.00 | 1.30 |
| 19 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 20 | Modulated Interferer uncertainty on absolute level | 2.9 | Normal | 2.00 | 1.45 |
| 21 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 22 | Amplifier uncertainties | 2.1 | Normal | 2.00 | 1.05 |
| 23 | Random uncertainty | 0.50 | Normal | 2.00 | 0.25 |
| 24 | Influence of the XPD | 0.01 | U-shaped | 1.41 | 0.00 |
| 25 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 26 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 27 | Multiple measurement antenna uncertainty (NOTE 3) | 0.15 | Actual | 1.00 | 0.15 |
| 28 | DUT repositioning | 0.08 | Rectangular | 1.73 | 0.05 |
| 29 | Influence of offset antenna (Std.Dev) (NOTE 5) | 0.00 | Normal | 2.00 | 0.00 |
| Stage 1: Calibration measurement (Wanted Signal contributions) |  |  |  |  |  |
| 30 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 31 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 32 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 33 | Uncertainty of the Network Analyzer | 0.73 | Normal | 2.00 | 0.37 |
| 34 | Uncertainty of the absolute gain of the calibration antenna | 0.60 | Normal | 2.00 | 0.30 |
| 35 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.01 | Rectangular | 1.73 | 0.00 |
| 36 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 37 | Quality of quiet zone for calibration process (NOTE 4) | 0.4 | Actual | 1.00 | 0.4 |
| 38 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 39 | Influence of the calibration antenna feed cable | 0.14 | Normal | 2.00 | 0.07 |
| 40 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| Stage 1: Calibration measurement (Modulated Interferer Signal contributions) |  |  |  |  |  |
| 41 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 42 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 43 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 44 | Uncertainty of the Network Analyzer | 0.73 | Normal | 2.00 | 0.37 |
| 45 | Uncertainty of the absolute gain of the calibration antenna | 0.60 | Normal | 2.00 | 0.30 |
| 46 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.01 | Rectangular | 1.73 | 0.00 |
| 47 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 48 | Quality of quiet zone for calibration process (NOTE 4) | 0.4 | Actual | 1.00 | 0.4 |
| 48 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 50 | Influence of the calibration antenna feed cable | 0.14 | Normal | 2.00 | 0.07 |
| 51 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
|  | Systematic uncertainties (NOTE 2) |  |  |  | Value |
| 52 | Systematic error related to beam peak search |  |  |  | 0.5 |
| 53 | Additional impact of interferer ACLR |  |  |  | 0.7 |
| 54 | Influence of offset antenna (mean error) (NOTE 5) |  |  |  | 0.00 |
| Total measurement uncertainty |  |  |  |  | Value |
| ACS Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 7.84 |
| NOTE 1: The analysis was done only for the case of operating at max output power, in-band, non-CA. NOTE 2: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 3: Applies to the system which has a structure of mechanical feed antenna positioning. NOTE 4: Value based on procedure defined in clause D.2 of TR 38.810 for Quiet Zone size less or equal to 30 cm. NOTE 5: For MTSU derivation purpose, this value is set to 0.0 (no offset antenna case). |  |  |  |  |  |

Table B.21.2-3: Uncertainty assessment for Adjacent Channel Selectivity measurement (f=23.45GHz, 32.125GHz, 40.8GHz, Quiet Zone size ≤ 30 cm) for PC1 UEs

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement (Wanted Signal contributions) |  |  |  |  |  |
| 1 | Positioning misalignment | 0.02 | Normal | 2.00 | 0.01 |
| 2 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 3 | Quality of Quiet Zone (NOTE 4) | 0.6 | Actual | 1.00 | 0.6 |
| 4 | Mismatch | 1.30 | Actual | 1.00 | 1.30 |
| 5 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | gNB uncertainty on absolute level | 2.9 | Normal | 2.00 | 1.45 |
| 7 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 8 | Amplifier uncertainties | 2.1 | Normal | 2.00 | 1.05 |
| 9 | Random uncertainty | 0.50 | Normal | 2.00 | 0.25 |
| 10 | Influence of the XPD | 0.01 | U-shaped | 1.41 | 0.00 |
| 11 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 13 | Multiple measurement antenna uncertainty (NOTE 3) | 0.15 | Actual | 1.00 | 0.15 |
| 14 | DUT repositioning | 0.35 | Rectangular | 1.73 | 0.20 |
| Stage 2: DUT measurement (Modulated Interferer Signal specific contributions) |  |  |  |  |  |
| 15 | Positioning misalignment | 0.02 | Normal | 2.00 | 0.01 |
| 16 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 17 | Quality of Quiet Zone (NOTE 4) | 0.6 | Actual | 1.00 | 0.6 |
| 18 | Mismatch | 1.30 | Actual | 1.00 | 1.30 |
| 19 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 20 | Modulated Interferer uncertainty on absolute level | 2.9 | Normal | 2.00 | 1.45 |
| 21 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 22 | Amplifier uncertainties | 2.1 | Normal | 2.00 | 1.05 |
| 23 | Random uncertainty | 0.50 | Normal | 2.00 | 0.25 |
| 24 | Influence of the XPD | 0.01 | U-shaped | 1.41 | 0.00 |
| 25 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 26 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 27 | Multiple measurement antenna uncertainty (NOTE 3) | 0.15 | Actual | 1.00 | 0.15 |
| 28 | DUT repositioning | 0.35 | Rectangular | 1.73 | 0.20 |
| 29 | Influence of offset antenna (Std.Dev) (NOTE 5) | 0.00 | Normal | 2.00 | 0.00 |
| Stage 1: Calibration measurement (Wanted Signal contributions) |  |  |  |  |  |
| 30 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 31 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 32 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 33 | Uncertainty of the Network Analyzer | 0.73 | Normal | 2.00 | 0.37 |
| 34 | Uncertainty of the absolute gain of the calibration antenna | 0.60 | Normal | 2.00 | 0.30 |
| 35 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.01 | Rectangular | 1.73 | 0.00 |
| 36 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 37 | Quality of quiet zone for calibration process (NOTE 4) | 0.4 | Actual | 1.00 | 0.4 |
| 38 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 39 | Influence of the calibration antenna feed cable | 0.14 | Normal | 2.00 | 0.07 |
| 40 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| Stage 1: Calibration measurement (Modulated Interferer Signal contributions) |  |  |  |  |  |
| 41 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 42 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 43 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 44 | Uncertainty of the Network Analyzer | 0.73 | Normal | 2.00 | 0.37 |
| 45 | Uncertainty of the absolute gain of the calibration antenna | 0.60 | Normal | 2.00 | 0.30 |
| 46 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.01 | Rectangular | 1.73 | 0.00 |
| 47 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 48 | Quality of quiet zone for calibration process (NOTE 4) | 0.4 | Actual | 1.00 | 0.4 |
| 48 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 50 | Influence of the calibration antenna feed cable | 0.14 | Normal | 2.00 | 0.07 |
| 51 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
|  | Systematic uncertainties (NOTE 2) |  |  |  | Value |
| 52 | Systematic error related to beam peak search |  |  |  | 0.7 |
| 53 | Additional impact of interferer ACLR |  |  |  | 0.7 |
| 54 | Influence of offset antenna (mean error) (NOTE 5) |  |  |  | 0.00 |
| Total measurement uncertainty |  |  |  |  | Value |
| ACS Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] (23.45GHz <= f <= 32.125GHz) |  |  |  |  | 8.06 |
| ACS Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] (32.125GHz < f <= 40.8GHz) |  |  |  |  | FFS |
| NOTE 1: The analysis was done only for the case of operating at max output power, in-band, non-CA. NOTE 2: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 3: Applies to the system which has a structure of mechanical feed antenna positioning. NOTE 4: Value based on procedure defined in clause D.2 of TR 38.810 for Quiet Zone size less or equal to 30 cm. NOTE 5: For MTSU derivation purpose, this value is set to 0.0 (no offset antenna case). |  |  |  |  |  |

B.22	In-Band Blocking

See B.21.

B.22.1	Uncertainty budget format and assessment for DFF

See B.21.1.

B.22.2	Uncertainty budget format and assessment for IFF

See B.21.2.

B.23

B.24

B.25	Receiver spurious emissions

Editor’s Note:

-	MU value analysis and offset value analysis for PC1, 2 and 4 are not complete.

-	MU value analysis for various test setups in subsection B.25.x is not complete for above 80 GHz for PC3

-	Offset value analysis is not complete as it is derived from MU value analysis for above 80 GHz for PC3

Test procedure of general spurious emission comprises 2 stages: coarse TRP measurement and fine TRP measurement BW. Coarse TRP measurement is introduced to reduce the measurement time by applying sparser grids and/or wider measurement BW than fine TRP measurement while having offset dB more stringent test requirement in order not to cause additional misjudgement risk. For the frequency ranges for which coarse TRP measurement does not PASS, the measurement is continued with fine TRP measurement procedure.

Table B.25-1 summarizes the MU threshold for fine TRP measurements for General spurious emissions. The origin MU values for fine TRP measurement for different test setups can be found in following subclauses.

Table B.25-1: MU threshold for TRP measurement for Rx spurious emission

| Power Class | Frequency | In-band BW | In-band Power (NOTE2) | Threshold MU value [dB] (NOTE1) |
| --- | --- | --- | --- | --- |
| PC3 | 6 GHz <= f <=12.75 GHz | BW <= 400MHz | P = Max Output Power | 5.50 |
|  | 12.75 GHz <= f <= 23.45 GHz |  |  | 5.46 |
|  | 23.45 GHz <= f <= 40.8 GHz |  |  | 6.11 |
|  | 40.8 GHz <= f <= 66 GHz |  |  | 7.65 |
|  | 66 GHz <= f <= 80 GHz |  |  | 7.95 |
| PC1 | 6 GHz <= f <=12.75 GHz | BW <= 400MHz | P = Max Output Power | [5.60] |
|  | 12.75 GHz <= f <= 23.45 GHz |  |  | [5.56] |
|  | 23.45 GHz <= f <= 40.8 GHz |  |  | [6.20] |
|  | 40.8 GHz <= f <= 66 GHz |  |  | [7.72] |
|  | 66 GHz <= f <= 80 GHz |  |  | FFS |
| NOTE 1: Total EIRP Expanded MU for IFF for Quiet Zone size ≤ 30cm in Table B.25.2-3 to Table B.25.2-11 for PC3 UEs and in Table B.25.2.12 to Table B.25.2.16 for PC1 UEs. |  |  |  |  |

Table B.25-2 provides valid coarse TRP measurement grids and corresponding offset dB value that may be used for UE general spurious emission test case.  The offset value is derived as 95%-tile TRP measurement uncertainty including the effect from uncertainty due to Coarse TRP measurement grid, excluding influence of noise.

Table B.25-2: Coarse TRP measurement grids and offset values for UE Rx spurious emission

| Power Class | Coarse TRP measurement grid | Frequency | Min Number of measurement points on the grid | Influence of coarse TRP measurement grid (dB) | Systematic error due to coarse TRP calculation/quadrature (dB) | Offset value (dB) |
| --- | --- | --- | --- | --- | --- | --- |
| PC3 | Constant density grid (charged particle based) | 6 GHz <= f <=12.75 GHz | 35 | 0.94 | 0.09 | 5.25 |
|  |  | 12.75 GHz <= f <= 23.45 GHz |  |  |  | 5.21 |
|  |  | 23.45 GHz <= f <= 40.8 GHz |  |  |  | 5.49 |
|  |  | 40.8 GHz <= f <= 66 GHz |  |  |  | 7.31 |
|  |  | 66 GHz <= f <= 80 GHz |  |  |  | 7.61 |
|  | Constant step size grid | 6 GHz <= f <=12.75 GHz | 62 | 0.97 | 0.2 | 5.38 |
|  |  | 12.75 GHz <= f <= 23.45 GHz |  |  |  | 5.34 |
|  |  | 23.45 GHz <= f <= 40.8 GHz |  |  |  | 5.62 |
|  |  | 40.8 GHz <= f <= 66 GHz |  |  |  | 7.43 |
|  |  | 66 GHz <= f <= 80 GHz |  |  |  | 7.73 |
| PC1 | Constant density grid (charged particle based) | 6 GHz <= f <=12.75 GHz | FFS | FFS | FFS | FFS |
|  |  | 12.75 GHz <= f <= 23.45 GHz |  |  |  | FFS |
|  |  | 23.45 GHz <= f <= 40.8 GHz |  |  |  | FFS |
|  |  | 40.8 GHz <= f <= 66 GHz |  |  |  | FFS |
|  |  | 66 GHz <= f <= 80 GHz |  |  |  | FFS |
|  | Constant step size grid | 6 GHz <= f <=12.75 GHz | FFS | FFS | FFS | FFS |
|  |  | 12.75 GHz <= f <= 23.45 GHz |  |  |  | FFS |
|  |  | 23.45 GHz <= f <= 40.8 GHz |  |  |  | FFS |
|  |  | 40.8 GHz <= f <= 66 GHz |  |  |  | FFS |
|  |  | 66 GHz <= f <= 80 GHz |  |  |  | FFS |
| NOTE 1: Based on Total TRP Expanded MU for IFF for Quiet Zone size ≤ 30cm in Table B.25.2-3 to Table B.25.2-11, replacing “Influence of TRP measurement grid” and “Systematic error due to TRP calculation/quadrature” by the values for coarse TRP grid, and excluding “Influence of noise”. |  |  |  |  |  |  |

<<Unchanged sections skipped>>

B.25.2	Uncertainty budget format and assessment for IFF

The uncertainty contributions that may impact the overall MU value are listed in Table B.25.2-1.

Table B.25.2-1: Uncertainty contributions for TRP measurement

| UID | Description of uncertainty contribution | Details in clause |
| --- | --- | --- |
| Stage 2: DUT measurement |  |  |
| 1 | Positioning misalignment | B.2.2.1 |
| 2 | Measure distance uncertainty | B.2.2.2 |
| 3 | Quality of Quiet Zone | B.2.2.3 |
| 4 | Mismatch | B.2.2.4 |
| 5 | Standing wave between the DUT and measurement antenna | B.2.2.5 |
| 6 | Uncertainty of the RF power measurement equipment | B.2.2.6 |
| 7 | Phase curvature | B.2.2.7 |
| 8 | Amplifier uncertainties | B.2.2.8 |
| 9 | Random uncertainty | B.2.2.9 |
| 10 | Influence of the XPD | B.2.2.10 |
| 11 | Insertion Loss Variation | B.2.2.11 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | B.2.2.12 |
| 13 | Influence of TRP measurement grid | B.2.2.22 |
| 14 | Influence of beam peak search grid | B.2.2.23 |
| 15 | Multiple measurement antenna uncertainty | B.2.2.25 |
| 16 | DUT repositioning | B.2.2.26 |
| 17 | Misalignment of DUT due to change of DUT orientation | B.2.2.31 |
| Stage 1: Calibration measurement |  |  |
| 18 | Mismatch | B.2.2.4 |
| 19 | Amplifier Uncertainties | B.2.2.8 |
| 20 | Misalignment of positioning System | B.2.2.13 |
| 21 | Uncertainty of the Network Analyzer | B.2.2.14 |
| 22 | Uncertainty of the absolute gain of the calibration antenna | B.2.2.15 |
| 23 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | B.2.2.16 |
| 24 | Phase centre offset of calibration antenna | B.2.2.18 |
| 25 | Quality of quiet zone for calibration process | B.2.2.19 |
| 26 | Standing wave between reference calibration antenna and measurement antenna | B.2.2.20 |
| 27 | Influence of the calibration antenna feed cable | B.2.2.21 |
| 28 | Insertion Loss Variation | B.2.2.11 |
| Systematic uncertainties |  |  |
| 29 | Systematic error due to TRP calculation/quadrature | B.2.2.24 |
| 30 | Influence of noise | B.2.2.27 |

The uncertainty assessment tables are organized as follows:

-	For the purpose of uncertainty assessment, the radiating antenna aperture of the DUT is denoted as D

-	The uncertainty assessment has been derived for the case of Quiet Zone size ≤ 30 cm, f = {6 GHz to 80 GHz}, P =Receiver Spurious Core Requirement Level + Relaxation(For n257, 10.2dB for 6GHz  f < 20GHz, 17.2 dB for 20GHz f < 40GHz, 33.1dB for 40GHz  f  2nd harmonic)

-	The uncertainty assessment for TRP is provided from Table B.25.2-2 to Table B.25.2-11 for PC3 UEs and from Table B.25.2.12 to Table B.25.2.16 for PC1 UEs.

Table B.25.2-2: Void

Table B.25.2-3: Uncertainty assessment for TRP measurement (f=6 GHz to 12.75 GHz, Quiet Zone size ≤ 30 cm) for PC3 UEs

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment | 0.00 | Normal | 2.00 | 0.00 |
| 2 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 3 | Quality of Quiet Zone (NOTE 4) | 0.70 | Actual | 1.00 | 0.70 |
| 4 | Mismatch | 1.60 | Actual | 1.00 | 1.60 |
| 5 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | Uncertainty of the RF power measurement equipment | 2.00 | Normal | 2.00 | 1.00 |
| 7 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 8 | Amplifier uncertainties | 2.1 | Normal | 2.00 | 1.05 |
| 9 | Random uncertainty | 0.5 | Normal | 2.00 | 0.25 |
| 10 | Influence of the XPD | 0.09 | U-shaped | 1.41 | 0.064 |
| 11 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 13 | Influence of TRP measurement grid (NOTE 1) | 0.32 | Actual | 1 | 0.32 |
| 14 | Influence of beam peak search grid (NOTE 2) | N/A | Actual | 1 | N/A |
| 15 | Multiple measurement antenna uncertainty (NOTE 5) | 0.15 | Actual | 1 | 0.15 |
| 16 | DUT repositioning | 0.00 | Rectangular | 1.73 | 0.00 |
| 17 | Misalignment of DUT due to change of DUT orientation | 0.10 | Actual | 1 | 0.10 |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 18 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 19 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 20 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 21 | Uncertainty of the Network Analyzer | 0.90 | Normal | 2.00 | 0.45 |
| 22 | Uncertainty of the absolute gain of the calibration antenna | 0.60 | Normal | 2.00 | 0.30 |
| 23 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.05 | Rectangular | 1.73 | 0.03 |
| 24 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 25 | Quality of quiet zone for calibration process (NOTE 4) | 0.70 | Actual | 1.00 | 0.70 |
| 26 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 27 | Influence of the calibration antenna feed cable | 0.14 | Normal | 2.00 | 0.07 |
| 28 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
|  | Expanded uncertainty (1.96σ - confidence interval of 95 %) |  |  |  | Value |
|  | TRP Expanded uncertainty (6 GHz < f <= 12.75 GHz) [dB] (a) |  |  |  | 4.86 |
|  | Systematic uncertainties (NOTE 3) |  |  |  | Value |
| 29 | Systematic error due to TRP calculation/quadrature (NOTE 1) (b) |  |  |  | 0.0 |
| 30 | Influence of noise (6 GHz < f <= 12.75 GHz) (c) |  |  |  | 0.64 |
| 31 | Systematic error related to beam peak search (NOTE 2) |  |  |  | N/A |
| Total measurement uncertainty |  |  |  |  | Value |
| Total measurement uncertainty (a)+(b)+(c) [dB] |  |  |  |  | 5.50 |
| NOTE 1: This contributor shall only be considered for TRP measurements. NOTE 2: This contributor shall only be considered for EIRP measurements. NOTE 3: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 4: Value based on procedure defined in clause D.2 of TR 38.810 for Quiet Zone size of less or equal to 30 cm. NOTE 5: Applies to the system which has a structure of mechanical feed antenna positioning. |  |  |  |  |  |

Table B.25.2-4: void

Table B.25.2-5: Uncertainty assessment for TRP measurement (f=12.75 GHz to 23.45 GHz, Quiet Zone size ≤ 30 cm) for PC3 UEs

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment | 0.00 | Normal | 2.00 | 0.00 |
| 2 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 3 | Quality of Quiet Zone (NOTE 4) | 0.60 | Actual | 1.00 | 0.60 |
| 4 | Mismatch | 1.60 | Actual | 1.00 | 1.60 |
| 5 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | Uncertainty of the RF power measurement equipment | 2.16 | Normal | 2.00 | 1.08 |
| 7 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 8 | Amplifier uncertainties | 2.1 | Normal | 2.00 | 1.05 |
| 9 | Random uncertainty | 0.5 | Normal | 2.00 | 0.25 |
| 10 | Influence of the XPD | 0.09 | U-shaped | 1.41 | 0.064 |
| 11 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 13 | Influence of TRP measurement grid (NOTE 1) | 0.32 | Actual | 1 | 0.32 |
| 14 | Influence of beam peak search grid (NOTE 2) | N/A | Actual | 1 | N/A |
| 15 | Multiple measurement antenna uncertainty (NOTE 5) | 0.15 | Actual | 1 | 0.15 |
| 16 | DUT repositioning | 0.00 | Rectangular | 1.73 | 0.00 |
| 17 | Misalignment of DUT due to change of DUT orientation | 0.10 | Actual | 1 | 0.10 |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 18 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 19 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 20 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 21 | Uncertainty of the Network Analyzer | 0.90 | Normal | 2.00 | 0.45 |
| 22 | Uncertainty of the absolute gain of the calibration antenna | 0.60 | Normal | 2.00 | 0.30 |
| 23 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.05 | Rectangular | 1.73 | 0.03 |
| 24 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 25 | Quality of quiet zone for calibration process (NOTE 4) | 0.60 | Actual | 1.00 | 0.60 |
| 26 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 27 | Influence of the calibration antenna feed cable | 0.14 | Normal | 2.00 | 0.07 |
| 28 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
|  | Expanded uncertainty (1.96σ - confidence interval of 95 %) |  |  |  | Value |
|  | TRP Expanded uncertainty ([12.75] GHz < f <= 23.45 GHz) [dB] (a) |  |  |  | 4.82 |
|  | Systematic uncertainties (NOTE 3) |  |  |  | Value |
| 29 | Systematic error due to TRP calculation/quadrature (NOTE 1) (b) |  |  |  | 0.0 |
| 30 | Influence of noise (12.75 GHz < f <= 23.45 GHz) (c) |  |  |  | 0.64 |
| 31 | Systematic error related to beam peak search (NOTE 2) |  |  |  | N/A |
| Total measurement uncertainty |  |  |  |  | Value |
| Total measurement uncertainty (a)+(b)+(c) [dB] |  |  |  |  | 5.46 |
| NOTE 1: This contributor shall only be considered for TRP measurements. NOTE 2: This contributor shall only be considered for EIRP measurements. NOTE 3: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 4: Value based on procedure defined in clause D.2 of TR 38.810 for Quiet Zone size of less or equal to 30 cm. NOTE 5: Applies to the system which has a structure of mechanical feed antenna positioning. |  |  |  |  |  |

Table B.25.2-6: Void

Table B.25.2-7: Uncertainty assessment for TRP measurement (f=23.45 GHz to 40.8 GHz, Quiet Zone size ≤ 30 cm) for PC3 UEs

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment | 0.00 | Normal | 2.00 | 0.00 |
| 2 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 3 | Quality of Quiet Zone (NOTE 4) | 0.6 | Actual | 1.00 | 0.6 |
| 4 | Mismatch | 1.50 | Actual | 1.00 | 1.50 |
| 5 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | Uncertainty of the RF power measurement equipment | 2.73 | Normal | 2.00 | 1.37 |
| 7 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 8 | Amplifier uncertainties | 2.1 | Normal | 2.00 | 1.05 |
| 9 | Random uncertainty | 0.5 | Normal | 2.00 | 0.25 |
| 10 | Influence of the XPD | 0.01 | U-shaped | 1.41 | 0.00 |
| 11 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 13 | Influence of TRP measurement grid (NOTE 1) | 0.32 | Actual | 1 | 0.32 |
| 14 | Influence of beam peak search grid (NOTE 2) | N/A | Actual | 1 | N/A |
| 15 | Multiple measurement antenna uncertainty (NOTE 5) | 0.15 | Actual | 1 | 0.15 |
| 16 | DUT repositioning | 0.00 | Rectangular | 1.73 | 0.00 |
| 17 | Misalignment of DUT due to change of DUT orientation | 0.10 | Actual | 1 | 0.10 |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 18 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 19 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 20 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 21 | Uncertainty of the Network Analyzer | 1.5 | Normal | 2.00 | 0.75 |
| 22 | Uncertainty of the absolute gain of the calibration antenna | 0.6 | Normal | 2.00 | 0.3 |
| 23 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.05 | Rectangular | 1.73 | 0.03 |
| 24 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 25 | Quality of quiet zone for calibration process (NOTE 4) | 0.6 | Actual | 1.00 | 0.6 |
| 26 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 27 | Influence of the calibration antenna feed cable | 0.14 | Normal | 2.00 | 0.07 |
| 28 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
|  | Expanded uncertainty (1.96σ - confidence interval of 95 %) |  |  |  | Value |
|  | TRP Expanded uncertainty (23.45 GHz < f <= 40.8 GHz) [dB] (a) |  |  |  | 5.11 |
|  | Systematic uncertainties (NOTE 3) |  |  |  | Value |
| 29 | Systematic error due to TRP calculation/quadrature (NOTE 1) (b) |  |  |  | 0.0 |
| 30 | Influence of noise (23.45 GHz < f <= 40.8 GHz) (c) |  |  |  | 1.0 |
| 31 | Systematic error related to beam peak search (NOTE 2) |  |  |  | N/A |
| Total measurement uncertainty |  |  |  |  | Value |
| Total measurement uncertainty (a)+(b)+(c) [dB] |  |  |  |  | 6.11 |
| NOTE 1: This contributor shall only be considered for TRP measurements. NOTE 2: This contributor shall only be considered for EIRP measurements. NOTE 3: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 4: Value based on procedure defined in clause D.2 of TR 38.810 for Quiet Zone size of less or equal to 30 cm. NOTE 5: Applies to the system which has a structure of mechanical feed antenna positioning. |  |  |  |  |  |

Table B.25.2-8: Void

Table B.25.2-9: Uncertainty assessment for TRP measurement (f= 40.8 GHz to 66 GHz, Quiet Zone size ≤ 30 cm) for PC3 UEs

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positi0.0oning misalignment |  | Normal | 2.00 | 0.0 |
| 2 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 3 | Quality of Quiet Zone (NOTE 4) | 0.6 | Actual | 1.00 | 0.6 |
| 4 | Mismatch | 2.30 | Actual | 1.00 | 2.30 |
| 5 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | Uncertainty of the RF power measurement equipment | 4.0 | Normal | 2.00 | 2.00 |
| 7 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 8 | Amplifier uncertainties | 2.1 | Normal | 2.00 | 1.05 |
| 9 | Random uncertainty | 0.5 | Normal | 2.00 | 0.25 |
| 10 | Influence of the XPD | 0.09 | U-shaped | 1.41 | 0.064 |
| 11 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 13 | Influence of TRP measurement grid (NOTE 1) | 0.32 | Actual | 1 | 0.32 |
| 14 | Influence of beam peak search grid (NOTE 2) | N/A | Actual | 1 | N/A |
| 15 | Multiple measurement antenna uncertainty (NOTE 5) | 0.15 | Actual | 1 | 0.15 |
| 16 | DUT repositioning | 0.00 | Rectangular | 1.73 | 0.00 |
| 17 | Misalignment of DUT due to change of DUT orientation | 0.10 | Actual | 1 | 0.10 |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 18 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 19 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 20 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 21 | Uncertainty of the Network Analyzer | 1.7 | Normal | 2.00 | 0.85 |
| 22 | Uncertainty of the absolute gain of the calibration antenna | 1.70 | Normal | 2.00 | 0.85 |
| 23 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.05 | Rectangular | 1.73 | 0.03 |
| 24 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 25 | Quality of quiet zone for calibration process (NOTE 4) | 0.6 | Actual | 1.00 | 0.6 |
| 26 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 27 | Influence of the calibration antenna feed cable | 0.28 | Normal | 2.00 | 0.14 |
| 28 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
|  | Expanded uncertainty (1.96σ - confidence interval of 95 %) |  |  |  | Value |
|  | TRP Expanded uncertainty ( 40.8 GHz < f <= 66 GHz) [dB] (a) |  |  |  | 7.01 |
|  | Systematic uncertainties (NOTE 3) |  |  |  | Value |
| 29 | Systematic error due to TRP calculation/quadrature (NOTE 1) (b) |  |  |  | 0.0 |
| 30 | Influence of noise ( 40.8 GHz < f <= 66 GHz) (c) |  |  |  | 0.64 |
| 31 | Systematic error related to beam peak search (NOTE 2) |  |  |  | N/A |
| Total measurement uncertainty |  |  |  |  | Value |
| Total measurement uncertainty (a)+(b)+(c) [dB] |  |  |  |  | 7.65 |
| NOTE 1: This contributor shall only be considered for TRP measurements. NOTE 2: This contributor shall only be considered for EIRP measurements. NOTE 3: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 4: Value based on procedure defined in clause D.2 of TR 38.810 for Quiet Zone size of less or equal to 30 cm. NOTE 5: Applies to the system which has a structure of mechanical feed antenna positioning. |  |  |  |  |  |

Table B.25.2-10: Void

Table B.25.2-11: Uncertainty assessment for TRP measurement (f= 66 GHz to 80 GHz, Quiet Zone size ≤ 30 cm) for PC3 UEs

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment | 0.00 | Normal | 2.00 | 0.00 |
| 2 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 3 | Quality of Quiet Zone (NOTE 4) | 0.6 | Actual | 1.00 | 0.6 |
| 4 | Mismatch | 2.30 | Actual | 1.00 | 2.30 |
| 5 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | Uncertainty of the RF power measurement equipment | 4.0 | Normal | 2.00 | 2.0 |
| 7 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 8 | Amplifier uncertainties | 3.0 | Normal | 2.00 | 1.50 |
| 9 | Random uncertainty | 0.5 | Normal | 2.00 | 0.25 |
| 10 | Influence of the XPD | 0.09 | U-shaped | 1.41 | 0.064 |
| 11 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 13 | Influence of TRP measurement grid (NOTE 1) | 0.32 | Actual | 1 | 0.32 |
| 14 | Influence of beam peak search grid (NOTE 2) | N/A | Actual | 1 | N/A |
| 15 | Multiple measurement antenna uncertainty (NOTE 5) | 0.15 | Actual | 1 | 0.15 |
| 16 | DUT repositioning | 0.00 | Rectangular | 1.73 | 0.00 |
| 17 | Misalignment of DUT due to change of DUT orientation | 0.10 | Actual | 1 | 0.10 |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 18 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 19 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 20 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 21 | Uncertainty of the Network Analyzer | 1.7 | Normal | 2.00 | 0.85 |
| 22 | Uncertainty of the absolute gain of the calibration antenna | 1.70 | Normal | 2.00 | 0.85 |
| 23 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.05 | Rectangular | 1.73 | 0.03 |
| 24 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 25 | Quality of quiet zone for calibration process (NOTE 4) | 0.60 | Actual | 1.00 | 0.60 |
| 26 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 27 | Influence of the calibration antenna feed cable | 0.28 | Normal | 2.00 | 0.14 |
| 28 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
|  | Expanded uncertainty (1.96σ - confidence interval of 95 %) |  |  |  | Value |
|  | TRP Expanded uncertainty ( 66 GHz < f <= 80 GHz) [dB] (a) |  |  |  | 7.31 |
|  | Systematic uncertainties (NOTE 3) |  |  |  | Value |
| 29 | Systematic error due to TRP calculation/quadrature (NOTE 1) (b) |  |  |  | 0.0 |
| 30 | Influence of noise ( 66 GHz < f <= 80 GHz) (c) |  |  |  | 0.64 |
| 31 | Systematic error related to beam peak search (NOTE 2) |  |  |  | N/A |
| Total measurement uncertainty |  |  |  |  | Value |
| Total measurement uncertainty (a)+(b)+(c) [dB] |  |  |  |  | 7.95 |
| NOTE 1: This contributor shall only be considered for TRP measurements. NOTE 2: This contributor shall only be considered for EIRP measurements. NOTE 3: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 4: Value based on procedure defined in clause D.2 of TR 38.810 for Quiet Zone size of less or equal to 30 cm. NOTE 5: Applies to the system which has a structure of mechanical feed antenna positioning. |  |  |  |  |  |

Table B.25.2-12: Uncertainty assessment for TRP measurement (f=6 GHz to 12.75 GHz, Quiet Zone size ≤ 30 cm) for PC1 UEs

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment | 0.02 | Normal | 2.00 | 0.01 |
| 2 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 3 | Quality of Quiet Zone (NOTE 4) | 0.70 | Actual | 1.00 | 0.70 |
| 4 | Mismatch | 1.60 | Actual | 1.00 | 1.60 |
| 5 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | Uncertainty of the RF power measurement equipment | 2.00 | Normal | 2.00 | 1.00 |
| 7 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 8 | Amplifier uncertainties | 2.1 | Normal | 2.00 | 1.05 |
| 9 | Random uncertainty | 0.5 | Normal | 2.00 | 0.25 |
| 10 | Influence of the XPD | 0.09 | U-shaped | 1.41 | 0.064 |
| 11 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 13 | Influence of TRP measurement grid (NOTE 1) | 0.32 | Actual | 1 | 0.32 |
| 14 | Influence of beam peak search grid (NOTE 2) | N/A | Actual | 1 | N/A |
| 15 | Multiple measurement antenna uncertainty (NOTE 5) | 0.15 | Actual | 1 | 0.15 |
| 16 | DUT repositioning | 0.00 | Rectangular | 1.73 | 0.00 |
| 17 | Misalignment of DUT due to change of DUT orientation | [0.50] | Actual | 1 | [0.50] |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 18 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 19 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 20 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 21 | Uncertainty of the Network Analyzer | 0.90 | Normal | 2.00 | 0.45 |
| 22 | Uncertainty of the absolute gain of the calibration antenna | 0.60 | Normal | 2.00 | 0.30 |
| 23 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.05 | Rectangular | 1.73 | 0.03 |
| 24 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 25 | Quality of quiet zone for calibration process (NOTE 4) | 0.70 | Actual | 1.00 | 0.70 |
| 26 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 27 | Influence of the calibration antenna feed cable | 0.14 | Normal | 2.00 | 0.07 |
| 28 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
|  | Expanded uncertainty (1.96σ - confidence interval of 95 %) |  |  |  | Value |
|  | TRP Expanded uncertainty (6 GHz < f <= 12.75 GHz) [dB] (a) |  |  |  | [4.96] |
|  | Systematic uncertainties (NOTE 3) |  |  |  | Value |
| 29 | Systematic error due to TRP calculation/quadrature (NOTE 1) (b) |  |  |  | 0.0 |
| 30 | Influence of noise (6 GHz < f <= 12.75 GHz) (c) |  |  |  | 0.64 |
| 31 | Systematic error related to beam peak search (NOTE 2) |  |  |  | N/A |
| Total measurement uncertainty |  |  |  |  | Value |
| Total measurement uncertainty (a)+(b)+(c) [dB] |  |  |  |  | [5.60] |
| NOTE 1: This contributor shall only be considered for TRP measurements. NOTE 2: This contributor shall only be considered for EIRP measurements. NOTE 3: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 4: Value based on procedure defined in clause D.2 of TR 38.810 for Quiet Zone size of less or equal to 30 cm. NOTE 5: Applies to the system which has a structure of mechanical feed antenna positioning. |  |  |  |  |  |

Table B.25.2-13: Uncertainty assessment for TRP measurement (f=12.75 GHz to 23.45 GHz, Quiet Zone size ≤ 30 cm) for PC1 UEs

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment | 0.02 | Normal | 2.00 | 0.01 |
| 2 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 3 | Quality of Quiet Zone (NOTE 4) | 0.60 | Actual | 1.00 | 0.60 |
| 4 | Mismatch | 1.60 | Actual | 1.00 | 1.60 |
| 5 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | Uncertainty of the RF power measurement equipment | 2.16 | Normal | 2.00 | 1.08 |
| 7 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 8 | Amplifier uncertainties | 2.1 | Normal | 2.00 | 1.05 |
| 9 | Random uncertainty | 0.5 | Normal | 2.00 | 0.25 |
| 10 | Influence of the XPD | 0.09 | U-shaped | 1.41 | 0.064 |
| 11 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 13 | Influence of TRP measurement grid (NOTE 1) | 0.32 | Actual | 1 | 0.32 |
| 14 | Influence of beam peak search grid (NOTE 2) | N/A | Actual | 1 | N/A |
| 15 | Multiple measurement antenna uncertainty (NOTE 5) | 0.15 | Actual | 1 | 0.15 |
| 16 | DUT repositioning | 0.00 | Rectangular | 1.73 | 0.00 |
| 17 | Misalignment of DUT due to change of DUT orientation | [0.5] | Actual | 1 | [0.5] |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 18 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 19 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 20 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 21 | Uncertainty of the Network Analyzer | 0.90 | Normal | 2.00 | 0.45 |
| 22 | Uncertainty of the absolute gain of the calibration antenna | 0.60 | Normal | 2.00 | 0.30 |
| 23 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.05 | Rectangular | 1.73 | 0.03 |
| 24 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 25 | Quality of quiet zone for calibration process (NOTE 4) | 0.60 | Actual | 1.00 | 0.60 |
| 26 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 27 | Influence of the calibration antenna feed cable | 0.14 | Normal | 2.00 | 0.07 |
| 28 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
|  | Expanded uncertainty (1.96σ - confidence interval of 95 %) |  |  |  | Value |
|  | TRP Expanded uncertainty (12.75 GHz < f <= 23.45 GHz) [dB] (a) |  |  |  | [4.92] |
|  | Systematic uncertainties (NOTE 3) |  |  |  | Value |
| 29 | Systematic error due to TRP calculation/quadrature (NOTE 1) (b) |  |  |  | 0.0 |
| 30 | Influence of noise (12.75 GHz < f <= 23.45 GHz) (c) |  |  |  | 0.64 |
| 31 | Systematic error related to beam peak search (NOTE 2) |  |  |  | N/A |
| Total measurement uncertainty |  |  |  |  | Value |
| Total measurement uncertainty (a)+(b)+(c) [dB] |  |  |  |  | [5.56] |
| NOTE 1: This contributor shall only be considered for TRP measurements. NOTE 2: This contributor shall only be considered for EIRP measurements. NOTE 3: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 4: Value based on procedure defined in clause D.2 of TR 38.810 for Quiet Zone size of less or equal to 30 cm. NOTE 5: Applies to the system which has a structure of mechanical feed antenna positioning. |  |  |  |  |  |

Table B.25.2-14: Uncertainty assessment for TRP measurement (f=23.45 GHz to 40.8 GHz, Quiet Zone size ≤ 30 cm) for PC1 UEs

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment | 0.02 | Normal | 2.00 | 0.01 |
| 2 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 3 | Quality of Quiet Zone (NOTE 4) | 0.6 | Actual | 1.00 | 0.6 |
| 4 | Mismatch | 1.50 | Actual | 1.00 | 1.50 |
| 5 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | Uncertainty of the RF power measurement equipment | 2.73 | Normal | 2.00 | 1.37 |
| 7 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 8 | Amplifier uncertainties | 2.1 | Normal | 2.00 | 1.05 |
| 9 | Random uncertainty | 0.5 | Normal | 2.00 | 0.25 |
| 10 | Influence of the XPD | 0.01 | U-shaped | 1.41 | 0.00 |
| 11 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 13 | Influence of TRP measurement grid (NOTE 1) | 0.32 | Actual | 1 | 0.32 |
| 14 | Influence of beam peak search grid (NOTE 2) | N/A | Actual | 1 | N/A |
| 15 | Multiple measurement antenna uncertainty (NOTE 5) | 0.15 | Actual | 1 | 0.15 |
| 16 | DUT repositioning | 0.00 | Rectangular | 1.73 | 0.00 |
| 17 | Misalignment of DUT due to change of DUT orientation | [0.5] | Actual | 1 | [0.5] |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 18 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 19 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 20 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 21 | Uncertainty of the Network Analyzer | 1.5 | Normal | 2.00 | 0.75 |
| 22 | Uncertainty of the absolute gain of the calibration antenna | 0.6 | Normal | 2.00 | 0.3 |
| 23 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.05 | Rectangular | 1.73 | 0.03 |
| 24 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 25 | Quality of quiet zone for calibration process (NOTE 4) | 0.6 | Actual | 1.00 | 0.6 |
| 26 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 27 | Influence of the calibration antenna feed cable | 0.14 | Normal | 2.00 | 0.07 |
| 28 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
|  | Expanded uncertainty (1.96σ - confidence interval of 95 %) |  |  |  | Value |
|  | TRP Expanded uncertainty (23.45 GHz < f <= 40.8 GHz) [dB] (a) |  |  |  | [5.20] |
|  | Systematic uncertainties (NOTE 3) |  |  |  | Value |
| 29 | Systematic error due to TRP calculation/quadrature (NOTE 1) (b) |  |  |  | 0.0 |
| 30 | Influence of noise (23.45 GHz < f <= 40.8 GHz) (c) |  |  |  | 1.0 |
| 31 | Systematic error related to beam peak search (NOTE 2) |  |  |  | N/A |
| Total measurement uncertainty |  |  |  |  | Value |
| Total measurement uncertainty (a)+(b)+(c) [dB] |  |  |  |  | [6.20] |
| NOTE 1: This contributor shall only be considered for TRP measurements. NOTE 2: This contributor shall only be considered for EIRP measurements. NOTE 3: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 4: Value based on procedure defined in clause D.2 of TR 38.810 for Quiet Zone size of less or equal to 30 cm. NOTE 5: Applies to the system which has a structure of mechanical feed antenna positioning. |  |  |  |  |  |

Table B.25.2-15: Uncertainty assessment for TRP measurement (f= 40.8 GHz to 66 GHz, Quiet Zone size ≤ 30 cm) for PC1 UEs

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment | 0.02 | Normal | 2.00 | 0.01 |
| 2 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 3 | Quality of Quiet Zone (NOTE 4) | 0.6 | Actual | 1.00 | 0.6 |
| 4 | Mismatch | 2.30 | Actual | 1.00 | 2.30 |
| 5 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | Uncertainty of the RF power measurement equipment | 4.0 | Normal | 2.00 | 2.00 |
| 7 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 8 | Amplifier uncertainties | 2.1 | Normal | 2.00 | 1.05 |
| 9 | Random uncertainty | 0.5 | Normal | 2.00 | 0.25 |
| 10 | Influence of the XPD | 0.09 | U-shaped | 1.41 | 0.064 |
| 11 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 13 | Influence of TRP measurement grid (NOTE 1) | 0.32 | Actual | 1 | 0.32 |
| 14 | Influence of beam peak search grid (NOTE 2) | N/A | Actual | 1 | N/A |
| 15 | Multiple measurement antenna uncertainty (NOTE 5) | 0.15 | Actual | 1 | 0.15 |
| 16 | DUT repositioning | 0.00 | Rectangular | 1.73 | 0.00 |
| 17 | Misalignment of DUT due to change of DUT orientation | [0.5] | Actual | 1 | [0.5] |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 18 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 19 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 20 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 21 | Uncertainty of the Network Analyzer | 1.7 | Normal | 2.00 | 0.85 |
| 22 | Uncertainty of the absolute gain of the calibration antenna | 1.70 | Normal | 2.00 | 0.85 |
| 23 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.05 | Rectangular | 1.73 | 0.03 |
| 24 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 25 | Quality of quiet zone for calibration process (NOTE 4) | 0.6 | Actual | 1.00 | 0.6 |
| 26 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 27 | Influence of the calibration antenna feed cable | 0.28 | Normal | 2.00 | 0.14 |
| 28 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
|  | Expanded uncertainty (1.96σ - confidence interval of 95 %) |  |  |  | Value |
|  | TRP Expanded uncertainty ( 40.8 GHz < f <= 66 GHz) [dB] (a) |  |  |  | [7.08] |
|  | Systematic uncertainties (NOTE 3) |  |  |  | Value |
| 29 | Systematic error due to TRP calculation/quadrature (NOTE 1) (b) |  |  |  | 0.0 |
| 30 | Influence of noise ( 40.8 GHz < f <= 66 GHz) (c) |  |  |  | 0.64 |
| 31 | Systematic error related to beam peak search (NOTE 2) |  |  |  | N/A |
| Total measurement uncertainty |  |  |  |  | Value |
| Total measurement uncertainty (a)+(b)+(c) [dB] |  |  |  |  | [7.72] |
| NOTE 1: This contributor shall only be considered for TRP measurements. NOTE 2: This contributor shall only be considered for EIRP measurements. NOTE 3: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 4: Value based on procedure defined in clause D.2 of TR 38.810 for Quiet Zone size of less or equal to 30 cm. NOTE 5: Applies to the system which has a structure of mechanical feed antenna positioning. |  |  |  |  |  |

Table B.25.2-16: Uncertainty assessment for TRP measurement (f= 66 GHz to 80 GHz, Quiet Zone size ≤ 30 cm) for PC1 UEs

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment | FFS | Normal | 2.00 | FFS |
| 2 | Measure distance uncertainty | FFS | Rectangular | 1.73 | FFS |
| 3 | Quality of Quiet Zone (NOTE 4) | FFS | Actual | 1.00 | FFS |
| 4 | Mismatch | FFS | Actual | 1.00 | FFS |
| 5 | Standing wave between the DUT and measurement antenna | FFS | U-shaped | 1.41 | FFS |
| 6 | Uncertainty of the RF power measurement equipment | FFS | Normal | 2.00 | FFS |
| 7 | Phase curvature | FFS | U-shaped | 1.41 | FFS |
| 8 | Amplifier uncertainties | FFS | Normal | 2.00 | FFS |
| 9 | Random uncertainty | FFS | Normal | 2.00 | FFS |
| 10 | Influence of the XPD | FFS | U-shaped | 1.41 | FFS |
| 11 | Insertion Loss Variation | FFS | Rectangular | 1.73 | FFS |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | FFS | Actual | 1.00 | FFS |
| 13 | Influence of TRP measurement grid (NOTE 1) | FFS | Actual | 1 | FFS |
| 14 | Influence of beam peak search grid (NOTE 2) | N/A | Actual | 1 | N/A |
| 15 | Multiple measurement antenna uncertainty (NOTE 5) | FFS | Actual | 1 | FFS |
| 16 | DUT repositioning | FFS | Rectangular | 1.73 | FFS |
| 17 | Misalignment of DUT due to change of DUT orientation | FFS | Actual | 1 | FFS |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 18 | Mismatch | FFS | U-shaped | 1.41 | FFS |
| 19 | Amplifier Uncertainties | FFS | Normal | 2.00 | FFS |
| 20 | Misalignment of positioning System | FFS | Normal | 2.00 | FFS |
| 21 | Uncertainty of the Network Analyzer | FFS | Normal | 2.00 | FFS |
| 22 | Uncertainty of the absolute gain of the calibration antenna | FFS | Normal | 2.00 | FFS |
| 23 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | FFS | Rectangular | 1.73 | FFS |
| 24 | Phase centre offset of calibration antenna | FFS | Rectangular | 1.73 | FFS |
| 25 | Quality of quiet zone for calibration process (NOTE 4) | FFS | Actual | 1.00 | FFS |
| 26 | Standing wave between reference calibration antenna and measurement antenna | FFS | U-shaped | 1.41 | FFS |
| 27 | Influence of the calibration antenna feed cable | FFS | Normal | 2.00 | FFS |
| 28 | Insertion Loss Variation | FFS | Rectangular | 1.73 | FFS |
|  | Expanded uncertainty (1.96σ - confidence interval of 95 %) |  |  |  | Value |
|  | TRP Expanded uncertainty ( 66 GHz < f <= 80 GHz) [dB] (a) |  |  |  | FFS |
|  | Systematic uncertainties (NOTE 3) |  |  |  | Value |
| 29 | Systematic error due to TRP calculation/quadrature (NOTE 1) (b) |  |  |  | FFS |
| 30 | Influence of noise ( 66 GHz < f <= 80 GHz) (c) |  |  |  | FFS |
| 31 | Systematic error related to beam peak search (NOTE 2) |  |  |  | N/A |
| Total measurement uncertainty |  |  |  |  | Value |
| Total measurement uncertainty (a)+(b)+(c) [dB] |  |  |  |  | FFS |
| NOTE 1: This contributor shall only be considered for TRP measurements. NOTE 2: This contributor shall only be considered for EIRP measurements. NOTE 3: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 4: Value based on procedure defined in clause D.2 of TR 38.810 for Quiet Zone size of less or equal to 30 cm. NOTE 5: Applies to the system which has a structure of mechanical feed antenna positioning. |  |  |  |  |  |

<<End of change>>