3GPP TSG RAN WG5 Meeting #97-e 	R5-227229r1

Toulouse, France, 14th – 18th November 2022

Source:		Rohde & Schwarz

Title:	Discussion on QoQZ and XPD MUs for 40cm

Agenda Item:		5.4.17

Document for:	Discussion and Endorsement

# Introduction

Following the discussion during last meetings regarding 40cm QZ size and the impact on MU, we present here our updated results and proposals.

# QoQZ measurement results

As presented in [3], the results provided here are based on an EIRP and TRP QoQZ measurement campaign with 238 measurements points (assuming re-positioning approach) from the combination of reference points (P1 to P7), reference antenna orientations (and) and reference antenna polarization (pol). No ETC enclosure was used for this measurement campaign. The reference points are as defined in clause O.2.5 of TS 38.521-2 [1] with the following coordinates:

Table 2-1: AUT reference points coordinates for 40cm QZ

| Position | x | y | z |
| --- | --- | --- | --- |
| P1 | 0 | 0 | 0 |
| P2 | 20 cm | 0 | 0 |
| P3 | - 20 cm | 0 | 0 |
| P4 | 0 | 20 cm | 0 |
| P5 | 0 | - 20 cm | 0 |
| P6 | 0 | 0 | 20 cm |
| P7 | 0 | 0 | - 15 cm |

The following table show the EIRP and TRP standard deviation calculated from full campaign measurement results for 40cm QZ (i.e. 238 points, P1 to P7, all antenna orientations, 2 polarizations):

Table 2-2: EIRP QoQZ MU for 40cm QZ

|  |  |  |  |  | EIRP Std Deviation [dB] / Frequency |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Set of points | # meas. | γpol [º] | α [º] | β [º] | 23.45 GHz | 32.125 GHz | 40.8 GHz | 44.3 GHz |
| P1 to P7 | 238 | all | {0} | {0,45,90,270,315} | 0.56 | 0.68 | 0.64 | 0.74 |
| P1 | 34 | all | {0} | {0,45,90,270,315} | 0.16 | 0.13 | 0.29 | 0.22 |

Table 2-3: TRP QoQZ MU for 40cm QZ

|  |  |  |  |  | TRP Std Deviation [dB] / Frequency |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Set of points | # meas. | γpol [º] | α [º] | β [º] | 23.45 GHz | 32.125 GHz | 40.8 GHz | 44.3 GHz |
| P1 to P7 | 238 | all | {0} | {0,45,90,270,315} | 0.48 | 0.61 | 0.54 | 0.69 |
| P1 | 34 | all | {0} | {0,45,90,270,315} | 0.18 | 0.10 | 0.24 | 0.25 |

Following the discussion and comments received during RAN5#96-e meeting, several test points were reassessed to ensure the consistency of results. As it can be seen from the comparison between the results for a 30cm QZ chamber [3] and the upgrade to 40cm, there is a degradation in the QoQZ MU as anticipated in [8].

Observation 1: the QZ extension of current test systems from 30cm to 40cm introduce degradation in QoQZ.

In Table 2-4, a summarized view of the QoQZ MU target value per frequency used in TR 38.903 [2], with the additional value for 44.3GHz corresponding to n259 as endorsed in [7] is presented:

Table 2-4: EIRP and TRP QoQZ MU target value for 30cm QZ

|  | Std Deviation [dB] / Frequency |  |  |  |
| --- | --- | --- | --- | --- |
| Stage | 23.45 GHz | 32.125 GHz | 40.8 GHz | 44.3 GHz |
| Stage 2 (P1 to P7) | 0.60 | 0.60 | 0.60 | 0.70 |
| Stage 1 (P1) | 0.40 | 0.40 | 0.40 | 0.50 |

Following the results presented above, taking the worst case found for 32.125GHz and 44.3GHz as a reference, the values in Table 2-5 are proposed as QoQZ MU for EIRP and TRP for 40cm QZ:

Table 2-5: Proposed EIRP and TRP QoQZ MU target value for 40cm QZ

|  | Std Deviation [dB] / Frequency |  |  |  |
| --- | --- | --- | --- | --- |
| Stage | 23.45 GHz | 32.125 GHz | 40.8 GHz | 44.3 GHz |
| Stage 2 (P1 to P7) | 0.70 | 0.70 | 0.70 | 0.80 |
| Stage 1 (P1) | 0.40 | 0.40 | 0.40 | 0.50 |

Proposal 1: Capture in TR 38.903 a Stage 2 QoQZ MU for 40cm QZ- of 0.70dB for frequencies ≤ 40.8GHz and 0.8dB for 44.3GHz for NTC. Other frequencies and ETC are FFS.

# XPD measurement results

In previous analyses [6][8], but also confirmed with measurements in [4][5], a degradation of the cross polar discrimination (XPD) is also expected on systems upgraded to 40cm QZ.

Table 3-1 presents a summarized view of the Influence of XPD MU target value per frequency used in TR 38.903 [2],

Table 3-1: Influence of XPD MU target value for 30cm QZ

| Influence of XPD MU [dB] / Frequency |  |  |  |
| --- | --- | --- | --- |
| 23.45 GHz | 32.125 GHz | 40.8 GHz | 44.3 GHz |
| 0.01 | 0.01 | 0.01 | Not defined |

Using the results from the QoQZ measurement campaign for described in previous section, a similar degradation is observed for 40cm:

Table 3-2: Proposed Influence of XPD MU target value for 40cm QZ

| Influence of XPD MU [dB] / Frequency |  |  |  |
| --- | --- | --- | --- |
| 23.45 GHz | 32.125 GHz | 40.8 GHz | 44.3 GHz |
| 0.03 | 0.03 | 0.03 | 0.03 |

The analysis provided in [5] showed a maximum Influence of XPD MU of 0.07dB among frequencies below 44.3GHz, and therefore it is proposed to capture this value in TR 38.903.

Proposal 2: Capture in TR 38.903 an Influence of XPD MU for 40cm QZ- of 0.07dB for frequencies ≤ 44.3GHz. Other frequencies are FFS.

# Impact on overall MTSU

As presented before in [3][6], the impact on the overall MU for 40cm QZ, taking MOP and REFSENS for PC3 ≤40.8GHz as example, is evaluated here for information using the proposed values above.

Table 4-1: Summary example of impact on EIRP MTSU for 40cm QZ

|  | Expanded Uncertainty (1.96σ - confidence interval of 95 %) [dB] / Frequency |  |
| --- | --- | --- |
|  | 23.45GHz <= f <= 32.125GHz | 32.125GHz < f <= 40.8GHz |
| 30cm QZ, PC3 UEs, NTC | 4.89 | 5.09 |
| 40cm QZ, PC3 UEs, NTC | 4.95 | 5.15 |
| Increase 30cm to 40cm QZ | 0.06 | 0.06 |

Table 4-2: Summary example of impact on TRP MTSU for 40cm QZ

|  | Expanded Uncertainty (1.96σ - confidence interval of 95 %) [dB] / Frequency |  |
| --- | --- | --- |
|  | 23.45GHz <= f <= 32.125GHz | 32.125GHz < f <= 40.8GHz |
| 30cm QZ, PC3 UEs, NTC | 4.42 | 4.62 |
| 40cm QZ, PC3 UEs, NTC | 4.48 | 4.62 |
| Increase 30cm to 40cm QZ | 0.06 | 0.06 |

Table 4-3: Summary example of impact on EIS MTSU for 40cm QZ

|  | Expanded Uncertainty (1.96σ - confidence interval of 95 %) [dB] / Frequency |
| --- | --- |
|  | 23.45GHz <= f <= 40.8GHz |
| 30cm QZ, PC3 UEs, NTC | 5.19 |
| 40cm QZ, PC3 UEs, NTC | 5.24 |
| Increase 30cm to 40cm QZ | 0.05 |

Following the discussion and proposals made during RAN5#96-e meeting, and according to the updated results presented in this contribution, we share the opinion that the potential increase in MTSU between 30cm and 40cm QZ is not enough to justify the burden to create an maintain separate MTSU, TT and test requirements for 40cm QZ. Therefore, we propose to keep the same MTSU for 40cm QZ, while capturing the degradation as reflected in Proposals 1 and 2.

Proposal 3: Define the MTSU (and correspondingly the TT) for the 40 cm QZ to be the same as for the 30 cm QZ for FR2a, FR2b, and FR2c. Other frequencies and ETC are FFS.

# Conclusion

In this contribution, the following observations and proposals were made:

Observation 1: the QZ extension of current test systems from 30cm to 40cm introduce degradation in QoQZ.

Proposal 1: Capture in TR 38.903 a Stage 2 QoQZ MU for 40cm QZ of 0.70dB for frequencies ≤ 40.8GHz and 0.8dB for 44.3GHz for NTC. Other frequencies and ETC are FFS.

Proposal 2: Capture in TR 38.903 an Influence of XPD MU for 40cm QZ of 0.07dB for frequencies ≤ 44.3GHz. Other frequencies are FFS.

Proposal 3: Define the MTSU (and correspondingly the TT) for the 40 cm QZ to be the same as for the 30 cm QZ for FR2a, FR2b, and FR2c. Other frequencies and ETC are FFS.

-

-

A CR is provided in [12] to implement Proposals 1, 2 and 3 in TR 38.903.

# References

- 3GPP TS 38.521-2, v17.0.0 (2022-10)

- 3GPP TR 38.903, v16.13.0 (2022-09)

- R5-225646, “Discussion on QoQZ and XPD MUs for 40cm”, Rohde & Schwarz, 3GPP TSG RAN WG5 Meeting #96-e, August 2022

- R5-223613, “MU discussion on 40 cm Quiet Zone”, Anritsu, 3GPP TSG RAN WG5 Meeting #95-e, May 2022

- R5-223614, “On QoQZ for 40cm QZ”, Keysight Technologies, 3GPP TSG RAN WG5 Meeting #95-e, May 2022

- R5-213193, “Discussion on the size of Quiet Zone above 30cm”, Rohde & Schwarz, 3GPP TSG RAN WG5 Meeting #90-e, May 2021

- R5-211656, “On the MU for n259”, Rohde & Schwarz, 3GPP TSG RAN WG5 Meeting #90-e, February 2021

- R5-211235, “Discussion on the size of Quiet Zone above 30cm”, Rohde & Schwarz, 3GPP TSG RAN WG5 Meeting #90-e, February 2021

- R5-206130, “Discussion on size of Quiet Zone above 30cm”, Rohde & Schwarz, 3GPP TSG RAN WG5 Meeting #89-e, November 2020

- R5-204200, “Discussion on size of Quiet Zone above 30cm”, Rohde & Schwarz, 3GPP TSG RAN WG5 Meeting #88-e, August 2020

- R5-202082, “Discussion on size of Quiet Zone above 30cm”, Rohde & Schwarz, 3GPP TSG RAN WG5 Meeting #87-e, May 2020

- R5-227230, “40cm QoQZ and XPD MU in 38.903”, Rohde & Schwarz, 3GPP TSG RAN WG5 Meeting #97, November 2022