3GPP TSG RAN WG5 Meeting #100 	R5-2315155r1

Toulouse, France, 21 – 25 August 2023

Source:		Bluetest

Title:	RC MU Analysis for NR FR1 TRP-TRS Enhancement (Rel-18)

Agenda Item:		5.3.26.2 Discussion Papers (Measurement Uncertainty (MU) assessment proposals for TR 38.870)

Document for:	Discussion and Endorsement

# Introduction

## Continuing the work on the RAN5 MU Work Plan for NR FR1 TRP-TRS Enhancement (Rel-18) [1] and the earlier document R5-233679 from 3GPP TSG RAN WG5 Meeting #99 [2], this document proposes a reformatting of the MU based on feedback from delegates. Specifically, to address Talk mode and Browser mode versions of TRP and TRS.

## Approach

## This draft is based primarily on the existing RC method MU tables for TRP and TRS such as Annex A in TR 25.914 [2], particularly Tables A.3a, A.3b, A.3c, A.7, and A.8. Similar tables can be found in [3] and are considered for this draft. In addition, as many terms as possible are aligned with the ongoing work on the reference method AC, in particular Tables B.4.1-1 and B.5.1-1 from [3].

## Common MU Definitions

## Some of the terms for measurement uncertainty in the RC method are identical to the reference method AC. Following recent updates to the AC reference approach to MU terms and concerns [3], we add and update the MU tables for the RC approach to align with the AC approach.

## RC Specific MU Definitions

## In this draft, we align the definitions of MU from the AC approach from [4] with those from the RC approach. And provide initial tables for MU consideration of the TRP and TRS in the reverberation chamber.

# References

- R5-225924, “MU workplan for NR FR1 TRP-TRS Enhancement (Rel-18)”, Rhode & Schwarz, vivo, 3GPP TSG RAN5 Meeting #97-e, November 2022

- R5-233679, “RC MU Analysis for NR FR1 TRP-TRS Enhancement (Rel-18)”, Bluetest AB, Electronic meeting, 22-26 May 2023.

- 3GPP TR 38.810 V16.6.0 (2020-10)

- 3GPP TR 38.870 V0.40.0 (2023-05)

# Conclusion

The text submission should be considered for the development of the MU sections in TR 37.870

Proposal 1: Endorse the text proposal below

# Appendix - Text Proposal to TR 38.870

< Unchanged Text Deleted >

< Beginning of Changes >

# B.3	MU contribution descriptions for Alternative method

Some of the terms for measurement uncertainty in the Reverberation Chamber method are identical to the reference method (i.e. Anechoic Chamber). Therefore, this section only lists the additional MU contribution description that are unique to the Reverberation Chamber method.

## B.3.1	Additional Power Loss in EUT Chassis

When the EUT is small and does not add noticeable loss to the chamber, the calibration procedure outlined in section 8.4, is performed without the EUT present in the chamber. The possible difference in average chamber transmission level between the EUT measurement and the reference measurement must in this case be considered in the uncertainty evaluation. The uncertainty value for this contribution can be tested empirically by choosing a unit within a set of samples which is considered to incur the highest amount of loss (normally the largest unit), and measure the average transmission loss in the chamber with and without the test unit present in the chamber including phantoms if applicable. The difference between the two cases shall be used in the uncertainty calculation and the distribution should be assumed to be rectangular.

Alternatively, a fixed value of 0.2 dB with a rectangular distribution can be used in the uncertainty calculations.

## B.3.3   Quality of Spatial Uniformity

The standard deviation over calibration antenna positions and rotations of the transfer function, chamber reference or chamber loss S212. This measures the lack of spatial uniformity over the test zone in the reverberation chamber which is also referred as lack of isotropy, statistical ripple, or uniformity of the transfer function. This procedure is detailed in section 8.3.3.

According to the significance tests performed in [23], the uncertainty due to lack of spatial uniformity is not negligible. In this case, equation 12 in [23] holds and leads to the following uncertainty:

U=σPrefdBT

The expanded uncertainty covers the 95th percentile and is given by

U95 = K95 *U

Where K95 is the coverage factor for the two-sided 95th percentile of Student’s t-distribution given T-1 degrees of freedom and corresponds to 2.201 at T=12 (11 degrees of freedom) or 2.069 at T=24 (23 degrees of freedom).

# B.4	MU Assessment for TRP

## B.4.1	MU Assessment for TRP in Anechoic Chamber

< Unchanged Text Deleted >

< Beginning of Changes >

### B.4.2	MU Assessment for TRP in Reverberation Chamber

Table B.4.2-1 Uncertainty contributions in TRP measurement for reverberation chamber method

| UID | Description of uncertainty contribution | Details in clause |
| --- | --- | --- |
| Stage 2: DUT measurement |  |  |
| 1 | Mismatch of receiver chain | B.2.1 |
| 2 | Insertion loss of receiver chain | B.2.2 |
| 3 | Influence of the measurement antenna cable | B.2.3 |
| 4 | Measurement Receiver: uncertainty of the absolute level | B.2.4 |
| 5 | Quality of Spatial Uniformity | B.3.3 |
| 6 | Additional power loss in EUT chassis | B.3.1 |
| 7 | DUT Tx-power drift | B.2.9 |
| 8 | Uncertainty related to the use of phantoms | B.2.11 |
| 9 | Random uncertainty | B.2.13 |
| Stage 1: Calibration measurement |  |  |
| 10 | Uncertainty of network analyzer | B.2.15 |
| 11 | Mismatch of receiver chain | B.2.1 |
| 12 | Insertion loss of receiver chain | B.2.2 |
| 13 | Mismatch in the connection of calibration antenna | B.2.1 |
| 14 | Influence of the calibration antenna feed cable | B.2.3 |
| 15 | Influence of the measurement antenna cable | B.2.3 |
| 16 | Uncertainty of the absolute gain/ radiation efficiency of the calibration antenna | B.2.16 |
| 17 | Uniformity of transfer function | B.3.3 |

Table B.4.2-2 Preliminary example of uncertainty budget for TRP measurement hand only (browsing mode) for reverberation chamber method for NR FR1 bands

| UID | Uncertainty Source | Comment | Uncertainty Value [dB] | Prob Distr | Div | ci | Standard Uncertainty [dB] |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |
| 1 | Mismatch of receiver chain | Гpower meter <0.05 Гfixed measurement antenna connection <0.16 | 0.07 | U-shaped | 1.41 | 1 | [0.05] |
| 2 | Insertion loss of receiver chain | Systematic with Stage 2 (=> cancels) | 0 | Rectangular | 1.73 | 1 | 0.00 |
| 3 | Influence of the fixed measurement antenna cable | Systematic with Stage 2 (=> cancels) | 0 | Rectangular | 1.73 | 1 | 0.00 |
| 4 | Measurement Receiver: uncertainty of the absolute level | Power Meter | 0.06 | Rectangular | 1.73 | 1 | [0.03] |
| 5 | Quality of Spatial Uniformity | Statistics of chamber | 0.5 | Gaussian | 1 | 1 | [0.5] |
| 6 | Additional power loss in EUT chassis | The EUT not present in the chamber during calibration measurement | 0.1 | Rectangular | 1.73 | 1 | [0.06] |
| 7 | DUT Tx-power drift | Drift | 0.2 | Rectangular | 1.73 | 1 | [0.12] |
| 8 | Uncertainty related to the use of phantoms | U [dB] = 0.20 U [dB] = 0.15 | 0.32 | Rectangular | 1.73 | 1 | [0.18] |
| 9 | Repeatability | Using the same setup and stirring sequence | 0.4 | Rectangular | 1.73 | 1 | [0.23] |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |
| 10 | Uncertainty of network analyzer | Manufacturer’s uncertainty calculator, covers whole NA setup | 0.5 | Rectangular | 1.73 | 1 | [0.29] |
| 11 | Mismatch of receiver chain | Taken in to account in NA setup uncertainty | 0 | U-shaped | 1.41 | 1 | 0.00 |
| 12 | Insertion loss of receiver chain | Systematic with Stage 1 (=> cancels) | 0 | Rectangular | 1.73 | 1 | 0.00 |
| 13 | Mismatch in the connection of calibration antenna | Taken in to account in NA setup uncertainty | 0 | U-shaped | 1.41 | 1 | 0.00 |
| 14 | Influence of the feed cable of the calibration antenna | Gain calibration with a dipole | 0.3 | Rectangular | 1.73 | 1 | [0.17] |
| 15 | Influence of the measurement antenna cable | Systematic with Stage 1 (=> cancels) | 0 | Rectangular | 1.73 | 1 | 0.00 |
| 16 | Uncertainty of the absolute gain/ radiation efficiency of the calibration antenna | Calibration certificate | 0.5 | Rectangular | 1.73 | 1 | [0.29] |
| 17 | Uniformity of transfer function | Statistics of chamber | Gaussian | 1.00 | 1 | 1 | [0.5] |
| Combined standard uncertainty |  |  |  |  |  |  | [0.74] |
| Expanded uncertainty (Confidence interval of 95 %) |  |  |  |  |  |  | [1.45] |

Table B.4.2-2 Preliminary example of uncertainty budget for TRP Beside Head and Hand (Talk mode) measurement for reverberation chamber method for NR FR1 bands

| UID | Uncertainty Source | Comment | Uncertainty Value [dB] | Prob Distr | Div | ci | Standard Uncertainty [dB] |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |
| 1 | Mismatch of receiver chain | Гpower meter <0.05 Гfixed measurement antenna connection <0.16 | 0.07 | U-shaped | 1.41 | 1 | [0.05] |
| 2 | Insertion loss of receiver chain | Systematic with Stage 2 (=> cancels) | 0 | Rectangular | 1.73 | 1 | 0.00 |
| 3 | Influence of the fixed measurement antenna cable | Systematic with Stage 2 (=> cancels) | 0 | Rectangular | 1.73 | 1 | 0.00 |
| 4 | Measurement Receiver: uncertainty of the absolute level | Power Meter | 0.06 | Rectangular | 1.73 | 1 | [0.03] |
| 5 | Quality of Spatial Uniformity | Statistics of chamber | 0.5 | Gaussian | 1 | 1 | [0.5] |
| 6 | Additional power loss in EUT chassis | The EUT not present in the chamber during calibration measurement | 0.1 | Rectangular | 1.73 | 1 | [0.06] |
| 7 | DUT Tx-power drift | Drift | 0.2 | Rectangular | 1.73 | 1 | [0.12] |
| 8 | Uncertainty related to the use of phantoms | U [dB] = 0.20 U [dB] = 0.15 | 0.32 | Rectangular | 1.73 | 1 | [0.18] |
| 9 | Repeatability | Using the same setup and stirring sequence | 0.4 | Rectangular | 1.73 | 1 | [0.23] |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |
| 10 | Uncertainty of network analyzer | Manufacturer’s uncertainty calculator, covers whole NA setup | 0.5 | Rectangular | 1.73 | 1 | [0.29] |
| 11 | Mismatch of receiver chain | Taken in to account in NA setup uncertainty | 0 | U-shaped | 1.41 | 1 | 0.00 |
| 12 | Insertion loss of receiver chain | Systematic with Stage 1 (=> cancels) | 0 | Rectangular | 1.73 | 1 | 0.00 |
| 13 | Mismatch in the connection of calibration antenna | Taken in to account in NA setup uncertainty | 0 | U-shaped | 1.41 | 1 | 0.00 |
| 14 | Influence of the feed cable of the calibration antenna | Gain calibration with a dipole | 0.3 | Rectangular | 1.73 | 1 | [0.17] |
| 15 | Influence of the measurement antenna cable | Systematic with Stage 1 (=> cancels) | 0 | Rectangular | 1.73 | 1 | 0.00 |
| 16 | Uncertainty of the absolute gain/ radiation efficiency of the calibration antenna | Calibration certificate | 0.5 | Rectangular | 1.73 | 1 | [0.29] |
| 17 | Uniformity of transfer function | Statistics of chamber | Gaussian | 1.00 | 1 | 1 | [0.5] |
| Combined standard uncertainty |  |  |  |  |  |  | [0.74] |
| Expanded uncertainty (Confidence interval of 95 %) |  |  |  |  |  |  | [1.45] |

# B.5	MU Assessment for TRS

## B.5.1	MU Assessment for TRS in Anechoic Chamber

< Unchanged Text Deleted >

< Beginning of Changes >

### B.5.2	MU Assessment for TRS in Reverberation Chamber

Table B.5.2-1 Uncertainty contributions in TRS measurement for reverberation chamber method

| UID | Description of uncertainty contribution | Details in clause |
| --- | --- | --- |
| Stage 2: DUT measurement |  |  |
| 1 | Mismatch of transmitter chain | B.2.1 |
| 2 | Insertion loss of transmitter chain | B.2.2 |
| 3 | Influence of the measurement antenna cable | B.2.3 |
| 4 | Communication Tester: uncertainty of the absolute output level | B.2.5 |
| 5 | Sensitivity measurement: output level step resolution | B.2.6 |
| 6 | Statistical uncertainty of BER measurement | B.5.4 |
| 7 | BER data rate normalizations | B.5.5 |
| 8 | DUT sensitivity drift | B.2.10 |
| 9 | Quality of Spatial Uniformity | B.5.3 |
| 10 | Uncertainty related to the use of phantoms | B.2.11 |
| 11 | Random uncertainty | B.2.13 |
| Stage 1: Calibration measurement |  |  |
| 12 | Uncertainty of network analyzer | B.2.15 |
| 13 | Mismatch of transmitter chain | B.2.1 |
| 14 | Insertion loss of transmitter chain | B.2.2 |
| 15 | Mismatch in the connection of calibration antenna | B.2.1 |
| 16 | Influence of the calibration antenna feed cable | B.2.3 |
| 17 | Influence of the measurement antenna cable | B.2.3 |
| 18 | Uncertainty of the absolute gain/radiation efficiency of the calibration antenna | B.2.16 |
| 19 | Quality of Spatial Uniformity | B.5.3 |

Table B.5.2-2: Preliminary example of uncertainty budget for TRS hand only (browsing mode) measurement for reverberation chamber method for NR FR1 bands

| UID | Uncertainty Source | Comment | Uncertainty Value [dB] | Prob Distr | Div | ci | Standard Uncertainty [dB] |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |
| 1 | Mismatch of transmitter chain | ГBSS <0.13 Г antenna connection <0.03 | 0.02 | Gaussian | 1 | 1 | [0.02] |
| 2 | Insertion loss of transmitter chain | Systematic with Stage 1 (=> cancels) | 0 | Rectangular | 1.73 | 1 | 0 |
| 3 | Influence of the fixed measurement antenna cable | Systematic with Stage 2 (=> cancels) | 0 | Rectangular | 1.73 | 1 | 0 |
| 5 | Base station simulator: uncertainty of the absolute level |  | 1 | Rectangular | 1.73 | 1 | [0.58] |
| 6 | BER measurement: output level step resolution | Step 0.1dB | 0.05 | Rectangular | 1.73 | 1 | [0.03] |
| 7 | Statistical uncertainty of the BER measurement | BER target 10%±2% , 20000 tested bits , N=60 | 0.12 | Gaussian | 1 | 1 | [0.12] |
| 8 | TRS data rate normalization | 4 reference points measured | 0.12 | Gaussian | 1 | 1 | [0.12] |
| 9 | Quality of Spatial Uniformity | Statistics of chamber | 0.5 | Gaussian | 1 | 1 | [0.5] |
| 10 | Additional power loss in EUT chassis | The EUT not present in the chamber during calibration measurement | 0.1 | Rectangular | 1.73 | 1 | [0.06] |
| 11 | DUT sensitivity drift | Drift measurement | 0.2 | Rectangular | 1.73 | 1 | [0.12] |
| 12 | Uncertainty related to the use of phantoms | Standard SAM with standard tissue simulant | 0 | Rectangular | 1.73 | 1 | 0 |
| 13 | Repeatability | Using the same setup and stirring sequence | 0.4 | Rectangular | 1.73 | 1 | [0.23] |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |
| 14 | Uncertainty of network analyzer | Manufacturer’s uncertainty calculator, covers NA setup | 0.5 | Rectangular | 1.73 | 1 | [0.29] |
| 15 | Mismatch of transmitter chain | Taken in to account in NA setup uncertainty | 0 | U-shaped | 1.41 | 1 | 0 |
| 16 | Insertion loss of transmitter chain | Systematic with Stage 1 (=> cancels) | 0 | Rectangular | 1.73 | 1 | 0 |
| 17 | Mismatch in the connection of calibration antenna | Taken in to account in NA setup uncertainty | 0 | Rectangular | 1.73 | 1 | 0 |
| 18 | Influence of the feed cable of the calibration antenna | Gain calibration with dipole | 0.3 | Rectangular | 1.73 | 1 | [0.17] |
| 19 | Influence of the fixed measurement antenna cable | Systematic with Stage 1 (=> cancels) | 0 | Rectangular | 1.73 | 1 | 0 |
| 20 | Uncertainty of the absolute gain of the calibration antenna | Calibration certificate | 0.5 | Rectangular | 1.73 | 1 | [0.29] |
| 21 | Quality of Spatial Uniformity | Statistics of chamber | 0.5 | Gaussian | 1 | 1 | [0.5] |
| Combined standard uncertainty |  |  |  |  |  |  | [1.02] |
| Expanded uncertainty (Confidence interval of 95 %) |  |  |  |  |  |  | [2.00] |

Table B.5.2-2: Preliminary example of uncertainty budget for TRS Beside Head and Hand (Talk mode) measurement for reverberation chamber method for NR FR1 bands

| UID | Uncertainty Source | Comment | Uncertainty Value [dB] | Prob Distr | Div | ci | Standard Uncertainty [dB] |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |
| 1 | Mismatch of transmitter chain | ГBSS <0.13 Г antenna connection <0.03 | 0.02 | Gaussian | 1 | 1 | [0.02] |
| 2 | Insertion loss of transmitter chain | Systematic with Stage 1 (=> cancels) | 0 | Rectangular | 1.73 | 1 | 0 |
| 3 | Influence of the fixed measurement antenna cable | Systematic with Stage 2 (=> cancels) | 0 | Rectangular | 1.73 | 1 | 0 |
| 5 | Base station simulator: uncertainty of the absolute level |  | 1 | Rectangular | 1.73 | 1 | [0.58] |
| 6 | BER measurement: output level step resolution | Step 0.1dB | 0.05 | Rectangular | 1.73 | 1 | [0.03] |
| 7 | Statistical uncertainty of the BER measurement | BER target 10%±2% , 20000 tested bits , N=60 | 0.12 | Gaussian | 1 | 1 | [0.12] |
| 8 | TRS data rate normalization | 4 reference points measured | 0.12 | Gaussian | 1 | 1 | [0.12] |
| 9 | Quality of Spatial Uniformity | Statistics of chamber | 0.5 | Gaussian | 1 | 1 | [0.5] |
| 10 | Additional power loss in EUT chassis | The EUT not present in the chamber during calibration measurement | 0.1 | Rectangular | 1.73 | 1 | [0.06] |
| 11 | DUT sensitivity drift | Drift measurement | 0.2 | Rectangular | 1.73 | 1 | [0.12] |
| 12 | Uncertainty related to the use of phantoms | Standard SAM with standard tissue simulant | 0 | Rectangular | 1.73 | 1 | 0 |
| 13 | Repeatability | Using the same setup and stirring sequence | 0.4 | Rectangular | 1.73 | 1 | [0.23] |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |
| 14 | Uncertainty of network analyzer | Manufacturer’s uncertainty calculator, covers NA setup | 0.5 | Rectangular | 1.73 | 1 | [0.29] |
| 15 | Mismatch of transmitter chain | Taken in to account in NA setup uncertainty | 0 | U-shaped | 1.41 | 1 | 0 |
| 16 | Insertion loss of transmitter chain | Systematic with Stage 1 (=> cancels) | 0 | Rectangular | 1.73 | 1 | 0 |
| 17 | Mismatch in the connection of calibration antenna | Taken in to account in NA setup uncertainty | 0 | Rectangular | 1.73 | 1 | 0 |
| 18 | Influence of the feed cable of the calibration antenna | Gain calibration with dipole | 0.3 | Rectangular | 1.73 | 1 | [0.17] |
| 19 | Influence of the fixed measurement antenna cable | Systematic with Stage 1 (=> cancels) | 0 | Rectangular | 1.73 | 1 | 0 |
| 20 | Uncertainty of the absolute gain of the calibration antenna | Calibration certificate | 0.5 | Rectangular | 1.73 | 1 | [0.29] |
| 21 | Quality of Spatial Uniformity | Statistics of chamber | 0.5 | Gaussian | 1 | 1 | [0.5] |
| Combined standard uncertainty |  |  |  |  |  |  | [1.02] |
| Expanded uncertainty (Confidence interval of 95 %) |  |  |  |  |  |  | [2.00] |

< End of Changes >