3GPP TSG RAN WG5 Meeting #101 	R5-237211

Chicago, USA, 13 – 17 November 2023

Source:				Rohde & Schwarz

Title:					TP to TR 38.870 Annex B on forearm phantom and other MU changes

Agenda Item:		5.3.22.2

Document for:	Discussion and Endorsement

# Introduction

The work item led by RAN4 for NR FR1 TRP-TRS enhancements was approved [1] during the September 2022 plenary meeting, revised [2] in December 2022, and includes a secondary responsibility for RAN5 to develop the Measurement Uncertainty (MU) assessment.

The aim of this contribution is to introduce changes to the preliminary Measurement Uncertainty tables for Annex B of TR 38.870 [3] to reflect the updates based on the analysis made for TS 38.561 [4], introducing the contents for forearm phantom and other editorial corrections.

It is understood that RAN5 will finalize the MU estimation following the performance requirements definition in a future update of TS 38.561 [4].

Proposal 1: Endorse the text proposal below.

# Conclusion

The following proposal is made in this contribution:

Proposal 1: Endorse the text proposal below.

# References

- RP-222669, “New WID: Enhancement of UE TRP (Total Radiated Power) and TRS (Total Radiated Sensitivity) requirements and test methodologies”, vivo, 3GPP TSG RAN Meeting #97-e, September 2022

- RP-223112, “Revised WID: Enhancement of UE TRP (Total Radiated Power) and TRS (Total Radiated Sensitivity) requirements and test methodologies for FR1 (NR SA and EN-DC)”, vivo, 3GPP TSG RAN Meeting #98-e, December 2022

- 3GPP TR 38.870, v0.6.0 (2023-10)

- 3GPP TS 38.561, v1.0.0 (2023-09)

# Appendix – Text Proposal to TR 38.870 v0.3.0

< Unchanged Text Deleted >

< Beginning of Changes >

## B.2.11	Uncertainty related to the use of phantoms

The following elements of uncertainty contributions (when applicable) are combined by the root-sum-squares (RSS) method to derive the total combined uncertainty related to the use of phantoms.

### B.2.11.1	Uncertainty from using different types of SAM phantom

This uncertainty contribution originates from the fact that different laboratories may use the two different versions of SAM head: the SAM head phantom or the SAM phantom including the head and the shoulders. The standard SAM head is the specified phantom. However, the use of the other type of SAM is also allowed with the requirement that the resulting uncertainty contribution is considered in the uncertainty budget.

### B.2.11.2	Simulated tissue liquid uncertainty

This uncertainty will occur, if the laboratory uses a liquid which has dielectric parameters deviating more than ±15% of the target parameters.

### B.2.11.3	Uncertainty of dielectric properties and shape of the hand phantom

The hand phantom contributes to OTA measurement uncertainty due to the manufacturing tolerances of its dielectric properties and shape. The dielectric properties on the surface of the hand may differ from those of its interior, so both are included in the evaluation. The moulded exterior surface of the hand shall be measured directly with an open-ended coaxial probe. The interior hand material is evaluated indirectly, by substituting a cube-shaped sample moulded from the same material and having some exterior surfaces removed. Following procedure will be used to evaluate the dielectric properties of the hand phantom;

1.	Each hand shall be manufactured together with a reference cube of the same material. The sides of the reference cube shall be not less than 40 mm in length.

2.	The moulded surface on three orthogonal sides of the cube shall be sliced away to a depth of at least 3 mm, in order to expose interior material for evaluation. The remaining three sides of the cube shall be left untreated.

3.	Relative permittivity and conductivity shall be measured at ten different points on each of the three cut, exposed surfaces of the reference cube, and the combined interior averages (εintavg, σintavg, 30 points) and standard deviations (εintstd,σintstd,30 points) shall be calculated. Individual interior averages for each of these three sides (εinti,σinti,10 points) shall also be calculated.

4.	Relative permittivity and conductivity shall be measured at ten points on the hand phantom exterior. A measurement point shall be located to each fingertip or as close to the tip as applicable. One measurement point shall be located to the back of the hand and one to the inner surface of wrist area. The exterior averages (εext_avg, σext_avg, 10 points) and standard deviations (εext_std, σext_std, 10 points) calculated accordingly.

5.	The total averages (εavg , σavg) shall be calculated as the average of exterior and interior values by either evaluating all data points or using equations :

εavg=εext_avg+3·εint_avg4

σavg=σext_avg+3·σint_avg4

6.	The total standard deviations (εstd , σstd) shall be calculated as the statistical combination of exterior and interior values by either evaluating all data points or using equations:

εstd=14εext_std2+εext_avg2+3*εint_std2+εint_avg2-εavg2

σstd=14σext_std2+σext_avg2+3*σint_std2+σint_avg2-σavg2

7.	The hands are acceptable for radiated performance testing, i.e., meet the minimal requirements, if

a.	εavg deviates by less than 15% from the target values

b.	σavg deviates by less than 25% from the target values

c.	the difference between the averaged permittivity of each 10-point interior surface εint_avg deviates by less than 10% and εext_avg by less than 20% from the total average εavg

d.	the difference between the averaged conductivity of each 10-point interior surfaceσint_avg deviates by less than 20% and σext_avg by less than 30% from the total average σavg

e.	the standard deviation of the combined measurements (30 interior points and 10 exterior points) is less than 20% for permittivity εstd and less than 40% for conductivity σstd

8.	For the hands meeting the minimal requirements of step 7, the following approximations shall be used to determine the hand uncertainty due to dielectric properties.

UεdB=c1·10log101+∆εavg2+εunc2+a1εstd2ε

UσdB=c2·10log101+∆σavg2+σunc2+a1σstd2ε

∆εavg, ∆σavg, εstd, σstd are the values determined as defined above and εunc and σunc are expanded measurement uncertainties (k = 2) of the dielectric parameter measurement method. The cube will be provided together with the hand such that the user can evaluate if the interior (cube) properties of the hand has degenerated over time by performing the test above. Coefficient c1=0.78, c2=0.39 and a1=0.50 were determined by numeric simulations.

In case the hand phantoms are manufactured within CAD models, the tolerance is 2% and therefore the effects shape errors are negligible. If the tolerance is larger, a numerical study must be conducted.

The assessment of uncertainties for the phantoms defined in CTIA Certification OTA Test Plan 01.72 Section 2 [25], is further detailed in CTIA Certification OTA Test Plan 01.70 Section 2 [26]. Values are from the CTIA Certification Test Plan for Wireless Device Over-the-Air Performance © CTIA Certification. Reproduced with permission.

### B.2.11.4	Uncertainty from using different types of Laptop Ground Plane phantom

This uncertainty contribution originates from the fact that different laboratories may use different variations of Laptop Ground Plane phantom. The standard Laptop Ground Plane is the specified phantom.

### B.2.11.5	Positioning Uncertainty from using Phantoms

Some uncertainty also occurs from the positioning of the DUT against the SAM phantom, as the DUT cannot be attached exactly in the same way every time. This uncertainty depends on how much the DUT’s positioning against the SAM phantom and hand phantoms varies from the specified testing positions. It is noted that the uncertainty of the phone positioning depends on the phone holder and the measurement operator and is in fact difficult to distinguish from random uncertainty. Some uncertainty also occurs from the positioning of the DUT plugged into the Laptop Ground Plane phantom, as the DUT may not be plugged into the USB connector and positioned exactly in the same way every time. This uncertainty depends on how much the DUT’s position plugged into the Laptop Ground Plane phantom varies from the specified plug-in position. Therefore, the positioning uncertainty is included in random uncertainty.

To estimate this uncertainty for the SAM phantom, it is suggested to perform at least five evaluations of TRP/TRS whereby the device shall be dismounted and newly positioned with a fully charged battery before each test. This measurement set has to be carried out in mid channel of lowest and highest frequency bands utilized by the testing lab, for at least three phones with different type of mechanical design. The values have to be normalized by the mean for each measurement set. As a result, the uncertainty contribution entered to uncertainty budget is the difference between the maximum and minimum normalized values.

With head and hand phantoms, random uncertainty evaluation may be done separately for each measurement configuration i.e. head only, browsing mode or speech mode. A speech mode random uncertainty evaluation, were both head and hand phantoms are used, can reasonably be considered to be the worst-case scenario and thus random uncertainties in other configurations to be less.

To estimate this uncertainty for the Laptop Ground Plane phantom, it is suggested to perform at least five evaluations of TRP/TRS for the plug-in position whereby the device shall be dismounted and newly positioned before each test. This measurement set has to be carried out in mid channel of lowest and highest frequency bands utilized by the testing lab, for at least three USBs with different type of mechanical design. The values have to be normalized by the mean for each measurement set. As a result, the uncertainty contribution entered to uncertainty budget is the difference between the maximum and minimum normalized value.

For the forearm phantom, the assessment of uncertainties for the phantoms defined in CTIA Certification OTA Test Plan 01.72 Section 2 [25], is further detailed in CTIA Certification OTA Test Plan 01.70 Section 2 [26]. Values are from the CTIA Certification Test Plan for Wireless Device Over-the-Air Performance © CTIA Certification. Reproduced with permission.

### B.2.11.5	Uncertainty of dielectric properties and shape of the forearm phantom

The assessment of uncertainties for the phantoms defined in CTIA Certification OTA Test Plan 01.72 Section 2 [25], is further detailed in CTIA Certification OTA Test Plan 01.70 Section 2 [26]. Values are from the CTIA Certification Test Plan for Wireless Device Over-the-Air Performance © CTIA Certification. Reproduced with permission.

< Unchanged Text Deleted >

# B.4	MU Assessment for TRP

## B.4.1	MU Assessment for TRP in Anechoic Chamber

The uncertainty contributions related to TRP are listed in Table B.4.1-1. A preliminary example uncertainty budgets are presented in Table B.4.1-2 and Table B.4.1-3.

Table B.4.1-1 Uncertainty contributions in TRP measurement for anechoic chamber method

| UID | Description of uncertainty contribution | Details in clause |
| --- | --- | --- |
| Stage 2: DUT measurement (Figure 7.2-1, Figure 7.2-2) |  |  |
| 1 | Mismatch of receiver chain | B.2.1 |
| 2 | Insertion loss of receiver chain | B.2.2 |
| 3 | Influence of the measurement antenna cable | B.2.3 |
| 4 | Measurement Receiver: uncertainty of the absolute level | B.2.4 |
| 5 | Measurement distance | B.2.7 |
| 6 | Quality of quiet zone | B.2.8 |
| 7 | DUT Tx-power drift | B.2.9 |
| 8 | Uncertainty related to the use of phantoms | B.2.11 |
| 9 | Coarse sampling grid | B.2.12 |
| 10 | Random uncertainty | B.2.13 |
| 11 | Frequency Response | B.2.14 |
| Stage 1: Calibration measurement, network analyzer method (Figure 7.3-1) |  |  |
| 12 | Uncertainty of network analyzer | B.2.15 |
| 13 | Mismatch of receiver chain | B.2.1 |
| 14 | Insertion loss of receiver chain | B.2.2 |
| 15 | Mismatch in the connection of calibration antenna | B.2.1 |
| 16 | Influence of the calibration antenna feed cable | B.2.3 |
| 17 | Influence of the measurement antenna cable | B.2.3 |
| 18 | Uncertainty of the absolute gain/ radiation efficiency of the calibration antenna | B.2.16 |
| 19 | Measurement distance | B.2.7 |
| 20 | Quality of the Quiet Zone | B.2.8 |
| Systematic Errors |  |  |
| 21 | Systematic Error related to TRP grids | B.2.12 |

Table B.4.1-2 Preliminary example of uncertainty budget for TRP hand only (browsing mode) measurement for anechoic chamber method for NR FR1 bands

| UID | Uncertainty Source | Comment | Uncertainty Value [dB] |  | Prob Distr | Div | ci | Standard Uncertainty [dB] |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  | Below 3GHz | Above 3GHz |  |  |  | Below 3GHz | Above 3GHz |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |  |
| 1 | Mismatch of receiver chain | Гreceiver < 0.33 Гmeasurement antenna < 0.5 Cable attenuation > 3dB | 0.26 | 0.26 | U-shaped | 1.4 | 1 | 0.18 | 0.18 |
| 2 | Insertion loss of receiver chain | Systematic with Stage 1 (=> cancels) | 0 | 0 | Rectangular | 1.7 | 1 | 0.00 | 0.00 |
| 3 | Influence of the measurement antenna cable | Systematic with Stage 1 (=> cancels) | 0 | 0 | Rectangular | 1.7 | 1 | 0.00 | 0.00 |
| 4 | Measurement Receiver: uncertainty of the absolute level | Spectrum Analyzer | 0.42 | 0.54 | Normal | 2 | 1 | 0.21 | 0.27 |
| 5 | Measurement distance | DUT is not offset for hand-only phantom testing | 0 | 0 | Rectangular | 1.7 | 1 | 0.00 | 0.00 |
| 6 | Quality of quiet zone | Surface standard deviation of power measurements in ripple test | 0.5 | 0.5 | Actual | 1 | 1 | 0.50 | 0.50 |
| 7 | DUT Tx-power drift | Drift | 0.2 | 0.2 | Rectangular | 1.7 | 1 | 0.12 | 0.12 |
| 8 | Uncertainty related to the use of phantoms | Material Dielectric Constant, Material Conductivity, Geometry/Shape (incl. spacer), Data Mode Fixture | 0.64 | 0.64 | Rectangular | 1.7 | 1 | 0.37 | 0.37 |
| 9 | Coarse sampling grid | Sampling grids per Table Table B.2.12-1 | 0 | 0.11 | Actual | 1 | 1 | 0.00 | 0.11 |
| 10 | Random Uncertainty | Fixed MU to account for all the unknown, unquantifiable, etc. uncertainties | 0.25 | 0.25 | Normal | 2 | 1 | 0.13 | 0.13 |
| 11 | Frequency Response | Average path loss corrected | 0 | 0 | Rectangular | 1.7 | 1 | 0.00 | 0.00 |
| Stage 1: Calibration measurement, network analyzer method |  |  |  |  |  |  |  |  |  |
| 12 | Uncertainty of network analyzer | From datasheet of VNA with assessed transmission coefficients | 0.2 | 0.5 | Normal | 2 | 1 | 0.10 | 0.25 |
| 13 | Mismatch of receiver chain | Taken into account in VNA uncertainty term | 0 | 0 | U-shaped | 1.4 | 1 | 0.00 | 0.00 |
| 14 | Insertion loss of receiver chain | Systematic with Stage 2 (=> cancels) | 0 | 0 | Rectangular | 1.7 | 1 | 0.00 | 0.00 |
| 15 | Mismatch in the connection of calibration antenna | Taken in to account in VNA setup uncertainty | 0 | 0 | U-shaped | 1.4 | 1 | 0.00 | 0.00 |
| 16 | Influence of the calibration antenna feed cable | Gain calibration with a dipole | 0.3 | 0.3 | Rectangular | 1.7 | 1 | 0.17 | 0.17 |
| 17 | Influence of the measurement antenna cable | Systematic with Stage 2 (=> cancels) | 0 | 0 | Rectangular | 1.7 | 1 | 0.00 | 0.00 |
| 18 | Uncertainty of the absolute gain/ radiation efficiency of the calibration antenna | Calibration report with traceability to a National Metrology Institute | 0.58 | 0.58 | Normal | 2 | 1 | 0.29 | 0.29 |
| 19 | Measurement distance | Dipole: aligned with phase center | 0 | 0 | Rectangular | 1.7 | 1 | 0.00 | 0.00 |
| 20 | Quality of the Quiet Zone | Peak-to-null ripple | 0.5 | 0.5 | Rectangular | 1.7 | 1 | 0.29 | 0.29 |
| Combined standard uncertainty [dB] |  |  |  |  |  |  |  | [0.84] | [0.89] |
| Expanded uncertainty [dB] (Confidence interval of 95 %) |  |  |  |  |  |  |  | [1.64] | [1.75] |
| 21 | Systematic Error related to TRP grids | mean error | 0 | 0 | Actual | 1 | 1 | 0.00 | 0.00 |
| Total Expanded uncertainty [dB] (Confidence interval of 95 %) |  |  |  |  |  |  |  | [1.64] | [1.75] |

Table B.4.1-3 Preliminary example of uncertainty budget for TRP Beside Head and Hand (Talk mode) measurement for anechoic chamber method for NR FR1 bands

| UID | Uncertainty Source | Comment | Uncertainty Value [dB] |  | Prob Distr | Div | ci | Standard Uncertainty [dB] |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  | Below 3GHz | Above 3GHz |  |  |  | Below 3GHz | Above 3GHz |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |  |
| 1 | Mismatch of receiver chain | Гreceiver < 0.33 Гmeasurement antenna < 0.5 Cable attenuation > 3dB | 0.26 | 0.26 | U-shaped | 1.4 | 1 | 0.18 | 0.18 |
| 2 | Insertion loss of receiver chain | Systematic with Stage 1 (=> cancels) | 0 | 0 | Rectangular | 1.7 | 1 | 0.00 | 0.00 |
| 3 | Influence of the measurement antenna cable | Systematic with Stage 1 (=> cancels) | 0 | 0 | Rectangular | 1.7 | 1 | 0.00 | 0.00 |
| 4 | Measurement Receiver: uncertainty of the absolute level | Spectrum Analyzer | 0.42 | 0.54 | Normal | 2 | 1 | 0.21 | 0.27 |
| 5 | Measurement distance | d=1.6m, Δd=0.05m | 0.27 | 0.27 | Rectangular | 1.7 | 1 | 0.16 | 0.16 |
| 6 | Quality of quiet zone | Surface standard deviation of power measurements in ripple test | 0.5 | 0.5 | Actual | 1 | 1 | 0.50 | 0.50 |
| 7 | DUT Tx-power drift | Drift | 0.2 | 0.2 | Rectangular | 1.7 | 1 | 0.12 | 0.12 |
| 8 | Uncertainty related to the use of phantoms | Material Dielectric Constant, Material Conductivity, Geometry/Shape (incl. spacer), Beside Head and Hand | 0.99 | 0.99 | Rectangular | 1.7 | 1 | 0.57 | 0.57 |
| 9 | Coarse sampling grid | Sampling grids per Table Table B.2.12-1 | 0 | 0.11 | Actual | 1 | 1 | 0.00 | 0.11 |
| 10 | Random Uncertainty | Fixed MU to account for all the unknown, unquantifiable, etc. uncertainties | 0.25 | 0.25 | Normal | 2 | 1 | 0.13 | 0.13 |
| 11 | Frequency Response | Average path loss corrected | 0 | 0 | Rectangular | 1.7 | 1 | 0.00 | 0.00 |
| Stage 1: Calibration measurement, network analyzer method |  |  |  |  |  |  |  |  |  |
| 12 | Uncertainty of network analyzer | From datasheet of VNA with assessed transmission coefficients | 0.2 | 0.5 | Normal | 2 | 1 | 0.10 | 0.25 |
| 13 | Mismatch of receiver chain | Taken into account in VNA uncertainty term | 0 | 0 | U-shaped | 1.4 | 1 | 0.00 | 0.00 |
| 14 | Insertion loss of receiver chain | Systematic with Stage 2 (=> cancels) | 0 | 0 | Rectangular | 1.7 | 1 | 0.00 | 0.00 |
| 15 | Mismatch in the connection of calibration antenna | Taken in to account in VNA setup uncertainty | 0 | 0 | U-shaped | 1.4 | 1 | 0.00 | 0.00 |
| 16 | Influence of the calibration antenna feed cable | Gain calibration with a dipole | 0.3 | 0.3 | Rectangular | 1.7 | 1 | 0.17 | 0.17 |
| 17 | Influence of the measurement antenna cable | Systematic with Stage 2 (=> cancels) | 0 | 0 | Rectangular | 1.7 | 1 | 0.00 | 0.00 |
| 18 | Uncertainty of the absolute gain/ radiation efficiency of the calibration antenna | Calibration report with traceability to a National Metrology Institute | 0.58 | 0.58 | Normal | 2 | 1 | 0.29 | 0.29 |
| 19 | Measurement distance | Dipole: aligned with phase center | 0 | 0 | Rectangular | 1.7 | 1 | 0.00 | 0.00 |
| 20 | Quality of the Quiet Zone | Peak-to-null ripple | 0.5 | 0.5 | Rectangular | 1.7 | 1 | 0.29 | 0.29 |
| Combined standard uncertainty [dB] |  |  |  |  |  |  |  | [0.96] | [1.00] |
| Expanded uncertainty [dB] (Confidence interval of 95 %) |  |  |  |  |  |  |  | [1.88] | [1.97] |
| 21 | Systematic Error related to TRP grids | mean error | 0 | 0 | Actual | 1 | 1 | 0.00 | 0.00 |
| Total Expanded uncertainty [dB] (Confidence interval of 95 %) |  |  |  |  |  |  |  | [1.88] | [1.97] |

Table B.4.1-4 Preliminary example of uncertainty budget for TRP Wrist-Worn device measurement for anechoic chamber method for NR FR1 bands

| UID | Uncertainty Source | Comment | Uncertainty Value [dB] |  | Prob Distr | Div | ci | Standard Uncertainty [dB] |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  | Below 3GHz | Above 3GHz |  |  |  | Below 3GHz | Above 3GHz |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |  |
| 1 | Mismatch of receiver chain | Гreceiver < 0.33 Гmeasurement antenna < 0.5 Cable attenuation > 3dB | 0.26 | 0.26 | U-shaped | 1.4 | 1 | 0.18 | 0.18 |
| 2 | Insertion loss of receiver chain | Systematic with Stage 1 (=> cancels) | 0 | 0 | Rectangular | 1.7 | 1 | 0.00 | 0.00 |
| 3 | Influence of the measurement antenna cable | Systematic with Stage 1 (=> cancels) | 0 | 0 | Rectangular | 1.7 | 1 | 0.00 | 0.00 |
| 4 | Measurement Receiver: uncertainty of the absolute level | Spectrum Analyzer | 0.42 | 0.54 | Normal | 2 | 1 | 0.21 | 0.27 |
| 5 | Measurement distance | d=1.6m, Δd=0.05m | 0.27 | 0.27 | Rectangular | 1.7 | 1 | 0.16 | 0.16 |
| 6 | Quality of quiet zone | Surface standard deviation of power measurements in ripple test | 0.5 | 0.5 | Actual | 1 | 1 | 0.50 | 0.50 |
| 7 | DUT Tx-power drift | Drift | 0.2 | 0.2 | Rectangular | 1.7 | 1 | 0.12 | 0.12 |
| 8 | Uncertainty related to the use of phantoms | Material Dielectric Constant, Material Conductivity, Geometry/Shape (incl. spacer), Forearm | 0.74 | 0.74 | Rectangular | 1.7 | 1 | 0.43 | 0.43 |
| 9 | Coarse sampling grid | Sampling grids per Table Table B.2.12-1 | 0 | 0.11 | Actual | 1 | 1 | 0.00 | 0.11 |
| 10 | Random Uncertainty | Fixed MU to account for all the unknown, unquantifiable, etc. uncertainties | 0.25 | 0.25 | Normal | 2 | 1 | 0.13 | 0.13 |
| 11 | Frequency Response | Average path loss corrected | 0 | 0 | Rectangular | 1.7 | 1 | 0.00 | 0.00 |
| Stage 1: Calibration measurement, network analyzer method |  |  |  |  |  |  |  |  |  |
| 12 | Uncertainty of network analyzer | From datasheet of VNA with assessed transmission coefficients | 0.2 | 0.5 | Normal | 2 | 1 | 0.10 | 0.25 |
| 13 | Mismatch of receiver chain | Taken into account in VNA uncertainty term | 0 | 0 | U-shaped | 1.4 | 1 | 0.00 | 0.00 |
| 14 | Insertion loss of receiver chain | Systematic with Stage 2 (=> cancels) | 0 | 0 | Rectangular | 1.7 | 1 | 0.00 | 0.00 |
| 15 | Mismatch in the connection of calibration antenna | Taken in to account in VNA setup uncertainty | 0 | 0 | U-shaped | 1.4 | 1 | 0.00 | 0.00 |
| 16 | Influence of the calibration antenna feed cable | Gain calibration with a dipole | 0.3 | 0.3 | Rectangular | 1.7 | 1 | 0.17 | 0.17 |
| 17 | Influence of the measurement antenna cable | Systematic with Stage 2 (=> cancels) | 0 | 0 | Rectangular | 1.7 | 1 | 0.00 | 0.00 |
| 18 | Uncertainty of the absolute gain/ radiation efficiency of the calibration antenna | Calibration report with traceability to a National Metrology Institute | 0.58 | 0.58 | Normal | 2 | 1 | 0.29 | 0.29 |
| 19 | Measurement distance | Dipole: aligned with phase center | 0 | 0 | Rectangular | 1.7 | 1 | 0.00 | 0.00 |
| 20 | Quality of the Quiet Zone | Peak-to-null ripple | 0.5 | 0.5 | Rectangular | 1.7 | 1 | 0.29 | 0.29 |
| Combined standard uncertainty [dB] |  |  |  |  |  |  |  | [0.88] | [0.93] |
| Expanded uncertainty [dB] (Confidence interval of 95 %) |  |  |  |  |  |  |  | [1.72] | [1.82] |
| 21 | Systematic Error related to TRP grids | mean error | 0 | 0 | Actual | 1 | 1 | 0.00 | 0.00 |
| Total Expanded uncertainty [dB] (Confidence interval of 95 %) |  |  |  |  |  |  |  | [1.72] | [1.82] |

< Unchanged Text Deleted >

# B.5	MU Assessment for TRS

## B.5.1	MU Assessment for TRS in Anechoic Chamber

The uncertainty contributions related to TRS are listed in Table B.5.1-1. A preliminary example uncertainty budgets are presented in Table B.5.1-2 and Table B.5.1-3.

Table B.5.1-1 Uncertainty contributions in TRS measurement for anechoic chamber method

| UID | Description of uncertainty contribution | Details in clause |
| --- | --- | --- |
| Stage 2: DUT measurement (Figure 7.2-1, Figure 7.2-2) |  |  |
| 1 | Mismatch of transmitter chain | B.2.1 |
| 2 | Insertion loss of transmitter chain | B.2.2 |
| 3 | Influence of the measurement antenna cable | B.2.3 |
| 4 | Communication Tester: uncertainty of the absolute output level | B.2.5 |
| 5 | Sensitivity measurement: output level step resolution | B.2.6 |
| 6 | Measurement distance | B.2.7 |
| 7 | Quality of quiet zone | B.2.8 |
| 8 | DUT sensitivity drift | B.2.10 |
| 9 | Uncertainty related to the use of phantoms | B.2.11 |
| 10 | Coarse sampling grid | B.2.12 |
| 11 | Random uncertainty | B.2.13 |
| 12 | Frequency Response | B.2.14 |
| Stage 1: Calibration measurement, network analyzer method (Figure 7.3-1) |  |  |
| 13 | Uncertainty of network analyzer | B.2.15 |
| 14 | Mismatch of transmitter chain | B.2.1 |
| 15 | Insertion loss of transmitter chain | B.2.2 |
| 16 | Mismatch in the connection of calibration antenna | B.2.1 |
| 17 | Influence of the calibration antenna feed cable | B.2.3 |
| 18 | Influence of the measurement antenna cable | B.2.3 |
| 19 | Uncertainty of the absolute gain/radiation efficiency of the calibration antenna | B.2.16 |
| 20 | Measurement distance | B.2.7 |
| 21 | Quality of quiet zone | B.2.8 |
| Systematic Errors |  |  |
| 22 | Systematic Error related to TRS grids | B.2.12 |

Table B.5.1-2: Preliminary example of uncertainty budget for TRS hand only (browsing mode) measurement for anechoic chamber method for NR FR1 bands

| UID | Uncertainty Source | Comment | Uncertainty Value [dB] |  | Prob Distr |  | Div |  | ci |  | Standard Uncertainty [dB] |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  | Below 3GHz | Above 3GHz |  |  |  |  |  |  | Below 3GHz | Above 3GHz |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |  |  |  |  |
| 1 | Mismatch of receiver chain | Гcomm tester < 0.29 Гmeasurement antenna < 0.5 Cable attenuation > 3dB | 0.22 | 0.22 | U-shaped |  | 1.4 |  | 1 |  | 0.16 | 0.16 |
| 2 | Insertion loss of receiver chain | Systematic with Stage 1 (=> cancels) | 0 | 0 | Rectangular |  | 1.7 |  | 1 |  | 0.00 | 0.00 |
| 3 | Influence of the measurement antenna cable | Systematic with Stage 1 (=> cancels) | 0 | 0 | Rectangular |  | 1.7 |  | 1 |  | 0.00 | 0.00 |
| 4 | Communication Tester: uncertainty of the absolute output level | Manufacturer’s data sheet | [1.16] | [1.16] | Normal |  | 2 |  | 1 |  | [0.58] | [0.58] |
| 5 | Sensitivity measurement: output level step resolution | Systematic error that can be corrected | 0 | 0 | Rectangular |  | 1.73 |  | 1 |  | 0.00 | 0.00 |
| 6 | Measurement distance | DUT is not offset for hand-only phantom testing | 0 | 0 | Rectangular |  | 1.7 |  | 1 |  | 0.00 | 0.00 |
| 7 | Quality of quiet zone | Surface standard deviation of power measurements in ripple test | 0.5 | 0.5 | Actual |  | 1 |  | 1 |  | 0.50 | 0.50 |
| 8 | DUT sensitivity drift | Drift | 0.2 | 0.2 | Rectangular |  | 1.7 |  | 1 |  | 0.12 | 0.12 |
| 9 | Uncertainty related to the use of phantoms | Material Dielectric Constant, Material Conductivity, Geometry/Shape (incl. spacer), Data Mode Fixture | 0.64 | 0.64 | Rectangular |  | 1.7 |  | 1 |  | 0.37 | 0.37 |
| 10 | Coarse sampling grid | Sampling grids per Table Table B.2.12-1 | 0.04 | 0.25 | Actual |  | 1 |  | 1 |  | 0.04 | 0.25 |
| 11 | Random Uncertainty | Fixed MU to account for all the unknown, unquantifiable, etc. uncertainties including digital error rate | 0.4 | 0.4 | Normal |  | 2 |  | 1 |  | 0.20 | 0.20 |
| 12 | Frequency Response | Included in the output level step resolution | 0 | 0 | Rectangular |  | 1.7 |  | 1 |  | 0.00 | 0.00 |
| Stage 1: Calibration measurement, network analyzer method |  |  |  |  |  |  |  |  |  |  |  |  |
| 13 | Uncertainty of network analyzer | From datasheet of VNA with assessed transmission coefficients | 0.2 | 0.5 | Normal |  | 2 |  | 1 |  | 0.10 | 0.25 |
| 14 | Mismatch of receiver chain | Taken into account in VNA uncertainty term | 0 | 0 | U-shaped |  | 1.4 |  | 1 |  | 0.00 | 0.00 |
| 15 | Insertion loss of receiver chain | Systematic with Stage 2 (=> cancels) | 0 | 0 | Rectangular |  | 1.7 |  | 1 |  | 0.00 | 0.00 |
| 16 | Mismatch in the connection of calibration antenna | Taken in to account in VNA setup uncertainty | 0 | 0 | U-shaped |  | 1.4 |  | 1 |  | 0.00 | 0.00 |
| 17 | Influence of the calibration antenna feed cable | Gain calibration with a dipole | 0.3 | 0.3 | Rectangular |  | 1.7 |  | 1 |  | 0.17 | 0.17 |
| 18 | Influence of the measurement antenna cable | Systematic with Stage 2 (=> cancels) | 0 | 0 | Rectangular |  | 1.7 |  | 1 |  | 0.00 | 0.00 |
| 19 | Uncertainty of the absolute gain/ radiation efficiency of the calibration antenna | Calibration report with traceability to a National Metrology Institute | 0.58 | 0.58 | Normal |  | 2 |  | 1 |  | 0.29 | 0.29 |
| 20 | Measurement distance | Dipole: aligned with phase center | 0 | 0 | Rectangular |  | 1.7 |  | 1 |  | 0.00 | 0.00 |
| 21 | Quality of the Quiet Zone | Peak-to-null ripple | 0.5 | 0.5 | Rectangular |  | 1.7 |  | 1 |  | 0.29 | 0.29 |
| Combined standard uncertainty [dB] |  |  |  |  |  |  |  |  |  |  | [1.00] | [1.06] |
| Expanded uncertainty [dB] (Confidence interval of 95 %) |  |  |  |  |  |  |  |  |  |  | [1.97] | [2.08] |
| 22 | Systematic Error related to TRS grids | mean error | 0 | 0.08 |  | Actual |  | 1 |  | 1 | 0.00 | 0.08 |
| Total Expanded uncertainty [dB] (Confidence interval of 95 %) |  |  |  |  |  |  |  |  |  |  | [1.97] | [2.16] |

Table B.5.1-3: Preliminary example of uncertainty budget for TRS Beside Head and Hand (Talk mode) measurement for anechoic chamber method for NR FR1 bands

| UID | Uncertainty Source | Comment | Uncertainty Value [dB] |  | Prob Distr |  | Div |  | ci |  | Standard Uncertainty [dB] |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  | Below 3GHz | Above 3GHz |  |  |  |  |  |  | Below 3GHz | Above 3GHz |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |  |  |  |  |
| 1 | Mismatch of receiver chain | Гcomm tester < 0.29 Гmeasurement antenna < 0.5 Cable attenuation > 3dB | 0.22 | 0.22 | U-shaped |  | 1.4 |  | 1 |  | 0.16 | 0.16 |
| 2 | Insertion loss of receiver chain | Systematic with Stage 1 (=> cancels) | 0 | 0 | Rectangular |  | 1.7 |  | 1 |  | 0.00 | 0.00 |
| 3 | Influence of the measurement antenna cable | Systematic with Stage 1 (=> cancels) | 0 | 0 | Rectangular |  | 1.7 |  | 1 |  | 0.00 | 0.00 |
| 4 | Communication Tester: uncertainty of the absolute output level | Manufacturer’s data sheet | [1.16] | [1.16] | Normal |  | 2 |  | 1 |  | [0.58] | [0.58] |
| 5 | Sensitivity measurement: output level step resolution | Systematic error that can be corrected | 0 | 0 | Rectangular |  | 1.73 |  | 1 |  | 0.00 | 0.00 |
| 6 | Measurement distance | d=1.6m, Δd=0.05m | 0.27 | 0.27 | Rectangular |  | 1.7 |  | 1 |  | 0.16 | 0.16 |
| 7 | Quality of quiet zone | Surface standard deviation of power measurements in ripple test | 0.5 | 0.5 | Actual |  | 1 |  | 1 |  | 0.50 | 0.50 |
| 8 | DUT sensitivity drift | Drift | 0.2 | 0.2 | Rectangular |  | 1.7 |  | 1 |  | 0.12 | 0.12 |
| 9 | Uncertainty related to the use of phantoms | Material Dielectric Constant, Material Conductivity, Geometry/Shape (incl. spacer), Beside Head and Hand | 0.99 | 0.99 | Rectangular |  | 1.7 |  | 1 |  | 0.57 | 0.57 |
| 10 | Coarse sampling grid | Sampling grids per Table Table B.2.12-1 | 0.04 | 0.25 | Actual |  | 1 |  | 1 |  | 0.04 | 0.25 |
| 11 | Random Uncertainty | Fixed MU to account for all the unknown, unquantifiable, etc. uncertainties including digital error rate | 0.4 | 0.4 | Normal |  | 2 |  | 1 |  | 0.20 | 0.20 |
| 12 | Frequency Response | Included in the output level step resolution | 0 | 0 | Rectangular |  | 1.7 |  | 1 |  | 0.00 | 0.00 |
| Stage 1: Calibration measurement, network analyzer method |  |  |  |  |  |  |  |  |  |  |  |  |
| 13 | Uncertainty of network analyzer | From datasheet of VNA with assessed transmission coefficients | 0.2 | 0.5 | Normal |  | 2 |  | 1 |  | 0.10 | 0.25 |
| 14 | Mismatch of receiver chain | Taken into account in VNA uncertainty term | 0 | 0 | U-shaped |  | 1.4 |  | 1 |  | 0.00 | 0.00 |
| 15 | Insertion loss of receiver chain | Systematic with Stage 2 (=> cancels) | 0 | 0 | Rectangular |  | 1.7 |  | 1 |  | 0.00 | 0.00 |
| 16 | Mismatch in the connection of calibration antenna | Taken in to account in VNA setup uncertainty | 0 | 0 | U-shaped |  | 1.4 |  | 1 |  | 0.00 | 0.00 |
| 17 | Influence of the calibration antenna feed cable | Gain calibration with a dipole | 0.3 | 0.3 | Rectangular |  | 1.7 |  | 1 |  | 0.17 | 0.17 |
| 18 | Influence of the measurement antenna cable | Systematic with Stage 2 (=> cancels) | 0 | 0 | Rectangular |  | 1.7 |  | 1 |  | 0.00 | 0.00 |
| 19 | Uncertainty of the absolute gain/ radiation efficiency of the calibration antenna | Calibration report with traceability to a National Metrology Institute | 0.58 | 0.58 | Normal |  | 2 |  | 1 |  | 0.29 | 0.29 |
| 20 | Measurement distance | Dipole: aligned with phase center | 0 | 0 | Rectangular |  | 1.7 |  | 1 |  | 0.00 | 0.00 |
| 21 | Quality of the Quiet Zone | Peak-to-null ripple | 0.5 | 0.5 | Rectangular |  | 1.7 |  | 1 |  | 0.29 | 0.29 |
| Combined standard uncertainty [dB] |  |  |  |  |  |  |  |  |  |  | [1.11] | [1.16] |
| Expanded uncertainty [dB] (Confidence interval of 95 %) |  |  |  |  |  |  |  |  |  |  | [2.17] | [2.27] |
| 22 | Systematic Error related to TRS grids | mean error | 0 | 0.08 |  | Actual |  | 1 |  | 1 | 0.00 | 0.08 |
| Total Expanded uncertainty [dB] (Confidence interval of 95 %) |  |  |  |  |  |  |  |  |  |  | [2.17] | [2.35] |

Table B.5.1-3: Preliminary example of uncertainty budget for TRS Wrist-Worn device measurement for anechoic chamber method for NR FR1 bands

| UID | Uncertainty Source | Comment | Uncertainty Value [dB] |  | Prob Distr |  | Div |  | ci |  | Standard Uncertainty [dB] |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  | Below 3GHz | Above 3GHz |  |  |  |  |  |  | Below 3GHz | Above 3GHz |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |  |  |  |  |
| 1 | Mismatch of receiver chain | Гcomm tester < 0.29 Гmeasurement antenna < 0.5 Cable attenuation > 3dB | 0.22 | 0.22 | U-shaped |  | 1.4 |  | 1 |  | 0.16 | 0.16 |
| 2 | Insertion loss of receiver chain | Systematic with Stage 1 (=> cancels) | 0 | 0 | Rectangular |  | 1.7 |  | 1 |  | 0.00 | 0.00 |
| 3 | Influence of the measurement antenna cable | Systematic with Stage 1 (=> cancels) | 0 | 0 | Rectangular |  | 1.7 |  | 1 |  | 0.00 | 0.00 |
| 4 | Communication Tester: uncertainty of the absolute output level | Manufacturer’s data sheet | [1.16] | [1.16] | Normal |  | 2 |  | 1 |  | [0.58] | [0.58] |
| 5 | Sensitivity measurement: output level step resolution | Systematic error that can be corrected | 0 | 0 | Rectangular |  | 1.73 |  | 1 |  | 0.00 | 0.00 |
| 6 | Measurement distance | d=1.6m, Δd=0.05m | 0.27 | 0.27 | Rectangular |  | 1.7 |  | 1 |  | 0.16 | 0.16 |
| 7 | Quality of quiet zone | Surface standard deviation of power measurements in ripple test | 0.5 | 0.5 | Actual |  | 1 |  | 1 |  | 0.50 | 0.50 |
| 8 | DUT sensitivity drift | Drift | 0.2 | 0.2 | Rectangular |  | 1.7 |  | 1 |  | 0.12 | 0.12 |
| 9 | Uncertainty related to the use of phantoms | Material Dielectric Constant, Material Conductivity, Geometry/Shape (incl. spacer), Forearm | 0.74 | 0.74 | Rectangular |  | 1.7 |  | 1 |  | 0.43 | 0.43 |
| 10 | Coarse sampling grid | Sampling grids per Table Table B.2.12-1 | 0.04 | 0.25 | Actual |  | 1 |  | 1 |  | 0.04 | 0.25 |
| 11 | Random Uncertainty | Fixed MU to account for all the unknown, unquantifiable, etc. uncertainties including digital error rate | 0.4 | 0.4 | Normal |  | 2 |  | 1 |  | 0.20 | 0.20 |
| 12 | Frequency Response | Included in the output level step resolution | 0 | 0 | Rectangular |  | 1.7 |  | 1 |  | 0.00 | 0.00 |
| Stage 1: Calibration measurement, network analyzer method |  |  |  |  |  |  |  |  |  |  |  |  |
| 13 | Uncertainty of network analyzer | From datasheet of VNA with assessed transmission coefficients | 0.2 | 0.5 | Normal |  | 2 |  | 1 |  | 0.10 | 0.25 |
| 14 | Mismatch of receiver chain | Taken into account in VNA uncertainty term | 0 | 0 | U-shaped |  | 1.4 |  | 1 |  | 0.00 | 0.00 |
| 15 | Insertion loss of receiver chain | Systematic with Stage 2 (=> cancels) | 0 | 0 | Rectangular |  | 1.7 |  | 1 |  | 0.00 | 0.00 |
| 16 | Mismatch in the connection of calibration antenna | Taken in to account in VNA setup uncertainty | 0 | 0 | U-shaped |  | 1.4 |  | 1 |  | 0.00 | 0.00 |
| 17 | Influence of the calibration antenna feed cable | Gain calibration with a dipole | 0.3 | 0.3 | Rectangular |  | 1.7 |  | 1 |  | 0.17 | 0.17 |
| 18 | Influence of the measurement antenna cable | Systematic with Stage 2 (=> cancels) | 0 | 0 | Rectangular |  | 1.7 |  | 1 |  | 0.00 | 0.00 |
| 19 | Uncertainty of the absolute gain/ radiation efficiency of the calibration antenna | Calibration report with traceability to a National Metrology Institute | 0.58 | 0.58 | Normal |  | 2 |  | 1 |  | 0.29 | 0.29 |
| 20 | Measurement distance | Dipole: aligned with phase center | 0 | 0 | Rectangular |  | 1.7 |  | 1 |  | 0.00 | 0.00 |
| 21 | Quality of the Quiet Zone | Peak-to-null ripple | 0.5 | 0.5 | Rectangular |  | 1.7 |  | 1 |  | 0.29 | 0.29 |
| Combined standard uncertainty [dB] |  |  |  |  |  |  |  |  |  |  | [1.04] | [1.09] |
| Expanded uncertainty [dB] (Confidence interval of 95 %) |  |  |  |  |  |  |  |  |  |  | [2.04] | [2.14] |
| 22 | Systematic Error related to TRS grids | mean error | 0 | 0.08 |  | Actual |  | 1 |  | 1 | 0.00 | 0.08 |
| Total Expanded uncertainty [dB] (Confidence interval of 95 %) |  |  |  |  |  |  |  |  |  |  | [2.04] | [2.22] |

< End of Changes >