3GPP TSG-RAN5 Meeting #101 	R5-237735

# Chicago, USA, 13th Nov 2023 - 17th Nov 2023

| CR-Form-v12.2 |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CHANGE REQUEST |  |  |  |  |  |  |  |  |
|  | 38.903 | CR | 0623 | rev | 1 | Current version: | 18.0.0 |  |
| For HELP on using this form: comprehensive instructions can be found at http://www.3gpp.org/Change-Requests. |  |  |  |  |  |  |  |  |

| Proposed change affects: | UICC apps |  | ME |  | Radio Access Network |  | Core Network |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |

| Title: | Update for FR2c MU |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Source to WG: | Anritsu |  |  |  |  |  |  |  |  |  |
| Source to TSG: | R5 |  |  |  |  |  |  |  |  |  |
| Work item code: | TEI16_Test, NR_bands_BW_R16-UEConTest |  |  |  |  |  | Date: |  |  | 2023-11-03 |
| Category: | F |  |  |  |  |  | Release: |  |  | Rel-18 |
|  | Use one of the following categories: F (correction) A (mirror corresponding to a change in an earlier release) B (addition of feature), C (functional modification of feature) D (editorial modification) Detailed explanations of the above categories can be found in 3GPP TR 21.900. |  |  |  |  |  |  |  | Use one of the following releases: Rel-8 (Release 8) Rel-9 (Release 9) Rel-10 (Release 10) Rel-11 (Release 11) … Rel-16 (Release 16) Rel-17 (Release 17) Rel-18 (Release 18) Rel-19 (Release 19) |  |
| Reason for change: |  | FR2c MU is missing in many test cases. |  |  |  |  |  |  |  |  |
| Summary of change: |  | Uncertainty values were defined for FR2c (up to 87 GHz for spurious): - B.2.2.6 Uncertainty of the RF power measurement equipment - B.2.2.10 Influence of XPD - B.2.2.15 Uncertainty of the absolute gain of the calibration antenna - B.2.2.17 gNB emulator uncertainty - B.2.2.21 Influence of the calibration antenna feed cable MTSU was defined as TBD for FR2c: - B.3 UE maximum output power (Spherical coverage) - B.4 UE maximum output power for modulation / channel bandwidth - B.7 Minimum Output power - B.17 Adjacent Channel Leakage Ratio - B.21 Adjacent Channel Selectivity |  |  |  |  |  |  |  |  |
| Consequences if not approved: |  | FR2c test cases will not be completed. |  |  |  |  |  |  |  |  |
| Clauses affected: |  | B.2.2, B.3, B.4, B.7, B.17, B.21 |  |  |  |  |  |  |  |  |
|  |  | Y | N |  |  |  |  |  |  |  |
| Other specs |  |  | x | Other core specifications |  |  |  | TS/TR ... CR ... |  |  |
| affected: |  |  | x | Test specifications |  |  |  | TS/TR ... CR ... |  |  |
| (show related CRs) |  |  | x | O&M Specifications |  |  |  | TS/TR ... CR ... |  |  |
| Other comments: |  | Depending on the outcome of MU discussion R5-237724 |  |  |  |  |  |  |  |  |
| This CR's revision history: |  | This is an update of R5-236791 and the outcome of 2 revisions to this document. r1: Updates to follow the outcome of MU discussion and solve the overlap with R5-236070. r2: Following changes were removed to solve the overlap with R5-237256. - Table B.3-3 - Table B.7-1 - Table B.7.2-4 |  |  |  |  |  |  |  |  |

<<Unchanged sections skipped>>

B.2.2	Measurement error contribution descriptions for IFF

<<Unchanged sections skipped>>

B.2.2.6	Uncertainty of the RF power measurement equipment

See B.2.1.6.

The uncertainty value of RF power measurement equipment is estimated as below table and used across clause B.

Table B.2.2.6-1: Uncertainty value for RF power measurement equipment for IFF

| Power class | Test case | Frequency range | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- | --- |
| PC1, PC3, PC5, PC6 | MOP, MPR, Configured output power with power boost,SEM, ACLR | FR2a, FR2b | 2.16 | Normal | 2.00 | 1.08 |
|  | Minimum output power, OFF power |  | 2.50 | Normal | 2.00 | 1.25 |
|  | SE | 6GHz to 12.75GHz | 2.00 | Normal | 2.00 | 1.00 |
|  |  | 12.75GHz to 23.45GHz | 2.16 | Normal | 2.00 | 1.08 |
|  |  | 23.45GHz to 40.8GHz | 2.73 | Normal | 2.00 | 1.37 |
|  |  | 40.8GHz to 66GHz | 4.00 | Normal | 2.00 | 2.00 |
|  |  | 66GHz to 80GHz | 4.00 | Normal | 2.00 | 2.00 |
| PC3 | MOP | FR2c | 3.6 | Normal | 2.00 | 1.8 |

<<Unchanged sections skipped>>

B.2.2.10	Influence of XPD

See B.2.1.10.

The uncertainty value of influence of the XPD is estimated as below table and used across clause B.

Table B.2.2.10-2: Uncertainty value for influence of the XPD for IFF

| QZ size | Power class | Test case | Frequency range | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- | --- | --- |
| All | All | ACLR | FR2a, FR2b, FR2c | 0.00 | U-shaped | 1.41 | 0.00 |
| ≤30cm | PC1, PC3, PC5, PC6 | Default | FR2a, FR2b, FR2c | 0.01 | U-shaped | 1.41 | 0.00 |
|  |  | SE | 6GHz to 12.75GHz | 0.09 | U-shaped | 1.41 | 0.064 |
|  |  |  | 12.75GHz to 23.45GHz | 0.09 | U-shaped | 1.41 | 0.064 |
|  |  |  | 23.45GHz to 40.8GHz | 0.01 | U-shaped | 1.41 | 0.00 |
|  |  |  | 40.8GHz to 66GHz | 0.09 | U-shaped | 1.41 | 0.064 |
|  |  |  | 66GHz to 87GHz | 0.09 | U-shaped | 1.41 | 0.064 |
| 40cm | PC1, PC3, PC5, PC6 | Default | FR2a, FR2b, FR2c | 0.07 (NOTE 1) | U-shaped | 1.41 | 0.02 |
|  |  | SE | 6GHz to 12.75GHz | TBD | U-shaped | 1.41 | TBD |
|  |  |  | 12.75GHz to 23.45GHz) | TBD | U-shaped | 1.41 | TBD |
|  |  |  | 23.45GHz to 40.8GHz | TBD | U-shaped | 1.41 | TBD |
|  |  |  | 40.8GHz to 66GHz | TBD | U-shaped | 1.41 | TBD |
|  |  |  | 66GHz to 87GHz | TBD | U-shaped | 1.41 | TBD |
| NOTE 1: For Uncertainty assessment in tables across clause B and MTSU calculation, influence of the XPD uncertainty value for QZ ≤ 30cm is used also for 40cm. |  |  |  |  |  |  |  |

<<Unchanged sections skipped>>

B.2.2.15	Uncertainty of the absolute gain of the calibration antenna

See B.2.1.15.

The uncertainty value of uncertainty of the absolute gain of the calibration antenna is estimated as below table and used across clause B.

Table B.2.2.15-1: Uncertainty value for uncertainty of the absolute gain of the calibration antenna for IFF

| Power class | Test case | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| PC1, PC3, PC5, PC6 | Default | 0.60 | Normal | 2.00 | 0.30 |
|  | SE (40.8GHz to 66GHz) | 1.70 | Normal | 2.00 | 0.85 |
|  | SE (66GHz to 87GHz) | 1.70 | Normal | 2.00 | 0.85 |

<<Unchanged sections skipped>>

B.2.2.17	gNB emulator uncertainty

See B.2.1.17.

The uncertainty value of gNB emulator uncertainty is estimated as below table and used across clause B.

Table B.2.2.17-1: Uncertainty value for gNB emulator uncertainty for IFF

| Power class | Uncertainty value | Frequency range | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| PC1, PC5 | 2.9 | FR2a, FR2b | Normal | 2.00 | 1.45 |
| PC3 | 2.9 | FR2a, FR2b | Normal | 2.00 | 1.45 |
|  | 3.6 | FR2c | Normal | 2.00 | 1.8 |
| PC6 | 2.9 | FR2a, FR2b | Normal | 2.00 | 1.45 |

<<Unchanged sections skipped>>

B.2.2.21	Influence of the calibration antenna feed cable (Flexing cables, adapters, attenuators, connector repeatability)

See B.2.1.21.

The uncertainty value of influence of the calibration antenna feed cable is estimated as below table and used across clause B.

Table B.2.2.21-1: Uncertainty value for influence of the calibration antenna feed cable for IFF

| Power class | Test case | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| PC1, PC3, PC5, PC6 | Default | 0.14 | Normal | 2.00 | 0.07 |
|  | SE (40.8GHz to 66GHz) | 0.28 | Normal | 2.00 | 0.14 |
|  | SE (66GHz to 87GHz) | 0.28 | Normal | 2.00 | 0.14 |

<<Unchanged sections skipped>>

B.3.2	Uncertainty budget format and assessment for IFF

The uncertainty contributions that may impact the overall MU value are listed in Table B.3.2-1.

Table B.3.2-1: Uncertainty contributions for EIRP and TRP measurement

| UID |  | Description of uncertainty contribution |  | Details in clause |  |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 |  | Positioning misalignment |  | B.2.2.1 |  |
| 2 |  | Measure distance uncertainty |  | B.2.2.2 |  |
| 3 |  | Quality of Quiet Zone |  | B.2.2.3 |  |
| 4 |  | Mismatch |  | B.2.2.4 |  |
| 5 |  | Standing wave between the DUT and measurement antenna |  | B.2.2.5 |  |
| 6 |  | Uncertainty of the RF power measurement equipment |  | B.2.2.6 |  |
| 7 |  | Phase curvature |  | B.2.2.7 |  |
| 8 |  | Amplifier uncertainties |  | B.2.2.8 |  |
| 9 |  | Random uncertainty |  | B.2.2.9 |  |
| 10 |  | Influence of the XPD |  | B.2.2.10 |  |
| 11 |  | Insertion Loss Variation |  | B.2.2.11 |  |
| 12 |  | RF leakage (from measurement antenna to the receiver/transmitter) |  | B.2.2.12 |  |
| 13 |  | Influence of TRP measurement grid |  | B.2.2.22 |  |
| 14 |  | Influence of beam peak search grid |  | B.2.2.23 |  |
| 15 |  | Multiple measurement antenna uncertainty |  | B.2.2.25 |  |
| 16 |  | DUT repositioning |  | B.2.2.26 |  |
| 17 |  | Influence of spherical coverage grid |  | B.2.2.29 |  |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 18 |  | Mismatch |  | B.2.2.4 |  |
| 19 |  | Amplifier Uncertainties |  | B.2.2.8 |  |
| 20 |  | Misalignment of positioning System |  | B.2.2.13 |  |
| 21 |  | Uncertainty of the Network Analyzer |  | B.2.2.14 |  |
| 22 |  | Uncertainty of the absolute gain of the calibration antenna |  | B.2.2.15 |  |
| 23 |  | Positioning and pointing misalignment between the reference antenna and the measurement antenna |  | B.2.2.16 |  |
| 24 |  | Phase centre offset of calibration antenna |  | B.2.2.18 |  |
| 25 |  | Quality of quiet zone for calibration process |  | B.2.2.19 |  |
| 26 |  | Standing wave between reference calibration antenna and measurement antenna |  | B.2.2.20 |  |
| 27 |  | Influence of the calibration antenna feed cable |  | B.2.2.21 |  |
| 28 |  | Insertion Loss Variation |  | B.2.1.11 |  |
| Systematic uncertainties |  |  |  |  |  |
| 29 |  | Systematic error due to TRP calculation/quadrature |  | B.2.2.24 |  |
| 30 |  | Influence of noise |  | B.2.1.27 |  |
| 31 |  | Systematic error related to beam peak search |  | B.2.2.28 |  |

The uncertainty assessment tables are organized as follows:

-	For the purpose of uncertainty assessment, the radiating antenna aperture of the DUT is denoted as D

-	The uncertainty assessment has been derived for the case of Quiet Zone size ≤ [30 cm], f = {23.45GHz, 32.125GHz, 40.8GHz, 44.3GHz}, [P = maximum output power].

-	The uncertainty assessment for EIRP and TRP in normal temperature conditions is provided in Table B.3.2-2 for PC3 UEs and in Table B.3.2-6 for PC1 and PC5 UEs.

-	The uncertainty assessment for EIRP and TRP in extreme temperature conditions is provided in Table B.3.2-8 for PC3 UEs and in Table B.3.2-9 for PC1 and PC5 UEs.

-	The uncertainty assessment for Spherical coverage is provided in Table B.3.2-4 for PC3 UEs in Table B.3.2-7 for PC1 and PC5 UEs.

Table B.3.2-2: Uncertainty assessment for EIRP and TRP measurement (f=23.45GHz, 32.125GHz, 40.8GHz, 44.3GHz, Quiet Zone size ≤ 30 cm) for PC3 UEs and normal temperature condition

| UID |  | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |
| 1 |  | Positioning misalignment | 0.00 | Normal | 2.00 | 0.00 |  |
| 2 |  | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |  |
| 3 |  | Quality of Quiet Zone (NOTE 1) (23.45GHz <= f <= 40.8GHz) | 0.6 | Actual | 1.00 | 0.6 |  |
| 3 |  | Quality of Quiet Zone (NOTE 1) (40.8GHz < f <= 44.3GHz) | 0.7 | Actual | 1.00 | 0.7 |  |
| 4 |  | Mismatch (23.45GHz <= f <= 40.8GHz) | 1.30 | Actual | 1.00 | 1.30 |  |
| 4 |  | Mismatch (40.8GHz < f <= 44.3GHz) | 1.70 | Actual | 1.00 | 1.70 |  |
| 5 |  | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |  |
| 6 |  | Uncertainty of the RF power measurement equipment (NOTE 3) (23.45GHz <= f <= 40.8GHz) | 2.16 | Normal | 2.00 | 1.08 |  |
| 6 |  | Uncertainty of the RF power measurement equipment (NOTE 3) (40.8GHz < f <= 44.3GHz) | [2.16] | Normal | 2.00 | [1.08] |  |
| 7 |  | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |  |
| 8 |  | Amplifier uncertainties | 2.10 | Normal | 2.00 | 1.05 |  |
| 9 |  | Random uncertainty | 0.50 | Normal | 2.00 | 0.25 |  |
| 10 |  | Influence of the XPD | 0.01 | U-shaped | 1.41 | 0.00 |  |
| 11 |  | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |  |
| 12 |  | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |  |
| 13 |  | Influence of TRP measurement grid (NOTE 4) | 0.25 | Actual | 1 | 0.25 |  |
| 14 |  | Influence of beam peak search grid (NOTE 5) | 0.00 | Actual | 1 | 0.00 |  |
| 15 |  | Multiple measurement antenna uncertainty (NOTE 9) | 0.15 | Actual | 1 | 0.15 |  |
| 16 |  | DUT repositioning | 0.00 (NOTE 4) 0.08 (NOTE 5) | Rectangular | 1.73 | 0.00 (NOTE 4) 0.05 (NOTE 5) |  |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |
| 17 |  | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |  |
| 18 |  | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |  |
| 19 |  | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |  |
| 20 |  | Uncertainty of the Network Analyzer (23.45GHz <= f <= 40.8GHz) | 1.50 | Normal | 2.00 | 0.75 |  |
| 20 |  | Uncertainty of the Network Analyzer (40.8GHz < f <= 44.3GHz) | 1.70 | Normal | 2.00 | 0.85 |  |
| 21 |  | Uncertainty of the absolute gain of the calibration antenna | 0.60 | Normal | 2.00 | 0.30 |  |
| 22 |  | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.01 | Rectangular | 1.73 | 0.00 |  |
| 23 |  | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |  |
| 24 |  | Quality of quiet zone for calibration process (NOTE 1) (23.45GHz <= f <= 40.8GHz) | 0.4 | Actual | 1.00 | 0.4 |  |
| 24 |  | Quality of quiet zone for calibration process (NOTE 1) (40.8GHz < f <= 44.3GHz) | 0.5 | Actual | 1.00 | 0.5 |  |
| 25 |  | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |  |
| 26 |  | Influence of the calibration antenna feed cable | 0.14 | Normal | 2.00 | 0.07 |  |
| 27 |  | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |  |
|  |  | Systematic uncertainties (NOTE 6) |  |  |  | Value |  |
| 28 |  | Systematic error due to TRP calculation/quadrature (NOTE 4) |  |  |  | 0.00 |  |
| 29 |  | Influence of noise (23.45GHz <= f <= 32.125GHz) |  |  |  | 0.1 |  |
| 29 |  | Influence of noise (32.125GHz < f <= 40.8GHz) |  |  |  | 0.3 |  |
| 29 |  | Influence of noise (40.8GHz < f <= 44.3GHz) |  |  |  | 0.3 |  |
| 30 |  | Systematic error related to beam peak search (NOTE 5) |  |  |  | 0.5 |  |
| Total measurement uncertainty |  |  |  |  |  | Value |  |
| EIRP Expanded uncertainty (23.45GHz <= f <= 32.125GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  | 5.08 |  |
| EIRP Expanded uncertainty (32.125GHz < f <= 40.8GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  | 5.28 |  |
| EIRP Expanded uncertainty (40.8GHz < f <= 44.3GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  | [5.91] |  |
| TRP Expanded uncertainty (23.45GHz <= f <= 32.125GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  | 4.61 |  |
| TRP Expanded uncertainty (32.125GHz < f <= 40.8GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  | 4.81 |  |
| TRP Expanded uncertainty (40.8GHz < f <= 44.3GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  | [5.44] |  |
| NOTE 1: The quality of quiet zone is the same for EIRP and TRP. Value based on procedure defined in clause D.2 of TR 38.810 for Quiet Zone size less or equal to 30 cm. NOTE 2: The analysis was done only for the case of operating at max output power, in-band, non-CA. NOTE 3: The assessment assumes maximum DUT output power. NOTE 4: This contributor shall only be considered for TRP measurements. NOTE 5: This contributor shall only be considered for EIRP measurements. NOTE 6: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 7: Void. NOTE 8: Void NOTE 9: Applies to the system which has a structure of mechanical feed antenna positioning. |  |  |  |  |  |  |  |

Table B.3.2-3: Void

Table B.3.2-4: Uncertainty assessment for Spherical coverage measurement (f=23.45GHz, 32.125GHz, 40.8GHz, 44.3GHz, Quiet Zone size ≤ 30 cm) for PC3 UEs and normal temperature condition

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment | 0.00 | Normal | 2.00 | 0.00 |
| 2 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 3 | Quality of Quiet Zone (NOTE 1) (23.45GHz <= f <= 40.8GHz) | 0.6 | Actual | 1.00 | 0.6 |
| 3 | Quality of Quiet Zone (NOTE 1) (40.8GHz < f <= 44.3GHz) | 0.7 | Actual | 1.00 | 0.7 |
| 4 | Mismatch (23.45GHz <= f <= 40.8GHz) | 1.30 | Actual | 1.00 | 1.30 |
| 4 | Mismatch (40.8GHz < f <= 44.3GHz) | 1.70 | Actual | 1.00 | 1.70 |
| 5 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | Uncertainty of the RF power measurement equipment (NOTE 3) (23.45GHz <= f <= 40.8GHz) | 2.16 | Normal | 2.00 | 1.08 |
| 6 | Uncertainty of the RF power measurement equipment (NOTE 3) (40.8GHz < f <= 44.3GHz) | 3.6 | Normal | 2.00 | 1.8 |
| 7 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 8 | Amplifier uncertainties | 2.1 | Normal | 2.00 | 1.05 |
| 9 | Random uncertainty | 0.50 | Normal | 2.00 | 0.25 |
| 10 | Influence of the XPD | 0.01 | U-shaped | 1.41 | 0.00 |
| 11 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 13 | Multiple measurement antenna uncertainty (NOTE 5) | 0.15 | Actual | 1 | 0.15 |
| 14 | DUT repositioning | 0.00 | Rectangular | 1.73 | 0.00 |
| 15 | Influence of spherical coverage grid | 0.12 | Actual | 1 | 0.12 |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 16 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 17 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 18 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 19 | Uncertainty of the Network Analyzer (23.45GHz <= f <= 40.8GHz) | 1.50 | Normal | 2.00 | 0.75 |
| 19 | Uncertainty of the Network Analyzer (40.8GHz < f <= 44.3GHz) | 1.70 | Normal | 2.00 | 0.85 |
| 20 | Uncertainty of the absolute gain of the calibration antenna | 0.60 | Normal | 2.00 | 0.30 |
| 21 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.01 | Rectangular | 1.73 | 0.00 |
| 22 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 23 | Quality of quiet zone for calibration process (NOTE 1) (23.45GHz <= f <= 40.8GHz) | 0.4 | Actual | 1.00 | 0.4 |
| 23 | Quality of quiet zone for calibration process (NOTE 1) (40.8GHz < f <= 44.3GHz) | 0.5 | Actual | 1.00 | 0.5 |
| 24 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 25 | Influence of the calibration antenna feed cable | 0.14 | Normal | 2.00 | 0.07 |
| 26 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
|  | Systematic uncertainties (NOTE 4) |  |  |  | Value |
| 27 | Influence of noise (23.45GHz <= f <= 32.125GHz) |  |  |  | 0.3 |
| 27 | Influence of noise (32.125GHz < f <= 40.8GHz) |  |  |  | 0.9 |
| 27 | Influence of noise (40.8GHz < f <= 44.3GHz) |  |  |  | TBD |
| Total measurement uncertainty |  |  |  |  | Value |
| Spherical coverage Expanded uncertainty (23.45GHz <= f <= 32.125GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 4.78 |
| Spherical coverage Expanded uncertainty (32.125GHz < f <= 40.8GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 5.38 |
| Spherical coverage Expanded uncertainty (40.8GHz < f <= 44.3GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | TBD |
| NOTE 1: The quality of quiet zone is the same for EIRP and TRP. Value based on procedure defined in clause D.2 of TR 38.810 for Quiet Zone size less or equal to 30 cm. NOTE 2: The analysis was done only for the case of operating at max output power, in-band, non-CA. NOTE 3: The assessment assumes maximum DUT output power. NOTE 4: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 5: Applies to the system which has a structure of mechanical feed antenna positioning. |  |  |  |  |  |

<<Unchanged sections skipped>>

B.4	UE maximum output power for modulation / channel bandwidth

Following tables summarize the MU threshold for EIRP measurements for UE maximum output power for modulation / channel bandwidth (a.k.a Maximum Power Reduction/MPR). The origin MU values for different test setups with varies parameters can be found in following clauses.

Table B.4-1: MU threshold for EIRP measurement for UE maximum output power for modulation / channel bandwidth

| Power Class | Frequency | MBW | Power (NOTE2) | Threshold MU value for NTC [dB] (NOTE1) | Threshold MU value for ETC [dB] (NOTE1) |
| --- | --- | --- | --- | --- | --- |
| PC3 | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = Max Output Power –MBR - MPR – T(MPR) | 5.11 | 5.38 |
|  | 32.125GHz < f <= 40.8GHz |  |  | 5.29 | 5.56 |
|  | 40.8GHz < f <= 44.3GHz |  |  | TBD | TBD |
| PC1 | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = Max Output Power –MPR – T(MPR) | 5.33 | 5.60 |
|  | 32.125GHz < f <= 40.8GHz |  |  | 5.50 | 5.77 |
| PC5 | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = Max Output Power –MPR – T(MPR) | 5.33 | 5.60 |
| NOTE 1: Total EIRP Expanded MU for IFF for Quiet Zone size ≤30cm in Table B.4.2-2 and Table B.4.2-3 for PC3 UEs and Table B.4.2-4 and Table B.4.2-5 for PC1 and PC5 UEs. NOTE 2: Max output power level for device with corresponding power class. |  |  |  |  |  |

<<Unchanged sections skipped>>

B.4.2	Uncertainty budget format and assessment for IFF

The uncertainty contributions that may impact the overall MU value are listed in Table B.4.2-1.

Table B.4.2-1: Uncertainty contributions for EIRP and TRP measurement

| UID | Description of uncertainty contribution | Details in annex |
| --- | --- | --- |
| Same as Stage 2 of Table 3.2-1 for EIRP |  |  |

The uncertainty assessment tables are organized as follows:

-	For the purpose of uncertainty assessment, the radiating antenna aperture of the DUT is denoted as D

-	The uncertainty assessment has been derived for the case of Quiet Zone size ≤ 30 cm, f = {23.45GHz, 32.125GHz, 40.8GHz, 44.3GHz}, P = maximum output power – MPR – T(MPR).

-	The uncertainty assessment for EIRP and TRP is provided in Table B.3.2-2 for PC3 UEs and in Table B.4.2-x for PC1 UEs.

Table B.4.2-2: Uncertainty assessment for EIRP measurement (f=23.45GHz, 32.125GHz, 40.8GHz, 44.3GHz, Quiet Zone size ≤ 30 cm) for PC3 UEs and normal temperature condition

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 to 16 | Same as Stage 2 of Table B.3.2-2 for EIRP |  |  |  |  |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 17 to 27 | Same as Stage 1 of Table B.3.2-2 for EIRP |  |  |  |  |
|  | Systematic uncertainties |  |  |  | Value |
| 28 | Systematic error due to TRP calculation/quadrature |  |  |  | N/A |
| 29 | Influence of noise (23.45GHz <= f <= 32.125GHz) |  |  |  | 0.13 |
| 29 | Influence of noise (32.125GHz < f <= 40.8GHz) |  |  |  | 0.31 |
| 29 | Influence of noise (40.8GHz < f <= 44.3GHz) |  |  |  | TBD |
| 30 | Systematic error related to beam peak search |  |  |  | Same as Table B.3.2-2 |
| Total measurement uncertainty |  |  |  |  | Value |
| EIRP Expanded uncertainty (23.45GHz <= f <= 32.125GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 5.11 |
| EIRP Expanded uncertainty (32.125GHz < f <= 40.8GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 5.29 |
| EIRP Expanded uncertainty (40.8GHz < f <= 44.3GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | TBD |
| NOTE 1: The assessment assumes maximum DUT output power – MBR - MPR – T(MPR) |  |  |  |  |  |

<<Unchanged sections skipped>>

B.7.1	Uncertainty budget format and assessment for DFF

The uncertainty contributions that may impact the overall MU value are listed in Table B.7.1-1.

Table B.7.1-1: Uncertainty contributions for EIRP measurement

| UID | Description of uncertainty contribution | Details in annex |
| --- | --- | --- |
| Stage 2: DUT measurement |  |  |
| 1 | Positioning misalignment | B.2.1.1 |
| 2 | Measure distance uncertainty | B.2.1.2 |
| 3 | Quality of quiet zone | B.2.1.3 |
| 4 | Mismatch | B.2.1.4 |
| 5 | Standing Wave Between the DUT and measurement antenna | B.2.1.5 |
| 6 | Uncertainty of the RF power measurement equipment | B.2.1.6 |
| 7 | Phase curvature | B.2.1.7 |
| 8 | Amplifier uncertainties | B.2.1.8 |
| 9 | Random uncertainty | B.2.1.9 |
| 10 | Influence of the XPD | B.2.1.10 |
| 11 | Insertion Loss Variation | B.2.1.11 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | B.2.1.12 |
| 13 | Influence of beam peak search grid | B.2.1.23 |
| 14 | Multiple measurement antenna uncertainty | B.2.1.25 |
| 15 | DUT repositioning | B.2.1.26 |
| Stage 1: Calibration measurement |  |  |
| 16 | Mismatch | B.2.1.4 |
| 17 | Amplifier uncertainties | B.2.1.8 |
| 18 | Misalignment of positioning System | B.2.1.13 |
| 19 | Uncertainty of the Network Analyzer | B.2.1.14 |
| 20 | Uncertainty of the absolute gain of the calibration antenna | B.2.1.15 |
| 21 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | B.2.1.16 |
| 22 | Phase centre offset of calibration antenna | B.2.1.18 |
| 23 | Quality of quiet zone for calibration process | B.2.1.19 |
| 24 | Standing wave between reference calibration antenna and measurement antenna | B.2.1.20 |
| 25 | Influence of the calibration antenna feed cable | B.2.1.21 |
| 26 | Insertion Loss Variation | B.2.1.11 |
| Systematic uncertainties |  |  |
| 27 | Systematic error related to beam peak search | B.2.1.28 |
| 28 | Influence of noise | B.2.1.27 |

The uncertainty assessment tables are organized as follows:

-	For the purpose of uncertainty assessment, the radiating antenna aperture of the DUT is denoted as D

-	The uncertainty assessment has been derived for the case of D = [5 cm], f = {23.45 GHz, 32.125 GHz, 40.8 GHz, 44.3 GHz}, P = [Minimum output power].

-	The uncertainty assessment for EIRP is provided in Table B.7.1-2.

Table B.7.1-2: Uncertainty assessment for EIRP measurement (f=TBD, D=TBD)

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment |  |  |  |  |
| 2 | Measure distance uncertainty |  |  |  |  |
| 3 | Quality of quiet zone (NOTE 2) |  |  |  |  |
| 4 | Mismatch (NOTE 3) |  |  |  |  |
| 5 | Standing Wave Between the DUT and measurement antenna |  |  |  |  |
| 6 | Uncertainty of the RF power measurement equipment (NOTE 4) |  |  |  |  |
| 7 | Phase curvature |  |  |  |  |
| 8 | Amplifier uncertainties |  |  |  |  |
| 9 | Random uncertainty |  |  |  |  |
| 10 | Influence of the XPD |  |  |  |  |
| 11 | Insertion Loss Variation |  |  |  |  |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) |  |  |  |  |
| 13 | Influence of beam peak search grid (NOTE 6) |  |  |  |  |
| 14 | Multiple measurement antenna uncertainty |  |  |  |  |
| 15 | DUT repositioning |  | Actual | 1 |  |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 16 | Mismatch |  |  |  |  |
| 17 | Amplifier uncertainties |  |  |  |  |
| 18 | Misalignment of positioning System |  |  |  |  |
| 19 | Uncertainty of the Network Analyzer |  |  |  |  |
| 20 | Uncertainty of the absolute gain of the calibration antenna |  |  |  |  |
| 21 | Positioning and pointing misalignment between the reference antenna and the measurement antenna |  |  |  |  |
| 22 | Phase centre offset of calibration antenna |  |  |  |  |
| 23 | Quality of quiet zone for calibration process (NOTE 2) |  |  |  |  |
| 24 | Standing wave between reference calibration antenna and measurement antenna |  |  |  |  |
| 25 | Influence of the calibration antenna feed cable |  |  |  |  |
| 26 | Insertion Loss Variation |  |  |  |  |
| EIRP Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  |
| Systematic uncertainties (NOTE 7) |  |  |  |  | Value |
| 27 | Systematic error related to beam peak search |  |  |  |  |
| 28 | Influence of noise |  |  |  |  |
| Total measurement uncertainty |  |  |  |  |  |
| EIRP total measurement uncertainty [dB] |  |  |  |  |  |
| NOTE 1: The analysis was done only for the case of operating at Minimum output power, in-band, non-CA. NOTE 2: The assessment assumes DUT Minimum output power. NOTE 3: This contributor shall only be considered for EIRP measurements. NOTE 4: Void NOTE 5: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 6: Void. NOTE 7: Void NOTE 8: Value based on procedure defined in Annex D.2 of TR 38.810 for Quiet Zone size less or equal to 30 cm. NOTE 9: Applies to the system which has a structure of mechanical feed antenna positioning. |  |  |  |  |  |

B.7.2	Uncertainty budget format and assessment for IFF

The uncertainty contributions that may impact the overall MU value are listed in Table B.7.2-1.

Table B.7.2-1: Uncertainty contributions for EIRP measurement

| UID | Description of uncertainty contribution | Details in annex |
| --- | --- | --- |
| Stage 2: DUT measurement |  |  |
| 1 | Positioning misalignment | # B.2.2.1 |
| 2 | Measure distance uncertainty | B.2.2.2 |
| 3 | Quality of Quiet Zone | B.2.2.3 |
| 4 | Mismatch | B.2.2.4 |
| 5 | Standing wave between the DUT and measurement antenna | B.2.2.5 |
| 6 | Uncertainty of the RF power measurement equipment | B.2.2.6 |
| 7 | Phase curvature | B.2.2.7 |
| 8 | Amplifier uncertainties | B.2.2.8 |
| 9 | Random uncertainty | B.2.2.9 |
| 10 | Influence of the XPD | B.2.2.10 |
| 11 | Insertion Loss Variation | B.2.2.11 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | B.2.2.12 |
| 13 | Influence of beam peak search grid | B.2.2.23 |
| 14 | Multiple measurement antenna uncertainty | B.2.2.25 |
| 15 | DUT repositioning | B.2.2.26 |
| Stage 1: Calibration measurement |  |  |
| 16 | Mismatch | B.2.2.4 |
| 17 | Amplifier Uncertainties | B.2.2.8 |
| 18 | Misalignment of positioning System | B.2.2.13 |
| 19 | Uncertainty of the Network Analyzer | B.2.2.14 |
| 20 | Uncertainty of the absolute gain of the calibration antenna | B.2.2.15 |
| 21 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | B.2.2.16 |
| 22 | Phase centre offset of calibration antenna | B.2.2.18 |
| 23 | Quality of quiet zone for calibration process | B.2.2.19 |
| 24 | Standing wave between reference calibration antenna and measurement antenna | B.2.2.20 |
| 25 | Influence of the calibration antenna feed cable | B.2.2.21 |
| 26 | Insertion Loss Variation | B.2.2.11 |
| Systematic uncertainties |  |  |
| 27 | Systematic error related to beam peak search | B.2.2.28 |
| 28 | Influence of noise | B.2.2.27 |

The uncertainty assessment tables are organized as follows:

-	For the purpose of uncertainty assessment, the radiating antenna aperture of the DUT is denoted as D

-	The uncertainty assessment has been derived for the case of Quiet Zone size ≤ 30 cm, f = {23.45GHz, 32.125GHz, 40.8GHz, 44.3GHz}, P = Minimum output power.

-	The uncertainty assessment for EIRP is provided in Table B.7.2-2 for PC3 UEs and in Table B.7.2-3 for PC1 UEs.

Table B.7.2-2: Uncertainty assessment for EIRP measurement (f=23.45GHz, 32.125GHz, 40.8GHz, 44.3GHz, Quiet Zone size ≤ 30 cm) for PC3 UEs and normal temperature condition

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment | 0.00 | Normal | 2.00 | 0.00 |
| 2 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 3 | Quality of Quiet Zone (NOTE 8) (23.45GHz <= f <= 40.8GHz) | 0.6 | Actual | 1.00 | 0.6 |
| 3 | Quality of Quiet Zone (NOTE 8) (40.8GHz < f <= 44.3GHz) | 0.7 | Actual | 1.00 | 0.7 |
| 4 | Mismatch (NOTE 1) (23.45GHz <= f <= 40.8GHz) | 1.30 | Actual | 1.00 | 1.30 |
| 4 | Mismatch (NOTE 1) (40.8GHz < f <= 44.3GHz) | 1.70 | Actual | 1.00 | 1.70 |
| 5 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | Uncertainty of the RF power measurement equipment (NOTE 2) (23.45GHz <= f <= 40.8GHz) | 2.50 | Normal | 2.00 | 1.25 |
| 6 | Uncertainty of the RF power measurement equipment (NOTE 2) (40.8GHz < f <= 44.3GHz) | 3.6 | Normal | 2.00 | 1.8 |
| 7 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 8 | Amplifier uncertainties | 2.10 | Normal | 2.00 | 1.05 |
| 9 | Random uncertainty | 0.50 | Normal | 2.00 | 0.25 |
| 10 | Influence of the XPD | 0.01 | U-shaped | 1.41 | 0.00 |
| 11 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 13 | Influence of beam peak search grid (NOTE 3) | 0.00 | Actual | 1 | 0.00 |
| 14 | Multiple measurement antenna uncertainty (NOTE 9) | 0.15 | Actual | 1 | 0.15 |
| 15 | DUT repositioning (NOTE 3) | 0.08 | Rectangular | 1.73 | 0.05 |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 16 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 17 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 18 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 19 | Uncertainty of the Network Analyzer (23.45GHz <= f <= 40.8GHz) | 1.50 | Normal | 2.00 | 0.75 |
| 19 | Uncertainty of the Network Analyzer (40.8GHz < f <= 44.3GHz) | 1.70 | Normal | 2.00 | 0.85 |
| 20 | Uncertainty of the absolute gain of the calibration antenna | 0.60 | Normal | 2.00 | 0.30 |
| 21 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.01 | Rectangular | 1.73 | 0.00 |
| 22 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 23 | Quality of quiet zone for calibration process (NOTE 8) (23.45GHz <= f <= 40.8GHz) | 0.4 | Actual | 1.00 | 0.4 |
| 23 | Quality of quiet zone for calibration process (NOTE 8) (40.8GHz < f <= 44.3GHz) | 0.5 | Actual | 1.00 | 0.5 |
| 24 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 25 | Influence of the calibration antenna feed cable | 0.14 | Normal | 2.00 | 0.07 |
| 26 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
|  | Systematic uncertainties (NOTE 5) |  |  |  | Value |
| 27 | Systematic error related to beam peak search |  |  |  | 0.5 |
| 28 | Influence of noise (23.45GHz <= f <= 32.125GHz) |  |  |  | 1.0 |
| 29 | Influence of noise (32.125GHz < f <= 40.8GHz) |  |  |  | 1.0 |
| 28 | Influence of noise (40.8GHz < f <= 44.3GHz) |  |  |  | TBD |
| Total measurement uncertainty |  |  |  |  | Value |
| EIRP Expanded uncertainty (23.45GHz <= f <= 32.125GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 6.15 |
| EIRP Expanded uncertainty (32.125GHz < f <= 40.8GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 6.15 |
| EIRP Expanded uncertainty (40.8GHz < f <= 44.3GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | TBD |
| NOTE 1: The analysis was done only for the case of operating at Minimum output power, in-band, non-CA. NOTE 2: The assessment assumes DUT Minimum output power. NOTE 3: This contributor shall only be considered for EIRP measurements. NOTE 4: Void NOTE 5: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 6: Void. NOTE 7: Void NOTE 8: Value based on procedure defined in Annex D.2 of TR 38.810 for Quiet Zone size less or equal to 30 cm. NOTE 9: Applies to the system which has a structure of mechanical feed antenna positioning. |  |  |  |  |  |

<<Unchanged sections skipped>>

B.17	Adjacent Channel Leakage Ratio

Editor’s Note: MU value analysis for PC1, 2 and 4 are not complete.

Following tables summarize the MU threshold for EIRP measurements for Adjacent Channel Leakage Ratio. The origin MU values for different test setups can be found in following subclauses.

Table B.17-1: MU threshold for TRP measurement for Spectrum emission mask

| Power Class | Frequency | MBW | Power | Threshold MU value (NOTE 1) |
| --- | --- | --- | --- | --- |
| PC3 | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = Max Output Power | TBD |
|  | 32.125GHz < f <= 40.8GHz |  |  | TBD |
|  | 40.8GHz < f <= 44.3GHz |  |  | TBD |
| PC1 | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = Max Output Power | TBD |
|  | 32.125GHz < f <= 40.8GHz |  |  | TBD |
| PC5 | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = Max Output Power | TBD |
| NOTE 1: Total Expanded MU for IFF for a Quiet Zone size ≤ 30 cm in Table B.17.2-2 for PC3 UEs and B.17.2-3 for PC1 UEs. |  |  |  |  |

Table B.17-1B: MU threshold for EIRP measurement for Adjacent Channel Leakage Ratio

| Power Class | Frequency | CBW | Power | Threshold MU value for NTC and ETC (NOTE 1) |
| --- | --- | --- | --- | --- |
| PC3 | 23.45GHz <= f <= 32.125GHz | 50MHz | P = Max Output Power | 5.63 |
|  |  | 100MHz |  | 6.09 |
|  |  | 200MHz |  | 6.09 (NOTE5) |
|  |  | 400MHz |  | 6.09 (NOTE2) |
|  | 32.125GHz < f <= 40.8GHz | 50MHz | P = Max Output Power | 6.09 (NOTE7) |
|  |  | 100MHz |  | 6.09 (NOTE6) |
|  |  | 200MHz |  | 6.09 (NOTE3) |
|  |  | 400MHz |  | 6.09 (NOTE4) |
|  | 40.8GHz < f <= 44.3GHz | 50MHz | P = Max Output Power | TBD |
|  |  | 100MHz |  | TBD |
|  |  | 200MHz |  | TBD |
|  |  | 400MHz |  | TBD |
| PC1 | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = Max Output Power | 6.04 |
|  | 32.125GHz < f <= 40.8GHz | BW <= 400MHz | P = Max Output Power | 6.04 |
| PC5 | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = Max Output Power | 6.04 |
| NOTE 1: Total Expanded MU for IFF for a Quiet Zone size ≤ 30 cm in Table B.17.2-2 for PC3 UEs and Table B.17.2-3 for PC1 and PC5 UEs. NOTE 2: This value is based on the relaxation of (MPR – 3.0) dB for MPR > 3.0dB. NOTE 3: Not applicable for MPR > 3.5dB NOTE 4: Not applicable for MPR > 2.0dB NOTE 5: This value is based on the relaxation of (MPR – 5.0) dB for MPR > 5.0dB. NOTE 6: Not applicable for MPR > 5.0dB NOTE 7: Not applicable for MPR >7. 5 dB |  |  |  |  |

B.17.1	Uncertainty budget format and assessment for DFF

The uncertainty contributions that may impact the overall MU value are listed in Table B.17.1-1.

Table B.17.1-1: Uncertainty contributions for EIRP measurement

| UID | Description of uncertainty contribution | Details in annex |
| --- | --- | --- |
| Stage 2: DUT measurement |  |  |
| 1 | Positioning misalignment | B.2.1.1 |
| 2 | Measure distance uncertainty | B.2.1.2 |
| 3 | Quality of quiet zone | B.2.1.3 |
| 4 | Mismatch | B.2.1.4 |
| 5 | Standing Wave Between the DUT and measurement antenna | B.2.1.5 |
| 6 | Uncertainty of the RF power measurement equipment | B.2.1.6 |
| 7 | Phase curvature | B.2.1.7 |
| 8 | Amplifier uncertainties | B.2.1.8 |
| 9 | Random uncertainty | B.2.1.9 |
| 10 | Influence of the XPD | B.2.1.10 |
| 11 | Insertion Loss Variation | B.2.1.11 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | B.2.1.12 |
| 13 | Influence of TRP measurement grid | B.2.1.22 |
| 14 | Influence of beam peak search grid | B.2.1.23 |
| 15 | Multiple measurement antenna uncertainty | B.2.1.25 |
| 16 | DUT repositioning | B.2.1.26 |
| Stage 1: Calibration measurement |  |  |
| 17 | Mismatch | B.2.1.4 |
| 18 | Amplifier uncertainties | B.2.1.8 |
| 19 | Misalignment of positioning System | B.2.1.13 |
| 20 | Uncertainty of the Network Analyzer | B.2.1.14 |
| 21 | Uncertainty of the absolute gain of the calibration antenna | B.2.1.15 |
| 22 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | B.2.1.16 |
| 23 | Phase centre offset of calibration antenna | B.2.1.18 |
| 24 | Quality of quiet zone for calibration process | B.2.1.19 |
| 25 | Standing wave between reference calibration antenna and measurement antenna | B.2.1.20 |
| 26 | Influence of the calibration antenna feed cable | B.2.1.21 |
| 27 | Insertion Loss Variation | B.2.1.11 |
| Systematic uncertainties |  |  |
| 28 | Systematic error due to TRP calculation/quadrature | B.2.1.24 |
| 29 | Influence of noise | B.2.1.27 |

The uncertainty assessment tables are organized as follows:

-	For the purpose of uncertainty assessment, the radiating antenna aperture of the DUT is denoted as D

-	The uncertainty assessment has been derived for the case of D = [5 cm], f = {23.45 GHz, 32.125 GHz, 40.8 GHz, 44.3 GHz}, P = [Maximum output power].

-	The uncertainty assessment for TRP is provided in Table B.17.1-2.

Table B.17.1-2: Uncertainty assessment for EIRP measurement (f=TBD, D=TBD)

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment |  |  |  |  |
| 2 | Measure distance uncertainty |  |  |  |  |
| 3 | Quality of quiet zone (NOTE 2) |  |  |  |  |
| 4 | Mismatch (NOTE 3) |  |  |  |  |
| 5 | Standing Wave Between the DUT and measurement antenna |  |  |  |  |
| 6 | Uncertainty of the RF power measurement equipment (NOTE 4) |  |  |  |  |
| 7 | Phase curvature |  |  |  |  |
| 8 | Amplifier uncertainties |  |  |  |  |
| 9 | Random uncertainty |  |  |  |  |
| 10 | Influence of the XPD |  |  |  |  |
| 11 | Insertion Loss Variation |  |  |  |  |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) |  |  |  |  |
| 13 | Influence of TRP measurement grid (NOTE 5) |  |  |  |  |
| 14 | Influence of beam peak search grid (NOTE 6) |  |  |  |  |
| 15 | Multiple measurement antenna uncertainty |  |  |  |  |
| 16 | DUT repositioning |  |  |  |  |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 17 | Mismatch |  |  |  |  |
| 18 | Amplifier uncertainties |  |  |  |  |
| 19 | Misalignment of positioning System |  |  |  |  |
| 20 | Uncertainty of the Network Analyzer |  |  |  |  |
| 21 | Uncertainty of the absolute gain of the calibration antenna |  |  |  |  |
| 22 | Positioning and pointing misalignment between the reference antenna and the measurement antenna |  |  |  |  |
| 23 | Phase centre offset of calibration antenna |  |  |  |  |
| 24 | Quality of quiet zone for calibration process (NOTE 2) |  |  |  |  |
| 25 | Standing wave between reference calibration antenna and measurement antenna |  |  |  |  |
| 26 | Influence of the calibration antenna feed cable |  |  |  |  |
| 27 | Insertion Loss Variation |  |  |  |  |
| EIRP Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  |
| Systematic uncertainties (NOTE 7) |  |  |  |  | Value |
| 28 | Systematic error due to TRP calculation/quadrature (NOTE 5) |  |  |  |  |
| 29 | Influence of noise |  |  |  |  |
| Total measurement uncertainty |  |  |  |  |  |
| TRP total measurement uncertainty [dB] |  |  |  |  |  |
| NOTE 1: The impact of phase variation on EIRP is FFS. NOTE 2: The quality of quiet zone is different for EIRP and TRP. For TRP, the standard uncertainty is FFS; for EIRP, the standard uncertainty of quiet zone is FFS. NOTE 3: The analysis was done only for the case of operating at max output power, in-band, non-CA. NOTE 4: The assessment assumes maximum DUT output power. NOTE 5: This contributor shall only be considered for TRP measurements. NOTE 6: This contributor shall only be considered for EIRP measurements. NOTE 7: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 8: Void |  |  |  |  |  |

B.17.2	Uncertainty budget format and assessment for IFF

The uncertainty contributions that may impact the overall MU value are listed in Table B.17.2-1.

Table B.17.2-1: Uncertainty contributions for EIRP measurement

| UID | Description of uncertainty contribution | Details in clause |
| --- | --- | --- |
| Stage 2: DUT measurement |  |  |
| 1 | Quality of Quiet Zone | B.2.2.3 |
| 2 | Mismatch | B.2.2.4 |
| 3 | Standing wave between the DUT and measurement antenna | B.2.2.5 |
| 4 | Uncertainty of the RF power measurement equipment | B.2.2.6 |
| 5 | Phase curvature | B.2.2.7 |
| 6 | Amplifier uncertainties | B.2.2.8 |
| 7 | Random uncertainty | B.2.2.9 |
| 8 | Influence of the XPD | B.2.2.10 |
| 9 | RF leakage (from measurement antenna to the receiver/transmitter) | B.2.2.12 |
| 10 | Multiple measurement antenna uncertainty | B.2.2.25 |
| Stage 1: Calibration measurement |  |  |
| 11 | Mismatch | B.2.2.4 |
| 12 | Amplifier Uncertainties | B.2.2.8 |
| 13 | Misalignment of positioning System | B.2.2.13 |
| 14 | Uncertainty of the Network Analyzer | B.2.2.14 |
| 15 | Uncertainty of the absolute gain of the calibration antenna | B.2.2.15 |
| 16 | Phase centre offset of calibration antenna | B.2.2.18 |
| 17 | Quality of quiet zone for calibration process | B.2.2.19 |
| 18 | Standing wave between reference calibration antenna and measurement antenna | B.2.2.20 |
| 19 | Influence of the calibration antenna feed cable | B.2.2.21 |
| 20 | Insertion Loss Variation | B.2.2.11 |
| Systematic uncertainties |  |  |
| 21 | Influence of noise | B.2.2.27 |

The uncertainty assessment tables are organized as follows:

-	For the purpose of uncertainty assessment, the radiating antenna aperture of the DUT is denoted as D

-	The uncertainty assessment has been derived for the case of Quiet Zone size ≤ 30 cm, f = {23.45GHz, 32.125GHz, 40.8GHz, 44.3GHz}, P = Maximum output power - MPR – MBR(Multi-band relaxation).

-	The uncertainty assessment for EIRP is provided in Table B.17.2-2 for PC3 UEs and Table B.17.2-3 for PC1 UEs.

Table B.17.2-2: Uncertainty assessment for EIRP measurement (f=23.45GHz, 32.125GHz, 40.8GHz, 44.3GHz, Quiet Zone size ≤ 30 cm) for PC3 UEs and normal and extreme temperature condition

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Quality of Quiet Zone (NOTE 10) | 0.52 | Actual | 1.00 | 0.52 |
| 2 | Mismatch (NOTE 2) (23.45GHz <= f <= 40.8GHz) | 1.84 | Actual | 1.00 | 1.84 |
| 2 | Mismatch (NOTE 2) (40.8GHz < f <= 44.3GHz) | 2.40 | Actual | 1.00 | 2.40 |
| 3 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 4 | Uncertainty of the RF power measurement equipment (NOTE 3, 7) (23.45GHz <= f <= 40.8GHz) | 2.16 | Normal | 2.00 | 1.08 |
| 4 | Uncertainty of the RF power measurement equipment (NOTE 3, 7) (40.8GHz < f <= 44.3GHz) | 3.6 | Normal | 2.00 | 1.8 |
| 5 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | Amplifier uncertainties | 2.1 | Normal | 2.00 | 1.05 |
| 7 | Random uncertainty | 0.50 | Normal | 2.00 | 0.25 |
| 8 | Influence of the XPD | 0.00 | U-shaped | 1.41 | 0.00 |
| 9 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 10 | Multiple measurement antenna uncertainty (NOTE 9) | 0.0 | Actual | 1 | 0.0 |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 11 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 12 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 13 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 14 | Uncertainty of the Network Analyzer (23.45GHz <= f <= 40.8GHz) | 1.5 | Normal | 2.00 | 0.75 |
| 14 | Uncertainty of the Network Analyzer (40.8GHz < f <= 44.3GHz) | 1.7 | Normal | 2.00 | 0.85 |
| 15 | Uncertainty of the absolute gain of the calibration antenna | 0.60 | Normal | 2.00 | 0.30 |
| 16 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 17 | Quality of quiet zone for calibration process (NOTE 10) | 0.32 | Actual | 1.00 | 0.32 |
| 18 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 19 | Influence of the calibration antenna feed cable | 0.00 | Normal | 2.00 | 0.00 |
| 20 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| EIRP Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] (23.45GHz <= f <= 40.8GHz) |  |  |  |  | 5.09 |
| EIRP Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] (40.8GHz < f <= 44.3GHz) |  |  |  |  | 6.61 |
|  | Systematic uncertainties (NOTE 6) |  |  |  | Value |
| 21 | Influence of noise |  |  |  | Table B.17.2-4 |
| Total measurement uncertainty |  |  |  |  | Value |
| EIRP total measurement uncertainty [dB] (23.45GHz <= f <= 40.8GHz) |  |  |  |  | 5.09 + Influence of Noise |
| EIRP total measurement uncertainty [dB] (40.8GHz < f <= 44.3GHz) |  |  |  |  | 6.61 + Influence of Noise |
| NOTE 1: Void NOTE 2: The analysis was done only for the case of operating at max output power – MPR – MBR(Multi-band relaxation)., in-band, non-CA. NOTE 3: The assessment assumes maximum DUT output power – MPR – MBR(Multi-band relaxation). NOTE 4: Void NOTE 5: Void NOTE 6: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 7: Void NOTE 8: Void NOTE 9: Void NOTE 10: Defined as fixed value MU contributor. |  |  |  |  |  |

<<Unchanged tables skipped>>

Table B.17.2-4: Influence of noise measurement (f=23.45GHz, 32.125GHz, 40.8GHz, 44.3GHz, Quiet Zone size ≤ 30 cm) for PC3 UEs

|  | FR2a | FR2b | FR2c |
| --- | --- | --- | --- |
| ChBW (50MHz) | 0.54 | 1.0 (NOTE 6) | TBD |
| ChBW (100MHz) | 1.0 | 1.0 (NOTE 5) | TBD |
| ChBW (200MHz) | 1.0 (NOTE 4) | 1.0 (NOTE 2) | TBD |
| ChBW (400MHz) | 1.0 (NOTE 1) | 1.0 (NOTE 3) | TBD |
| NOTE 1: This value is based on the relaxation of (MPR – 3.0) dB for MPR > 3.0dB. NOTE 2: Not applicable for MPR > 3.5dB NOTE 3: Not applicable for MPR > 2.0dB NOTE 4: This value is based on the relaxation of (MPR – 5.0) dB for MPR > 5.0dB. NOTE 5: Not applicable for MPR > 5.0dB NOTE 6: Not applicable for MPR > 7.5dB |  |  |  |

<<Unchanged sections skipped>>

B.21	Adjacent Channel Selectivity

Following tables summarize the MU threshold for Adjacent Channel Selectivity measurement. The origin MU values for different test setups with varies parameters can be found in following subclauses.

Table B.21-1: MU threshold for Adjacent Channel Selectivity

| Power Class | Frequency | MBW | Power | Threshold MU value (NOTE 1) |
| --- | --- | --- | --- | --- |
| PC3 | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = Max Output Power | 8.08 |
|  | 32.125GHz < f <= 40.8GHz |  |  | 8.08 |
|  | 40.8GHz < f <= 44.3GHz |  |  | TBD |
| PC1 | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = Max Output Power | 8.31 |
|  | 32.125GHz < f <= 40.8GHz |  |  | 8.31 |
| NOTE 1: Total Expanded MU for IFF for Quiet Zone size ≤ 30cm in Table B.21.2-2 for PC3 UEs and Table B.21.2-3 for PC1 UEs. |  |  |  |  |

B.21.1	Uncertainty budget format and assessment for DFF

FFS

B.21.2	Uncertainty budget format and assessment for IFF

The uncertainty contributions that may impact the overall MU value are listed in Table B.21.2-1.

Table B.21.2-1: Total Uncertainty contributions for Adjacent Channel Selectivity measurement

| UID | Description of uncertainty contribution | Details in clause |
| --- | --- | --- |
| Stage 2: DUT measurement (Wanted Signal contributions) |  |  |
| 1 | Positioning misalignment | B.2.2.1 |
| 2 | Measure distance uncertainty | B.2.2.2 |
| 3 | Quality of Quiet Zone | B.2.2.3 |
| 4 | Mismatch | B.2.2.4 |
| 5 | Standing wave between the DUT and measurement antenna | B.2.2.5 |
| 6 | gNB emulator uncertainty | B.2.2.17 |
| 7 | Phase curvature | B.2.2.7 |
| 8 | Amplifier uncertainties | B.2.2.8 |
| 9 | Random uncertainty | B.2.2.9 |
| 10 | Influence of the XPD | B.2.2.10 |
| 11 | Insertion Loss Variation | B.2.2.11 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | B.2.2.12 |
| 13 | Multiple measurement antenna uncertainty | B.2.2.25 |
| 14 | DUT repositioning | B.2.2.26 |
| Stage 2: DUT measurement (Modulated Interferer Signal specific contributions) |  |  |
| 15 | Positioning misalignment | B.2.2.1 |
| 16 | Measure distance uncertainty | B.2.2.2 |
| 17 | Quality of Quiet Zone | B.2.2.3 |
| 18 | Mismatch | B.2.2.4 |
| 19 | Standing wave between the DUT and measurement antenna | B.2.2.5 |
| 20 | Modulated Interferer uncertainty | B.2.2.33 |
| 21 | Phase curvature | B.2.2.7 |
| 22 | Amplifier uncertainties | B.2.2.8 |
| 23 | Random uncertainty | B.2.2.9 |
| 24 | Influence of the XPD | B.2.2.10 |
| 25 | Insertion Loss Variation | B.2.2.11 |
| 26 | RF leakage (from measurement antenna to the receiver/transmitter) | B.2.2.12 |
| 27 | Multiple measurement antenna uncertainty | B.2.2.25 |
| 28 | DUT repositioning | B.2.2.26 |
| 29 | Influence of offset antenna (Std.Dev) | B.2.2.35 |
| Stage 1: Calibration measurement (Wanted Signal contributions) |  |  |
| 30 | Mismatch | B.2.2.4 |
| 31 | Amplifier Uncertainties | B.2.2.8 |
| 32 | Misalignment of positioning System | B.2.2.13 |
| 33 | Uncertainty of the Network Analyzer | B.2.2.14 |
| 34 | Uncertainty of the absolute gain of the calibration antenna | B.2.2.15 |
| 35 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | B.2.2.16 |
| 36 | Phase centre offset of calibration antenna | B.2.2.18 |
| 37 | Quality of quiet zone for calibration process | B.2.2.19 |
| 38 | Standing wave between reference calibration antenna and measurement antenna | B.2.2.20 |
| 39 | Influence of the calibration antenna feed cable | B.2.2.21 |
| 40 | Insertion Loss Variation | B.2.2.11 |
| Stage 1: Calibration measurement (Modulated Interferer Signal contributions) |  |  |
| 41 | Mismatch | B.2.2.4 |
| 42 | Amplifier Uncertainties | B.2.2.8 |
| 43 | Misalignment of positioning System | B.2.2.13 |
| 44 | Uncertainty of the Network Analyzer | B.2.2.14 |
| 45 | Uncertainty of the absolute gain of the calibration antenna | B.2.2.15 |
| 46 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | B.2.2.16 |
| 47 | Phase centre offset of calibration antenna | B.2.2.18 |
| 48 | Quality of quiet zone for calibration process | B.2.2.19 |
| 48 | Standing wave between reference calibration antenna and measurement antenna | B.2.2.20 |
| 50 | Influence of the calibration antenna feed cable | B.2.2.21 |
| 51 | Insertion Loss Variation | B.2.2.11 |
| Systematic uncertainties |  |  |
| 52 | Systematic error related to beam peak search | B.2.2.28 |
| 53 | Additional impact of interferer ACLR | B.2.2.32 |
| 54 | Influence of offset antenna (mean error) | B.2.2.35 |

The uncertainty assessment tables are organized as follows:

-	For the purpose of uncertainty assessment, the radiating antenna aperture of the DUT is denoted as D

-	The uncertainty assessment has been derived for the case of Quiet Zone size ≤ [30 cm], f = {23.45GHz, 32.125GHz, 40.8GHz, 44.3GHz}, [P = maximum output power].

-	The uncertainty assessment for ACS is provided in Table B.21.2-2 for PC3 UEs and Table B.21.2-3 for PC1 UEs.

Table B.21.2-2: Uncertainty assessment for Adjacent Channel Selectivity measurement (f=23.45GHz, 32.125GHz, 40.8GHz, 44.3GHz, Quiet Zone size ≤ 30 cm) for PC3 UEs

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement (Wanted Signal contributions) |  |  |  |  |  |
| 1 | Positioning misalignment | 0.00 | Normal | 2.00 | 0.00 |
| 2 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 3 | Quality of Quiet Zone (NOTE 4) (23.45GHz <= f <= 40.8GHz) | 0.6 | Actual | 1.00 | 0.6 |
| 3 | Quality of Quiet Zone (NOTE 4) (40.8GHz < f <= 44.3GHz) | 0.7 | Actual | 1.00 | 0.7 |
| 4 | Mismatch (23.45GHz <= f <= 40.8GHz) | 1.30 | Actual | 1.00 | 1.30 |
| 4 | Mismatch (40.8GHz < f <= 44.3GHz) | 1.70 | Actual | 1.00 | 1.70 |
| 5 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | gNB uncertainty on absolute level (23.45GHz <= f <= 40.8GHz) | 2.9 | Normal | 2.00 | 1.45 |
| 6 | gNB uncertainty on absolute level (40.8GHz < f <= 44.3GHz) | 3.6 | Normal | 2.00 | 1.8 |
| 7 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 8 | Amplifier uncertainties | 2.1 | Normal | 2.00 | 1.05 |
| 9 | Random uncertainty | 0.50 | Normal | 2.00 | 0.25 |
| 10 | Influence of the XPD | 0.01 | U-shaped | 1.41 | 0.00 |
| 11 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 13 | Multiple measurement antenna uncertainty (NOTE 3) | 0.15 | Actual | 1.00 | 0.15 |
| 14 | DUT repositioning | 0.08 | Rectangular | 1.73 | 0.05 |
| Stage 2: DUT measurement (Modulated Interferer Signal specific contributions) |  |  |  |  |  |
| 15 | Positioning misalignment | 0.00 | Normal | 2.00 | 0.00 |
| 16 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 17 | Quality of Quiet Zone (NOTE 4) (23.45GHz <= f <= 40.8GHz) | 0.6 | Actual | 1.00 | 0.6 |
| 17 | Quality of Quiet Zone (NOTE 4) (40.8GHz < f <= 44.3GHz) | 0.7 | Actual | 1.00 | 0.7 |
| 18 | Mismatch (23.45GHz <= f <= 40.8GHz) | 1.30 | Actual | 1.00 | 1.30 |
| 18 | Mismatch (40.8GHz < f <= 44.3GHz) | 1.70 | Actual | 1.00 | 1.70 |
| 19 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 20 | Modulated Interferer uncertainty on absolute level | 2.9 | Normal | 2.00 | 1.45 |
| 21 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 22 | Amplifier uncertainties | 2.1 | Normal | 2.00 | 1.05 |
| 23 | Random uncertainty | 0.50 | Normal | 2.00 | 0.25 |
| 24 | Influence of the XPD | 0.01 | U-shaped | 1.41 | 0.00 |
| 25 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 26 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 27 | Multiple measurement antenna uncertainty (NOTE 3) | 0.15 | Actual | 1.00 | 0.15 |
| 28 | DUT repositioning | 0.08 | Rectangular | 1.73 | 0.05 |
| 29 | Influence of offset antenna (Std.Dev) (NOTE 5) | 0.00 | Normal | 2.00 | 0.00 |
| Stage 1: Calibration measurement (Wanted Signal contributions) |  |  |  |  |  |
| 30 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 31 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 32 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 33 | Uncertainty of the Network Analyzer (23.45GHz <= f <= 40.8GHz) | 1.50 | Normal | 2.00 | 0.75 |
| 33 | Uncertainty of the Network Analyzer (40.8GHz < f <= 44.3GHz) | 1.70 | Normal | 2.00 | 0.85 |
| 34 | Uncertainty of the absolute gain of the calibration antenna | 0.60 | Normal | 2.00 | 0.30 |
| 35 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.01 | Rectangular | 1.73 | 0.00 |
| 36 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 37 | Quality of quiet zone for calibration process (NOTE 4) (23.45GHz <= f <= 40.8GHz) | 0.4 | Actual | 1.00 | 0.4 |
| 37 | Quality of quiet zone for calibration process (NOTE 4) (40.8GHz < f <= 44.3GHz) | 0.5 | Actual | 1.00 | 0.5 |
| 38 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 39 | Influence of the calibration antenna feed cable | 0.14 | Normal | 2.00 | 0.07 |
| 40 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| Stage 1: Calibration measurement (Modulated Interferer Signal contributions) |  |  |  |  |  |
| 41 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 42 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 43 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 44 | Uncertainty of the Network Analyzer (23.45GHz <= f <= 40.8GHz) | 1.50 | Normal | 2.00 | 0.75 |
| 44 | Uncertainty of the Network Analyzer (40.8GHz < f <= 44.3GHz) | 1.70 | Normal | 2.00 | 0.85 |
| 45 | Uncertainty of the absolute gain of the calibration antenna | 0.60 | Normal | 2.00 | 0.30 |
| 46 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.01 | Rectangular | 1.73 | 0.00 |
| 47 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 48 | Quality of quiet zone for calibration process (NOTE 4) (23.45GHz <= f <= 40.8GHz) | 0.4 | Actual | 1.00 | 0.4 |
| 48 | Quality of quiet zone for calibration process (NOTE 4) (40.8GHz < f <= 44.3GHz) | 0.5 | Actual | 1.00 | 0.5 |
| 48 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 50 | Influence of the calibration antenna feed cable | 0.14 | Normal | 2.00 | 0.07 |
| 51 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
|  | Systematic uncertainties (NOTE 2) |  |  |  | Value |
| 52 | Systematic error related to beam peak search |  |  |  | 0.5 |
| 53 | Additional impact of interferer ACLR |  |  |  | 0.7 |
| 54 | Influence of offset antenna (mean error) (NOTE 5) |  |  |  | 0.00 |
| Total measurement uncertainty |  |  |  |  | Value |
| ACS Expanded uncertainty (23.45GHz <= f <= 40.8GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 8.08 |
| ACS Expanded uncertainty (40.8GHz < f <= 44.3GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | TBD |
| NOTE 1: The analysis was done only for the case of operating at max output power, in-band, non-CA. NOTE 2: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 3: Applies to the system which has a structure of mechanical feed antenna positioning. NOTE 4: Value based on procedure defined in clause D.2 of TR 38.810 for Quiet Zone size less or equal to 30 cm. NOTE 5: For MTSU derivation purpose, this value is set to 0.0 (no offset antenna case). |  |  |  |  |  |

<<End of change>>