3GPP TSG-RAN5 Meeting #106	R5-251045

# Athens, Greece, 17th Feb 2025 - 21st Feb 2025

| CR-Form-v12.3 |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CHANGE REQUEST |  |  |  |  |  |  |  |  |
|  | 38.561 | CR | 0020 | rev | - | Current version: | 18.1.0 |  |
| For HELP on using this form: comprehensive instructions can be found at http://www.3gpp.org/Change-Requests. |  |  |  |  |  |  |  |  |

| Proposed change affects: | UICC apps |  | ME |  | Radio Access Network |  | Core Network |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |

| Title: | MU contents for BHH and WR |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Source to WG: | ROHDE & SCHWARZ |  |  |  |  |  |  |  |  |  |
| Source to TSG: | R5 |  |  |  |  |  |  |  |  |  |
| Work item code: | NR_FR1_TRP_TRS_enh-UEConTest |  |  |  |  |  | Date: |  |  | 2025-02-07 |
| Category: | F |  |  |  |  |  | Release: |  |  | Rel-18 |
|  | Use one of the following categories: F (correction) A (mirror corresponding to a change in an earlier release) B (addition of feature), C (functional modification of feature) D (editorial modification) Detailed explanations of the above categories can be found in 3GPP TR 21.900. |  |  |  |  |  |  |  | Use one of the following releases: Rel-8 (Release 8) Rel-9 (Release 9) Rel-10 (Release 10) Rel-11 (Release 11) … Rel-17 (Release 17) Rel-18 (Release 18) Rel-19 (Release 19) Rel-20 (Release 20) |  |
| Reason for change: |  | Beside Head and Hand phantom (Talk mode) and Forearm phantoms (Wrist-worn mode) were introduced in the Release 18 version of core specifications, and the corresponding MU contents need to be incorporated in the test specifications. |  |  |  |  |  |  |  |  |
| Summary of change: |  | Update the MU section with the corresponding contents for Talk mode and Wrist-worn mode |  |  |  |  |  |  |  |  |
| Consequences if not approved: |  | Missing MU references in test specification for Release 18 |  |  |  |  |  |  |  |  |
| Clauses affected: |  | A.4.2.5, A.4.3.1, A.4.4.1, D.2 |  |  |  |  |  |  |  |  |
|  |  | Y | N |  |  |  |  |  |  |  |
| Other specs |  |  | X | Other core specifications |  |  |  | TS/TR ... CR ... |  |  |
| affected: |  |  | X | Test specifications |  |  |  | TS/TR ... CR ... |  |  |
| (show related CRs) |  |  | X | O&M Specifications |  |  |  | TS/TR ... CR ... |  |  |
| Other comments: |  | MU contents aligned with TR 38.870 V18.4.0 |  |  |  |  |  |  |  |  |
| This CR's revision history: |  | - |  |  |  |  |  |  |  |  |

< Unchanged Text Deleted >

< Beginning of Changes >

### A.4.2.5	Communication tester: uncertainty of the absolute level

The transmitter device (typically a communication tester or BS simulator) is used to drive a signal to the measurement antenna in sensitivity tests either as an absolute level or as a relative level. Receiving device used is the UE. Generally, there occurs uncertainty contribution from limited absolute level accuracy and non-linearity of the communication tester.

For practical reasons, the calibration measurement (Stage 1) should be only performed with the measurement antenna as a receiver. Hence, the uncertainty on the absolute level of the transmitter device cannot be assumed as systematic. This uncertainty will be determined from the manufacturer’s datasheet and the distribution used (see clause 5.1.2 in [5]) shall match that provided in the datasheet which are commonly quoting MUs/accuracies with a “95% confidence level” and/or a “coverage factor of 2”. In the absence of a declared distribution in the datasheet, the rectangular distribution should be used. Furthermore, the uncertainty of the non-linearity of the device is included in the absolute level uncertainty. Given the wide bandwidths of the modulated signal received during the testing, the level flatness and not just the CW level accuracy shall be taken into account.

### A.4.2.6	Sensitivity measurement: output level step resolution

When output power of the communication tester is swept to reach the throughput target that defines the sensitivity threshold, the final power step resolution represents an asymmetric uncertainty contribution that can be corrected since this uncertainty is device and test system independent. The lab shall correct the TRS results by -0.5·output level step resolution of the final power step search and note the correction in the test report.

Alternatively, if the alternate EIS search procedure is applied as described in A.3.3.3.1 to assess the EIS at each grid point, no correction of TRS shall be applied and the no uncertainty contribution shall be considered.

Please note that the digital error rater uncertainty is included in the random uncertainty, see A.4.2.13.

< Unchanged Text Deleted >

#### A.4.2.11.6	Uncertainty of dielectric properties and shape of the forearm phantom

The assessment of uncertainties for the phantoms defined in CTIA Certification OTA Test Plan 01.72 Section 2 [12], is further detailed in CTIA Certification OTA Test Plan 01.70 Section 2 [8]. Values are from the CTIA Certification Test Plan for Wireless Device Over-the-Air Performance © CTIA Certification. Reproduced with permission.

< Unchanged Text Deleted >

## A.4.3	Total Radiated Power (TRP)

### A.4.3.1	Anechoic Chamber Method

The uncertainty contributions related to TRP for the Anechoic Chamber method are listed in Table A.4.3.1-1. Examples of uncertainty budget are presented in Tables A.4.3.1-2 through Table A.4.3.1-4.

Table A.4.3.1-1: Uncertainty contributions in TRP measurement for anechoic chamber method

| UID | Description of uncertainty contribution | Details in clause |
| --- | --- | --- |
| Stage 2: DUT measurement (Figure A.3.1-1, Figure A.3.1-2) |  |  |
| 1 | Mismatch of receiver chain | A.4.2.1 |
| 2 | Insertion loss of receiver chain | A.4.2.2 |
| 3 | Influence of the measurement antenna cable | A.4.2.3 |
| 4 | Measurement Receiver: uncertainty of the absolute level | A.4.2.4 |
| 5 | Measurement distance | A.4.2.7.1 |
| 6 | Quality of quiet zone | A.4.2.8 |
| 7 | DUT Tx-power drift | A.4.2.9 |
| 8 | Uncertainty related to the use of phantoms | A.4.2.11 |
| 9 | Coarse sampling grid | A.4.2.12 |
| 10 | Random uncertainty | A.4.2.13 |
| 11 | Frequency Response | A.4.2.14 |
| Stage 1: Calibration measurement, network analyzer method (Figure A.3.2-1) |  |  |
| 12 | Uncertainty of network analyzer | A.4.2.15 |
| 13 | Mismatch of receiver chain | A.4.2.1 |
| 14 | Insertion loss of receiver chain | A.4.2.2 |
| 15 | Mismatch in the connection of calibration antenna | A.4.2.1 |
| 16 | Influence of the calibration antenna feed cable | A.4.2.3 |
| 17 | Influence of the measurement antenna cable | A.4.2.3 |
| 18 | Uncertainty of the absolute gain/ radiation efficiency of the calibration antenna | A.4.2.16 |
| 19 | Measurement distance | A.4.2.7.2 |
| 20 | Quality of the Quiet Zone | A.4.2.8 |
| Systematic Errors |  |  |
| 21 | Systematic Error related to TRP grids | A.4.2.12 |

Table A.4.3.1-2: Example of uncertainty budget for TRP hand only (browsing mode) measurement for anechoic chamber method for NR FR1 bands

| UID | Uncertainty Source | Comment | Uncertainty Value [dB] |  | Prob Distr | Div | ci | Standard Uncertainty [dB] |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  | Below 3GHz | Above 3GHz |  |  |  | Below 3GHz | Above 3GHz |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |  |
| 1 | Mismatch of receiver chain | Гreceiver < 0.33 Гmeasurement antenna < 0.5 Cable attenuation > 3dB | 0.26 | 0.26 | U-shaped | 1.41 | 1 | 0.18 | 0.18 |
| 2 | Insertion loss of receiver chain | Systematic with Stage 1 (=> cancels) | 0 | 0 | Rectangular | 1.73 | 1 | 0.00 | 0.00 |
| 3 | Influence of the measurement antenna cable | Systematic with Stage 1 (=> cancels) | 0 | 0 | Rectangular | 1.73 | 1 | 0.00 | 0.00 |
| 4 | Measurement Receiver: uncertainty of the absolute level | From datasheet of communication tester or spectrum analyser. | 0.42 | 0.54 | Normal | 2.00 | 1 | 0.21 | 0.27 |
| 5 | Measurement distance | DUT is not offset for hand-only phantom testing | 0 | 0 | Rectangular | 1.73 | 1 | 0.00 | 0.00 |
| 6 | Quality of quiet zone | Surface standard deviation of power measurements in ripple test | 0.5 | 0.5 | Actual | 1.00 | 1 | 0.50 | 0.50 |
| 7 | DUT Tx-power drift | Drift | 0.2 | 0.2 | Rectangular | 1.73 | 1 | 0.12 | 0.12 |
| 8 | Uncertainty related to the use of phantoms | Material Dielectric Constant, Material Conductivity, Geometry/Shape (incl. spacer), Data Mode Fixture | 0.64 | 0.64 | Rectangular | 1.73 | 1 | 0.37 | 0.37 |
| 9 | Coarse sampling grid | Sampling grids per Table A.4.2.12-1 | 0 | 0 | Actual | 1.00 | 1 | 0.0 | 0.0 |
| 10 | Random Uncertainty | Fixed MU to account for all the unknown, unquantifiable, etc. uncertainties | 0.25 | 0.25 | Normal | 2.00 | 1 | 0.13 | 0.13 |
| 11 | Frequency Response | Average path loss corrected | 0 | 0 | Rectangular | 1.73 | 1 | 0.00 | 0.00 |
| Stage 1: Calibration measurement, network analyzer method |  |  |  |  |  |  |  |  |  |
| 12 | Uncertainty of network analyzer | From datasheet of VNA with assessed transmission coefficients | 0.2 | 0.5 | Normal | 2.00 | 1 | 0.10 | 0.25 |
| 13 | Mismatch of receiver chain | Taken into account in VNA uncertainty term | 0 | 0 | U-shaped | 1.41 | 1 | 0.00 | 0.00 |
| 14 | Insertion loss of receiver chain | Systematic with Stage 2 (=> cancels) | 0 | 0 | Rectangular | 1.73 | 1 | 0.00 | 0.00 |
| 15 | Mismatch in the connection of calibration antenna | Taken in to account in VNA setup uncertainty | 0 | 0 | U-shaped | 1.41 | 1 | 0.00 | 0.00 |
| 16 | Influence of the calibration antenna feed cable | Gain calibration with a dipole | 0.3 | 0.3 | Rectangular | 1.73 | 1 | 0.17 | 0.17 |
| 17 | Influence of the measurement antenna cable | Systematic with Stage 2 (=> cancels) | 0 | 0 | Rectangular | 1.73 | 1 | 0.00 | 0.00 |
| 18 | Uncertainty of the absolute gain/ radiation efficiency of the calibration antenna | Calibration report with traceability to a National Metrology Institute | 0.58 | 0.58 | Normal | 2.00 | 1 | 0.29 | 0.29 |
| 19 | Measurement distance | Dipole: phase centre is aligned with axis of rotation | 0 | 0 | Rectangular | 1.73 | 1 | 0.00 | 0.00 |
| 20 | Quality of the Quiet Zone | Peak-to-null ripple | 0.5 | 0.5 | Rectangular | 1.73 | 1 | 0.29 | 0.29 |
| Combined Standard Uncertainty [dB] |  |  |  |  |  |  |  | 0.84 | 0.88 |
| Expanded Uncertainty [dB] (Confidence interval of 95 %) |  |  |  |  |  |  |  | 1.64 | 1.73 |
| 21 | Systematic Error related to TRP grids | mean error | 0 | 0 | Actual | 1.00 | 1 | 0.00 | 0.00 |
| Total Expanded Uncertainty [dB] (Confidence interval of 95 %) |  |  |  |  |  |  |  | 1.64 | 1.73 |

Table A.4.3.1-3: Example of uncertainty budget for TRP Beside Head and Hand (Talk mode) measurement for anechoic chamber method for NR FR1 bands

| UID | Uncertainty Source | Comment | Uncertainty Value [dB] |  | Prob Distr | Div | ci | Standard Uncertainty [dB] |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  | Below 3GHz | Above 3GHz |  |  |  | Below 3GHz | Above 3GHz |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |  |
| 1 | Mismatch of receiver chain | Гreceiver < 0.33 Гmeasurement antenna < 0.5 Cable attenuation > 3dB | 0.26 | 0.26 | U-shaped | 1.41 | 1 | 0.18 | 0.18 |
| 2 | Insertion loss of receiver chain | Systematic with Stage 1 (=> cancels) | 0 | 0 | Rectangular | 1.73 | 1 | 0.00 | 0.00 |
| 3 | Influence of the measurement antenna cable | Systematic with Stage 1 (=> cancels) | 0 | 0 | Rectangular | 1.73 | 1 | 0.00 | 0.00 |
| 4 | Measurement Receiver: uncertainty of the absolute level | From datasheet of communication tester or spectrum analyser. | 0.42 | 0.54 | Normal | 2.00 | 1 | 0.21 | 0.27 |
| 5 | Measurement distance | d=1.6m, Δd=0.05m | 0.27 | 0.27 | Rectangular | 1.73 | 1 | 0.16 | 0.16 |
| 6 | Quality of quiet zone | Surface standard deviation of power measurements in ripple test | 0.5 | 0.5 | Actual | 1.00 | 1 | 0.50 | 0.50 |
| 7 | DUT Tx-power drift | Drift | 0.2 | 0.2 | Rectangular | 1.73 | 1 | 0.12 | 0.12 |
| 8 | Uncertainty related to the use of phantoms | Material Dielectric Constant, Material Conductivity, Geometry/Shape (incl. spacer), Beside Head and Hand | 0.99 | 0.99 | Rectangular | 1.73 | 1 | 0.57 | 0.57 |
| 9 | Coarse sampling grid | Sampling grids per Table A.4.2.12-1 | 0 | 0 | Actual | 1.00 | 1 | 0.0 | 0.0 |
| 10 | Random Uncertainty | Fixed MU to account for all the unknown, unquantifiable, etc. uncertainties | 0.25 | 0.25 | Normal | 2.00 | 1 | 0.13 | 0.13 |
| 11 | Frequency Response | Average path loss corrected | 0 | 0 | Rectangular | 1.73 | 1 | 0.00 | 0.00 |
| Stage 1: Calibration measurement, network analyzer method |  |  |  |  |  |  |  |  |  |
| 12 | Uncertainty of network analyzer | From datasheet of VNA with assessed transmission coefficients | 0.2 | 0.5 | Normal | 2.00 | 1 | 0.10 | 0.25 |
| 13 | Mismatch of receiver chain | Taken into account in VNA uncertainty term | 0 | 0 | U-shaped | 1.41 | 1 | 0.00 | 0.00 |
| 14 | Insertion loss of receiver chain | Systematic with Stage 2 (=> cancels) | 0 | 0 | Rectangular | 1.73 | 1 | 0.00 | 0.00 |
| 15 | Mismatch in the connection of calibration antenna | Taken in to account in VNA setup uncertainty | 0 | 0 | U-shaped | 1.41 | 1 | 0.00 | 0.00 |
| 16 | Influence of the calibration antenna feed cable | Gain calibration with a dipole | 0.3 | 0.3 | Rectangular | 1.73 | 1 | 0.17 | 0.17 |
| 17 | Influence of the measurement antenna cable | Systematic with Stage 2 (=> cancels) | 0 | 0 | Rectangular | 1.73 | 1 | 0.00 | 0.00 |
| 18 | Uncertainty of the absolute gain/ radiation efficiency of the calibration antenna | Calibration report with traceability to a National Metrology Institute | 0.58 | 0.58 | Normal | 2.00 | 1 | 0.29 | 0.29 |
| 19 | Measurement distance | Dipole: phase centre is aligned with axis of rotation | 0 | 0 | Rectangular | 1.73 | 1 | 0.00 | 0.00 |
| 20 | Quality of the Quiet Zone | Peak-to-null ripple | 0.5 | 0.5 | Rectangular | 1.73 | 1 | 0.29 | 0.29 |
| Combined Standard Uncertainty [dB] |  |  |  |  |  |  |  | 0.96 | 1.00 |
| Expanded Uncertainty [dB] (Confidence interval of 95 %) |  |  |  |  |  |  |  | 1.88 | 1.96 |
| 21 | Systematic Error related to TRP grids | mean error | 0 | 0 | Actual | 1.00 | 1 | 0.00 | 0.00 |
| Total Expanded Uncertainty [dB] (Confidence interval of 95 %) |  |  |  |  |  |  |  | 1.88 | 1.96 |

Table A.4.3.1-4: Example of uncertainty budget for TRP Wrist-Worn device measurement for anechoic chamber method for NR FR1 bands

| UID | Uncertainty Source | Comment | Uncertainty Value [dB] |  | Prob Distr | Div | ci | Standard Uncertainty [dB] |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  | Below 3GHz | Above 3GHz |  |  |  | Below 3GHz | Above 3GHz |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |  |
| 1 | Mismatch of receiver chain | Гreceiver < 0.33 Гmeasurement antenna < 0.5 Cable attenuation > 3dB | 0.26 | 0.26 | U-shaped | 1.41 | 1 | 0.18 | 0.18 |
| 2 | Insertion loss of receiver chain | Systematic with Stage 1 (=> cancels) | 0 | 0 | Rectangular | 1.73 | 1 | 0.00 | 0.00 |
| 3 | Influence of the measurement antenna cable | Systematic with Stage 1 (=> cancels) | 0 | 0 | Rectangular | 1.73 | 1 | 0.00 | 0.00 |
| 4 | Measurement Receiver: uncertainty of the absolute level | From datasheet of communication tester or spectrum analyser. | 0.42 | 0.54 | Normal | 2.00 | 1 | 0.21 | 0.27 |
| 5 | Measurement distance | d=1.6m, Δd=0.05m | 0.27 | 0.27 | Rectangular | 1.73 | 1 | 0.16 | 0.16 |
| 6 | Quality of quiet zone | Surface standard deviation of power measurements in ripple test | 0.5 | 0.5 | Actual | 1.00 | 1 | 0.50 | 0.50 |
| 7 | DUT Tx-power drift | Drift | 0.2 | 0.2 | Rectangular | 1.73 | 1 | 0.12 | 0.12 |
| 8 | Uncertainty related to the use of phantoms | Material Dielectric Constant, Material Conductivity, Geometry/Shape (incl. spacer), Forearm | 0.74 | 0.74 | Rectangular | 1.73 | 1 | 0.43 | 0.43 |
| 9 | Coarse sampling grid | Sampling grids per Table A.4.2.12-1 | 0 | 0 | Actual | 1.00 | 1 | 0.0 | 0.0 |
| 10 | Random Uncertainty | Fixed MU to account for all the unknown, unquantifiable, etc. uncertainties | 0.25 | 0.25 | Normal | 2.00 | 1 | 0.13 | 0.13 |
| 11 | Frequency Response | Average path loss corrected | 0 | 0 | Rectangular | 1.73 | 1 | 0.00 | 0.00 |
| Stage 1: Calibration measurement, network analyzer method |  |  |  |  |  |  |  |  |  |
| 12 | Uncertainty of network analyzer | From datasheet of VNA with assessed transmission coefficients | 0.2 | 0.5 | Normal | 2.00 | 1 | 0.10 | 0.25 |
| 13 | Mismatch of receiver chain | Taken into account in VNA uncertainty term | 0 | 0 | U-shaped | 1.41 | 1 | 0.00 | 0.00 |
| 14 | Insertion loss of receiver chain | Systematic with Stage 2 (=> cancels) | 0 | 0 | Rectangular | 1.73 | 1 | 0.00 | 0.00 |
| 15 | Mismatch in the connection of calibration antenna | Taken in to account in VNA setup uncertainty | 0 | 0 | U-shaped | 1.41 | 1 | 0.00 | 0.00 |
| 16 | Influence of the calibration antenna feed cable | Gain calibration with a dipole | 0.3 | 0.3 | Rectangular | 1.73 | 1 | 0.17 | 0.17 |
| 17 | Influence of the measurement antenna cable | Systematic with Stage 2 (=> cancels) | 0 | 0 | Rectangular | 1.73 | 1 | 0.00 | 0.00 |
| 18 | Uncertainty of the absolute gain/ radiation efficiency of the calibration antenna | Calibration report with traceability to a National Metrology Institute | 0.58 | 0.58 | Normal | 2.00 | 1 | 0.29 | 0.29 |
| 19 | Measurement distance | Dipole: phase centre is aligned with axis of rotation | 0 | 0 | Rectangular | 1.73 | 1 | 0.00 | 0.00 |
| 20 | Quality of the Quiet Zone | Peak-to-null ripple | 0.5 | 0.5 | Rectangular | 1.73 | 1 | 0.29 | 0.29 |
| Combined Standard Uncertainty [dB] |  |  |  |  |  |  |  | 0.86 | 0.91 |
| Expanded Uncertainty [dB] (Confidence interval of 95 %) |  |  |  |  |  |  |  | 1.69 | 1.78 |
| 21 | Systematic Error related to TRP grids | mean error | 0 | 0 | Actual | 1.00 | 1 | 0.00 | 0.00 |
| Total Expanded Uncertainty [dB] (Confidence interval of 95 %) |  |  |  |  |  |  |  | 1.69 | 1.78 |

## A.4.4	Total Radiated Sensitivity (TRS)

### A.4.4.1	Anechoic Chamber Method

The uncertainty contributions related to TRS for the Anechoic Chamber method are listed in Table A.4.4.1-1. Examples of uncertainty budget are presented in Tables A.4.4.1-2 through A.4.4.1-4.

Table A.4.4.1-1: Uncertainty contributions in TRS measurement for anechoic chamber method

| UID | Description of uncertainty contribution | Details in clause |
| --- | --- | --- |
| Stage 2: DUT measurement (Figure A.3.1-1, Figure A.3.1-2) |  |  |
| 1 | Mismatch of transmitter chain | A.4.2.1 |
| 2 | Insertion loss of transmitter chain | A.4.2.2 |
| 3 | Influence of the measurement antenna cable | A.4.2.3 |
| 4 | Communication Tester: uncertainty of the absolute output level | A.4.2.5 |
| 5 | Sensitivity measurement: output level step resolution | A.4.2.6 |
| 6 | Measurement distance | A.4.2.7.1 |
| 7 | Quality of quiet zone | A.4.2.8 |
| 8 | DUT sensitivity drift | A.4.2.10 |
| 9 | Uncertainty related to the use of phantoms | A.4.2.11 |
| 10 | Coarse sampling grid | A.4.2.12 |
| 11 | Random uncertainty | A.4.2.13 |
| 12 | Frequency Response | A.4.2.14 |
| Stage 1: Calibration measurement, network analyzer method (Figure A.3.2-1) |  |  |
| 13 | Uncertainty of network analyzer | A.4.2.15 |
| 14 | Mismatch of transmitter chain | A.4.2.1 |
| 15 | Insertion loss of transmitter chain | A.4.2.2 |
| 16 | Mismatch in the connection of calibration antenna | A.4.2.1 |
| 17 | Influence of the calibration antenna feed cable | A.4.2.3 |
| 18 | Influence of the measurement antenna cable | A.4.2.3 |
| 19 | Uncertainty of the absolute gain/radiation efficiency of the calibration antenna | A.4.2.16 |
| 20 | Measurement distance | A.4.2.7.2 |
| 21 | Quality of quiet zone | A.4.2.8 |
| Systematic Errors |  |  |
| 22 | Systematic Error related to TRS grids | A.4.2.12 |

Table A.4.4.1-2: Example of uncertainty budget for TRS hand only (browsing mode) measurement for anechoic chamber method for NR FR1 bands

| UID | Uncertainty Source | Comment | Uncertainty Value [dB] |  | Prob Distr | Div | ci | Standard Uncertainty [dB] |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  | Below 3GHz | Above 3GHz |  |  |  | Below 3GHz | Above 3GHz |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |  |
| 1 | Mismatch of transmitter chain | Гreceiver < 0.29 Гmeasurement antenna < 0.5 Cable attenuation > 3dB | 0.22 | 0.22 | U-shaped | 1.41 | 1 | 0.16 | 0.16 |
| 2 | Insertion loss of transmitter chain | Systematic with Stage 1 (=> cancels) | 0 | 0 | Rectangular | 1.73 | 1 | 0.00 | 0.00 |
| 3 | Influence of the measurement antenna cable | Systematic with Stage 1 (=> cancels) | 0 | 0 | Rectangular | 1.73 | 1 | 0.00 | 0.00 |
| 4 | Communication Tester: uncertainty of the absolute output level | From datasheet of communication tester | 1.30 | 1.30 | Normal | 2.00 | 1 | 0.65 | 0.65 |
| 5 | Sensitivity measurement: output level step resolution | Systematic error that can be corrected | 0 | 0 | Rectangular | 1.73 | 1 | 0.00 | 0.00 |
| 6 | Measurement distance | DUT is not offset for hand-only phantom testing | 0 | 0 | Rectangular | 1.73 | 1 | 0 | 0 |
| 7 | Quality of quiet zone | Surface standard deviation of power measurements in ripple test | 0.5 | 0.5 | Actual | 1.00 | 1 | 0.50 | 0.50 |
| 8 | DUT sensitivity drift | Drift measurement | 0.2 | 0.2 | Rectangular | 1.73 | 1 | 0.12 | 0.12 |
| 9 | Uncertainty related to the use of phantoms | Material Dielectric Constant, Material Conductivity, Geometry/Shape (incl. spacer), Data Mode Fixture | 0.64 | 0.64 | Rectangular | 1.73 | 1 | 0.37 | 0.37 |
| 10 | Coarse sampling grid | Sampling grids per Table A.4.2.12-1 | 0.10 | 0.18 | Actual | 1.00 | 1 | 0.10 | 0.18 |
| 11 | Random uncertainty | Fixed MU to account for all the unknown, unquantifiable, etc. uncertainties including digital error rate | 0.4 | 0.4 | Normal | 2.00 | 1 | 0.20 | 0.20 |
| 12 | Frequency Response | Included in the output level step resolution | 0 | 0 | Rectangular | 1.73 | 1 | 0.00 | 0.00 |
| Stage 1: Calibration measurement, network analyzer method |  |  |  |  |  |  |  |  |  |
| 13 | Uncertainty of network analyzer | From datasheet of VNA with assessed transmission coefficients | 0.2 | 0.5 | Normal | 2.00 | 1 | 0.10 | 0.25 |
| 14 | Mismatch of transmitter chain | Taken into account in VNA setup uncertainty | 0 | 0 | U-shaped | 1.41 | 1 | 0.00 | 0.00 |
| 15 | Insertion loss of transmitter chain | Systematic with Stage 2 (=> cancels) | 0 | 0 | Rectangular | 1.73 | 1 | 0.00 | 0.00 |
| 16 | Mismatch in the connection of calibration antenna | Taken into account in VNA setup uncertainty | 0 | 0 | U-shaped | 1.41 | 1 | 0.00 | 0.00 |
| 17 | Influence of the calibration antenna feed cable | Gain calibration with dipole | 0.3 | 0.3 | Rectangular | 1.73 | 1 | 0.17 | 0.17 |
| 18 | Influence of the measurement antenna cable | Systematic with Stage 2 (=> cancels) | 0 | 0 | Rectangular | 1.73 | 1 | 0.00 | 0.00 |
| 19 | Uncertainty of the absolute gain/ radiation efficiency of the calibration antenna | Calibration report with traceability to a National Metrology Institute | 0.58 | 0.58 | Normal | 2.00 | 1 | 0.29 | 0.29 |
| 20 | Measurement distance | Dipole: phase centre is aligned with axis of rotation | 0 | 0 | Rectangular | 1.73 | 1 | 0.00 | 0.00 |
| 21 | Quality of quiet zone | Peak-to-null ripple | 0.5 | 0.5 | Rectangular | 1.73 | 1 | 0.29 | 0.29 |
| Combined Standard Uncertainty [dB] |  |  |  |  |  |  |  | 1.05 | 1.09 |
| Expanded Uncertainty [dB] (Confidence interval of 95 %) |  |  |  |  |  |  |  | 2.06 | 2.13 |
| 22 | Systematic Error related to TRS grids | mean error | 0 | 0.00 | Actual | 1.00 | 1 | 0.00 | 0.00 |
| Total Expanded Uncertainty [dB] (Confidence interval of 95 %) |  |  |  |  |  |  |  | 2.06 | 2.13 |

Table A.4.4.1-3: Example of uncertainty budget for TRS Beside Head and Hand (Talk mode) measurement for anechoic chamber method for NR FR1 bands

| UID | Uncertainty Source | Comment | Uncertainty Value [dB] |  | Prob Distr | Div | ci | Standard Uncertainty [dB] |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  | Below 3GHz | Above 3GHz |  |  |  | Below 3GHz | Above 3GHz |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |  |
| 1 | Mismatch of transmitter chain | Гreceiver < 0.29 Гmeasurement antenna < 0.5 Cable attenuation > 3dB | 0.22 | 0.22 | U-shaped | 1.41 | 1 | 0.16 | 0.16 |
| 2 | Insertion loss of transmitter chain | Systematic with Stage 1 (=> cancels) | 0 | 0 | Rectangular | 1.73 | 1 | 0.00 | 0.00 |
| 3 | Influence of the measurement antenna cable | Systematic with Stage 1 (=> cancels) | 0 | 0 | Rectangular | 1.73 | 1 | 0.00 | 0.00 |
| 4 | Communication Tester: uncertainty of the absolute output level | From datasheet of communication tester | 1.30 | 1.30 | Normal | 2.00 | 1 | 0.65 | 0.65 |
| 5 | Sensitivity measurement: output level step resolution | Systematic error that can be corrected | 0 | 0 | Rectangular | 1.73 | 1 | 0.00 | 0.00 |
| 6 | Measurement distance | d=1.6m, Δd=0.05m | 0.27 | 0.27 | Rectangular | 1.73 | 1 | 0.12 | 0.12 |
| 7 | Quality of quiet zone | Surface standard deviation of power measurements in ripple test | 0.5 | 0.5 | Actual | 1.00 | 1 | 0.50 | 0.50 |
| 8 | DUT sensitivity drift | Drift measurement | 0.2 | 0.2 | Rectangular | 1.73 | 1 | 0.12 | 0.12 |
| 9 | Uncertainty related to the use of phantoms | Material Dielectric Constant, Material Conductivity, Geometry/Shape (incl. spacer), Beside Head and Hand | 0.99 | 0.99 | Rectangular | 1.73 | 1 | 0.57 | 0.57 |
| 10 | Coarse sampling grid | Sampling grids per Table A.4.2.12-1 | 0.10 | 0.18 | Actual | 1.00 | 1 | 0.10 | 0.18 |
| 11 | Random uncertainty | Fixed MU to account for all the unknown, unquantifiable, etc. uncertainties including digital error rate | 0.4 | 0.4 | Normal | 2.00 | 1 | 0.20 | 0.20 |
| 12 | Frequency Response | Included in the output level step resolution | 0 | 0 | Rectangular | 1.73 | 1 | 0.00 | 0.00 |
| Stage 1: Calibration measurement, network analyzer method |  |  |  |  |  |  |  |  |  |
| 13 | Uncertainty of network analyzer | From datasheet of VNA with assessed transmission coefficients | 0.2 | 0.5 | Normal | 2.00 | 1 | 0.10 | 0.25 |
| 14 | Mismatch of transmitter chain | Taken into account in VNA setup uncertainty | 0 | 0 | U-shaped | 1.41 | 1 | 0.00 | 0.00 |
| 15 | Insertion loss of transmitter chain | Systematic with Stage 2 (=> cancels) | 0 | 0 | Rectangular | 1.73 | 1 | 0.00 | 0.00 |
| 16 | Mismatch in the connection of calibration antenna | Taken into account in VNA setup uncertainty | 0 | 0 | U-shaped | 1.41 | 1 | 0.00 | 0.00 |
| 17 | Influence of the calibration antenna feed cable | Gain calibration with dipole | 0.3 | 0.3 | Rectangular | 1.73 | 1 | 0.17 | 0.17 |
| 18 | Influence of the measurement antenna cable | Systematic with Stage 2 (=> cancels) | 0 | 0 | Rectangular | 1.73 | 1 | 0.00 | 0.00 |
| 19 | Uncertainty of the absolute gain/ radiation efficiency of the calibration antenna | Calibration report with traceability to a National Metrology Institute | 0.58 | 0.58 | Normal | 2.00 | 1 | 0.29 | 0.29 |
| 20 | Measurement distance | Dipole: phase centre is aligned with axis of rotation | 0 | 0 | Rectangular | 1.73 | 1 | 0.00 | 0.00 |
| 21 | Quality of quiet zone | Peak-to-null ripple | 0.5 | 0.5 | Rectangular | 1.73 | 1 | 0.29 | 0.29 |
| Combined Standard Uncertainty [dB] |  |  |  |  |  |  |  | 1.15 | 1.18 |
| Expanded Uncertainty [dB] (Confidence interval of 95 %) |  |  |  |  |  |  |  | 2.25 | 2.31 |
| 22 | Systematic Error related to TRS grids | mean error | 0 | 0.00 | Actual | 1.00 | 1 | 0.00 | 0.00 |
| Total Expanded Uncertainty [dB] (Confidence interval of 95 %) |  |  |  |  |  |  |  | 2.25 | 2.31 |

Table A.4.4.1-4: Example of uncertainty budget for TRS Wrist-Worn device measurement for anechoic chamber method for NR FR1 bands

| UID | Uncertainty Source | Comment | Uncertainty Value [dB] |  | Prob Distr | Div | ci | Standard Uncertainty [dB] |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  | Below 3GHz | Above 3GHz |  |  |  | Below 3GHz | Above 3GHz |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |  |
| 1 | Mismatch of transmitter chain | Гreceiver < 0.29 Гmeasurement antenna < 0.5 Cable attenuation > 3dB | 0.22 | 0.22 | U-shaped | 1.41 | 1 | 0.16 | 0.16 |
| 2 | Insertion loss of transmitter chain | Systematic with Stage 1 (=> cancels) | 0 | 0 | Rectangular | 1.73 | 1 | 0.00 | 0.00 |
| 3 | Influence of the measurement antenna cable | Systematic with Stage 1 (=> cancels) | 0 | 0 | Rectangular | 1.73 | 1 | 0.00 | 0.00 |
| 4 | Communication Tester: uncertainty of the absolute output level | From datasheet of communication tester | 1.30 | 1.30 | Normal | 2.00 | 1 | 0.65 | 0.65 |
| 5 | Sensitivity measurement: output level step resolution | Systematic error that can be corrected | 0 | 0 | Rectangular | 1.73 | 1 | 0.00 | 0.00 |
| 6 | Measurement distance | d=1.6m, Δd=0.05m | 0.27 | 0.27 | Rectangular | 1.73 | 1 | 0.12 | 0.12 |
| 7 | Quality of quiet zone | Surface standard deviation of power measurements in ripple test | 0.5 | 0.5 | Actual | 1.00 | 1 | 0.50 | 0.50 |
| 8 | DUT sensitivity drift | Drift measurement | 0.2 | 0.2 | Rectangular | 1.73 | 1 | 0.12 | 0.12 |
| 9 | Uncertainty related to the use of phantoms | Material Dielectric Constant, Material Conductivity, Geometry/Shape (incl. spacer), Forearm | 0.74 | 0.74 | Rectangular | 1.73 | 1 | 0.43 | 0.43 |
| 10 | Coarse sampling grid | Sampling grids per Table A.4.2.12-1 | 0.10 | 0.18 | Actual | 1.00 | 1 | 0.10 | 0.18 |
| 11 | Random uncertainty | Fixed MU to account for all the unknown, unquantifiable, etc. uncertainties including digital error rate | 0.4 | 0.4 | Normal | 2.00 | 1 | 0.20 | 0.20 |
| 12 | Frequency Response | Included in the output level step resolution | 0 | 0 | Rectangular | 1.73 | 1 | 0.00 | 0.00 |
| Stage 1: Calibration measurement, network analyzer method |  |  |  |  |  |  |  |  |  |
| 13 | Uncertainty of network analyzer | From datasheet of VNA with assessed transmission coefficients | 0.2 | 0.5 | Normal | 2.00 | 1 | 0.10 | 0.25 |
| 14 | Mismatch of transmitter chain | Taken into account in VNA setup uncertainty | 0 | 0 | U-shaped | 1.41 | 1 | 0.00 | 0.00 |
| 15 | Insertion loss of transmitter chain | Systematic with Stage 2 (=> cancels) | 0 | 0 | Rectangular | 1.73 | 1 | 0.00 | 0.00 |
| 16 | Mismatch in the connection of calibration antenna | Taken into account in VNA setup uncertainty | 0 | 0 | U-shaped | 1.41 | 1 | 0.00 | 0.00 |
| 17 | Influence of the calibration antenna feed cable | Gain calibration with dipole | 0.3 | 0.3 | Rectangular | 1.73 | 1 | 0.17 | 0.17 |
| 18 | Influence of the measurement antenna cable | Systematic with Stage 2 (=> cancels) | 0 | 0 | Rectangular | 1.73 | 1 | 0.00 | 0.00 |
| 19 | Uncertainty of the absolute gain/ radiation efficiency of the calibration antenna | Calibration report with traceability to a National Metrology Institute | 0.58 | 0.58 | Normal | 2.00 | 1 | 0.29 | 0.29 |
| 20 | Measurement distance | Dipole: phase centre is aligned with axis of rotation | 0 | 0 | Rectangular | 1.73 | 1 | 0.00 | 0.00 |
| 21 | Quality of quiet zone | Peak-to-null ripple | 0.5 | 0.5 | Rectangular | 1.73 | 1 | 0.29 | 0.29 |
| Combined Standard Uncertainty [dB] |  |  |  |  |  |  |  | 1.07 | 1.11 |
| Expanded Uncertainty [dB] (Confidence interval of 95 %) |  |  |  |  |  |  |  | 2.10 | 2.17 |
| 22 | Systematic Error related to TRS grids | mean error | 0 | 0.00 | Actual | 1.00 | 1 | 0.00 | 0.00 |
| Total Expanded Uncertainty [dB] (Confidence interval of 95 %) |  |  |  |  |  |  |  | 2.10 | 2.17 |

< Unchanged Text Deleted >

Annex D (normative): Maximum uncertainty of test system and test tolerance

# D.1	Maximum uncertainty of test system

- The maximum acceptable uncertainty of the Test System is specified in Table D.1-1 for each test, where appropriate. The Test System shall enable each test to be measured with an uncertainty not exceeding the specified values. All ranges and uncertainties are absolute values and are valid for a confidence level of 95 %.

- In Annex A.4, the estimation of measurement uncertainty is defined.

- Table D.1-1: Maximum Test System uncertainties for TRP hand only (browsing mode) tests using anechoic chamber method for NR FR1 bands

| - Clause | - Maximum Test System Uncertainty (MTSU) | - Derivation of Test System Uncertainty |
| --- | --- | --- |
| - 6.2.1.1.1 Total Radiated Power (TRP) for FR1 NR Standalone (SA) in Browsing Mode with Hand Phantom | 1.64 dB : < 3 GHz 1.73 dB >= 3 GHz | - Refer to Annex A.4 |
| - 7.2.1.1.1 Total Radiated Sensitivity (TRS) for FR1 NR Standalone (SA) in Browsing Mode with Hand Phantom | 2.06 dB : < 3 GHz 2.13 dB >= 3 GHz | - Refer to Annex A.4 |

- Table D.1-2: Maximum Test System uncertainties for TRP Beside Head and Hand (Talk mode) tests using anechoic chamber method for NR FR1 bands

| - Clause | - Maximum Test System Uncertainty (MTSU) | - Derivation of Test System Uncertainty |
| --- | --- | --- |
| - 6.2.1.2.1 Total Radiated Power (TRP) for FR1 NR Standalone (SA) in Talk Mode with Head and Hand Phantom | 1.88 dB : < 3 GHz 1.96 dB >= 3 GHz | - Refer to Annex A.4 |
| - 7.2.1.2.1 Total Radiated Sensitivity (TRS) for FR1 NR Standalone (SA) in Talk Mode with Head and Hand Phantom | 2.25 dB : < 3 GHz 2.31 dB >= 3 GHz | - Refer to Annex A.4 |

### D.2	Test Tolerances

- Test tolerances in Table D.2-1 are used to relax the Minimum Requirements in the present document and to derive the Test Requirements.

- Table D.2-1: Test Tolerances for TRP hand only (browsing mode) tests using anechoic chamber method for NR FR1 bands

| - Clause | - Test Tolerance (TT) |
| --- | --- |
| - 6.2.1.1.1 Total Radiated Power (TRP) for FR1 NR Standalone (SA) in Browsing Mode with Hand Phantom | TT = 0.62 x MTSU 1.02 dB : < 3 GHz 1.07 dB >= 3 GHz |
| - 7.2.1.1.1 Total Radiated Sensitivity (TRS) for FR1 NR Standalone (SA) in Browsing Mode with Hand Phantom | TT = 0.62 x MTSU 1.28 dB : < 3 GHz - 1.32 dB >= 3 GHz |

< End of Changes >