3GPP TSG RAN WG5 #80	R5-184429r1

Gothenburg, Sweeden, 20 - 24, August 2018

Agenda Item:	5.3.13.17

Source:	Anritsu

Title:	Offset in quality of quiet zone evaluation

Document for:	Agreement

1	Introduction

This paper points out a overlooked point of the current quality of quiet zone evaluation method specified in Annex D of TR38.810[1]. The quality of quiet zone is currently specified as a standard deviation, which is a deviation from a mean value, however, the mean value should have a offset from the reference level to be calibrated by a reference antenna. But the offset is not taken into account in the current MU evaluation. This paper proposes two ways to sort out the issue; one way is to adjust the signal level and the total MU, the other way is to change the definition of quality of quiet zone to calculate the deviation from the reference level. The issue pointed out by this paper is a common issue for both DFF and IFF.

2	Discussion

2.1	Overlooked point of current quality of quiet zone evaluation

According to Annex D of TR38.810[1], the quality of quiet zone is defined as a standard variation. The quality of quiet zone for EIRP and TRP would be calculated as sEIRP and sTRP respectively with the following equations. In all the equations below, all the notations of EIRP and TRP are in the unit of dB, i.e. log scale.

--------------------------------------------------------------------------------------------------------------------------

- Quality of quiet zone for EIRP:

EIRP=1KP×Kdirp=1KPdir=1Kdir10∙log1010EIRPθp,dir10+10EIRPϕp,dir10

sEIRP=1KP×Kdir-1p=1KPdir=1Kdir10∙log1010EIRPθp,dir10+10EIRPϕp,dir10-EIRP2

- Quality of quiet zone for TRP:

TRPp,dir= π2NMn=0N-1m=0M-110EIRPθp,dir;θn,ϕm10+10EIRPϕp,dir;θn,ϕm10∙sinθn

TRP=1KP×Kdirp=1KPdir=1KdirTRPp,dir

sTRP=1KP×Kdir-1p=1KPdir=1KdirTRPp,dir-TRP2

Here, the following notations are used.

EIRPθp,dir;θndir,ϕmdir≡EIRPθp,dir

θndir,ϕmdir→angle correspond to beam direction:dir

p		: index of reference-AUT position

dir		: index of reference-AUT beam direction

KP		: number of reference-AUT positions

Kdir	: number of reference-AUT beam directions

--------------------------------------------------------------------------------------------------------------------------

When calibration is performed with a reference antenna in the replacement method, the signal level is calibrated with regard to the centre of rotation. When the quality of quiet zone is evaluated using the setup as shown in Figure 2.1-1 or Figure 2.1-2, the signal level with the reference-AUT placed at P1 should correspond to the reference level. In other words, the signal levels “EIRPRef” and “TRPRef” calculated with the following equations are the reference level for EIRP and TRP respectively.

EIRPRef=1Kdirdir=1Kdir10∙log1010EIRPθP1,dir10+10EIRPϕP1,dir10

TRPRef=1Kdirdir=1KdirTRPP1,dir

As the measurement antenna gain would be higher at around the centre direction than at some off-centre directions, the signal levels corresponding to the reference levels “EIRPRef” and “TRPRef” should be higher than the mean levels “EIRP” and “TRP” which are averaged over all the reference-AUT positions. In short, the current definitions of quality of quiet zone are the deviations from the mean levels (EIRP, TRP) which have offsets from the reference levels(EIRPRef, TRPRef) as calculated in the equations below, but the offsets are not taken into account in the current MU estimations in TR38.810[1]. This issue was pointed out in [2] as well.

OFFSETQQZ_EIRP = EIRP- EIRPRef

OFFSETQQZ_TRP = TRP- TRPRef

Observation1: The current MU calculations in TR38.810[1] does not take account of the offsets of mean EIRP and TRP level from the reference level, which would not be negligible.

Calibration point

Figure 2.1-1  Influence of measurement-antenna beam pattern against reference-AUT positions in DFF

Calibration point

Figure 2.1-2  Influence of measurement-antenna beam pattern against reference-AUT positions in IFF

2.2	How to solve the issue

This section shows the following two ways to solve the issue pointed out in the previous section, whose pros and cons are shown in Tale 2.2-1 and examples are shown in A.1.

- Signal-level & total-MU adjustment

If the measurement results of EIRP and TRP w.r.t an actual DUT having antennas arranged in the same way as in the quality of quiet zone evaluation method are adjusted by the offset according to the equations below, the current values of quality of quiet zone can be used as it is for its MU evaluation. By this adjustment, the signal-level histogram shall be changed as shown in Figure 2.2-1, which means that the signal level deviates not only towards the minus direction but also the plus direction.

EIRPadjθθn,ϕm=EIRPθθn,ϕm- OFFSETQQZ_EIRP

EIRPadjϕθn,ϕm=EIRPϕθn,ϕm- OFFSETQQZ_EIRP

TRPadj= TRP- OFFSETQQZ_TRP

However, as the signal-level distribution should be changed depending on antenna arrangement in a DUT, the offset to be adjusted should also be dependent on each DUT. The offset in quality of quiet zone evaluation (OFFSETQQZ_EIRP ) can be regarded as the biggest. If an AUT is only placed at the centre of rotation, the offset shall be zero, which is the smallest offset. If we could assume all the variations of actual DUTs, the distribution of the offset could be estimated so that we can determine 95%-confidence lower and upper offset limit, which correspond to the 1.96-sigma standard deviation in the normal distribution.

Actually, since it would be difficult to estimate the distribution of the offsets across all the current and future DUTs in the world, we propose to perform EIRP/TRP-adjustment and to calculate the MU in the following way.

~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~

- Adjustment:

EIRPadjθθn,ϕm=EIRPθθn,ϕm- OFFSETQQZ_EIRP 2

EIRPadjϕθn,ϕm=EIRPϕθn,ϕm- OFFSETQQZ_EIRP 2

TRPadj= TRP- OFFSETQQZ_TRP 2

- MU calculation:

Final Total EIRP MU= Total expanded EIRP MU calculated in the budget table - OFFSETQQZ_EIRP 2

Final Total TRP MU= Total expanded TRP MU calculated in the budget table - OFFSETQQZ_TRP 2

~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~~

If the adjustment is performed by half the offsets “OFFSETQQZ_EIRP 2” and “OFFSETQQZ_TRP 2” for EIRP and TRP respectively, the maximum remaining offset would be “±OFFSETQQZ_EIRP 2” and “±OFFSETQQZ_TRP 2” w.r.t all the DUTs to be placed within the quiet zone.  In the MU-budget table, the standard deviations of the random values are to be combined by RSS, which is based on the idea that any random distribution should make a normal distribution as long as a lot of random values are summed up; the principal is called as “Central limit theorem”. However, even though offsets are summed up, it just shift the centre of the normal distribution. And also the standard deviation of the normal distribution shall be multiplied with 1.96 to derive an expanded uncertainty with 95%-confidence, but 1.96 should not be multiplied to the offset.  So in the MU calculation shown above, the maximum remaining offsets of “- OFFSETQQZ_EIRP 2” and “- OFFSETQQZ_TRP 2” are just added to the total expanded MUs as the last step to derive the final total MUs.

EIRP or TRP

Probability

Reference level

EIRP or TRP

Probability

Reference level

Signal level adjustment

Figure 2.2-1  Signal level adjustment to solve the offset issue

- Definition change of quality of quiet zone calculation

If the definitions of quality of quiet zone are to be changed so that they represent the deviations from the reference level as the equations below, the offsets shall be included within the evaluation of quality of quiet zone. But in this case, the values for quality of quiet zone would be increased than the current estimation. As for the derivation of the equations, refer to A.2.

- Quality of quiet zone for EIRP(modified):

sEIRP'=1KP×Kdir-1p=1KPdir=1Kdir10∙log1010EIRPθp,dir10+10EIRPϕp,dir10-EIRPRef2≈1KP×Kdir-1p=1KPdir=1Kdir10∙log1010EIRPθp,dir10+10EIRPϕp,dir10-EIRP2+OFFSETQQZ_EIRP 2= sEIRP2+OFFSETQQZ_EIRP 2

- Quality of quiet zone for TRP(modified):

sTRP'=1KP×Kdir-1p=1KPdir=1KdirTRPp,dir-TRPRef2≈1KP×Kdir-1p=1KPdir=1KdirTRPp,dir-TRP2+OFFSETQQZ_TRP 2= sTRP2+OFFSETQQZ_TRP 2

Table 2.2-1  Pros and Cons for each way of solving the issue

|  | Pros | Cons |
| --- | --- | --- |
| Signal-level & total-MU adjustment | It aligns with the theoretical background, such as the central limit theorem. The increase of the total MU would be as small as possible. | New process of signal adjustment shall be necessary. New MU-estimation step of the total-MU adjustment by the offset would be necessary. |
| Definition change of quality of quiet zone calculation | The signal level processing can be as it is. New MU-estimation step of the total-MU adjustment by the offset would not be necessary. | The values of quality of quiet zone should be bigger than the current estimation. It doesn’t align with the theoretical background, such as the central limit theorem. |

Proposal1: Following 2 aspects are further studied.

- Treatment of systematic error (offsets of mean EIRP and TRP level from the reference level) with the MU values.

- Applicability of the two solutions proposed in 2.2.

3	Conclusion

In this paper, the following observation and proposal are made.

Observation1: The current MU calculations in TR38.810[1] does not take account of the offsets of mean EIRP and TRP level from the reference level, which would not be negligible.

Proposal1: Following 2 aspects are further studied.

- Treatment of systematic error (offsets of mean EIRP and TRP level from the reference level) with the MU values.

- Applicability of the two solutions proposed in 2.2.

4	References

# [1] TR38.810 v2.2.0, “Study on test methods for New Radio”, RAN WG4, 2018-07

# [2] R4-1800327, “MU in mmWave Gray-box Measurements”, Anritsu, RAN4 AH-1801

A	Appendix

A.1	Examples of MU budget sheet modification

Here, the examples of MU budget tables modified from Table B.1.1.3-1 of [1] are shown in Table A.1-1 and Table A.1-2, assuming the offsets below.

OFFSETQQZ_EIRP = -0.88 [dB]

OFFSETQQZ_TRP = -0.58 [dB]

Table A.1-1: Uncertainty assessment for EIRP and TRP measurement (D = 5 cm)

modified by “Signal-level &total-MU adjustment”

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] | Offset in EIRP [dB] | Offset in TRP [dB] |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |
| 1 | Positioning misalignment | 0.50 | Rectangular | 1.73 | [0.29] |  |  |
| 2 | Measure distance uncertainty | 1.00 | Rectangular | 1.73 | [0.58] |  |  |
| 3 | Quality of quiet zone (NOTE 2) | 1.50 | Actual | 1.00 | [1.50] | -0.44 | -0.29 |
| 4 | Mismatch (NOTE 3) | 1.30 | Actual | 1.00 | [1.30] |  |  |
| 5 | Absolute antenna gain uncertainty of the measurement antenna | 0.00 | Normal | 2.00 | 0.00 |  |  |
| 6 | Uncertainty of the RF power measurement equipment (NOTE 4) | 2.16 | Normal | 2.00 | [1.08] |  |  |
| 7 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |  |  |
| 8 | Amplifier uncertainties | 2.00 | Normal | 2.00 | 1.00 |  |  |
| 9 | Random uncertainty | 0.40 | Rectangular | 1.73 | [0.23] |  |  |
| 10 | Influence of the XPD | 0.68 | U-shaped | 1.41 | 0.48 |  |  |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |
| 11 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |  |  |
| 12 | Reference antenna positioning misalignment | 0.29 | Rectangular | 1.73 | 0.17 |  |  |
| 13 | Quality of quiet zone for calibration process (NOTE 2) | 1.50 | Actual | 1.00 | [1.50] | 0.00 | 0.00 |
| 14 | Amplifier uncertainties | 0.00 | Normal | 2.00 | 0.00 |  |  |
| 15 | Uncertainty of the Network Analyzer | 0.40 | Normal | 2.00 | 0.20 |  |  |
| 16 | Reference antenna feed cable loss measurement uncertainty | 0.29 | Rectangular | 1.73 | 0.17 |  |  |
| 17 | Uncertainty of an absolute gain of the calibration antenna | 1.60 | Normal | 2.00 | [0.80] |  |  |
| 18 | Positioning and pointing misalignment between the reference antenna and the receiving antenna | 0.35 | Rectangular | 1.73 | [0.20] |  |  |
| 19 | Phase centre offset of calibration antenna | 0.62 | Rectangular | 1.73 | [0.36] |  |  |
| EIRP Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | [6.20] |  |  |
| EIRP Expanded uncertainty(1.96σ - confidence interval of 95 %) + abs(total offset) [dB] |  |  |  |  | [6.64] |  |  |
| TRP Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | [5.37] |  |  |
| TRP Expanded uncertainty(1.96σ - confidence interval of 95 %) + abs(total offset) [dB] |  |  |  |  | [5.66] |  |  |
| NOTE 1: The impact of phase variation on EIRP is FFS. NOTE 2: The quality of quiet zone is different for EIRP and TRP. For TRP, the standard uncertainty is [1dB]; for EIRP, the standard uncertainty of quiet zone is [1.5dB]. NOTE 3: The analysis was done only for the case of operating at max output power, in-band, non-CA. NOTE 4: The assessment assumes maximum DUT output power. |  |  |  |  |  |  |  |

sEIRP'= sEIRP2+OFFSETQQZ_EIRP 2=1.502+0.882=1.74 [dB]

sTRP'= sTRP2+OFFSETQQZ_TRP 2= 1.002+0.582=1.16 [dB]

Table A.1-2: Uncertainty assessment for EIRP and TRP measurement (D = 5 cm)

modified by “Definition change of quality of quiet zone calculation”

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment | 0.50 | Rectangular | 1.73 | [0.29] |
| 2 | Measure distance uncertainty | 1.00 | Rectangular | 1.73 | [0.58] |
| 3 | Quality of quiet zone (NOTE 2) | 1.50 | Actual | 1.00 | [1.74] |
| 4 | Mismatch (NOTE 3) | 1.30 | Actual | 1.00 | [1.30] |
| 5 | Absolute antenna gain uncertainty of the measurement antenna | 0.00 | Normal | 2.00 | 0.00 |
| 6 | Uncertainty of the RF power measurement equipment (NOTE 4) | 2.16 | Normal | 2.00 | [1.08] |
| 7 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 8 | Amplifier uncertainties | 2.00 | Normal | 2.00 | 1.00 |
| 9 | Random uncertainty | 0.40 | Rectangular | 1.73 | [0.23] |
| 10 | Influence of the XPD | 0.68 | U-shaped | 1.41 | 0.48 |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 11 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 12 | Reference antenna positioning misalignment | 0.29 | Rectangular | 1.73 | 0.17 |
| 13 | Quality of quiet zone for calibration process (NOTE 2) | 1.50 | Actual | 1.00 | [1.50] |
| 14 | Amplifier uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 15 | Uncertainty of the Network Analyzer | 0.40 | Normal | 2.00 | 0.20 |
| 16 | Reference antenna feed cable loss measurement uncertainty | 0.29 | Rectangular | 1.73 | 0.17 |
| 17 | Uncertainty of an absolute gain of the calibration antenna | 1.60 | Normal | 2.00 | [0.80] |
| 18 | Positioning and pointing misalignment between the reference antenna and the receiving antenna | 0.35 | Rectangular | 1.73 | [0.20] |
| 19 | Phase centre offset of calibration antenna | 0.62 | Rectangular | 1.73 | [0.36] |
| EIRP Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | [6.42] |
| TRP Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | [5.48] |
| NOTE 1: The impact of phase variation on EIRP is FFS. NOTE 2: The quality of quiet zone is different for EIRP and TRP. For TRP, the standard uncertainty is [1dB]; for EIRP, the standard uncertainty of quiet zone is [1.5dB]. NOTE 3: The analysis was done only for the case of operating at max output power, in-band, non-CA. NOTE 4: The assessment assumes maximum DUT output power. |  |  |  |  |  |

A.2	Derivation of Equations for sEIRP' and sTRP'

By replacing EIRP or TRP with x, we can derive the following equations.

σx2=Ex-x2

=Ex2-2xx+x2

=Ex2-2xEx+x2

=Ex2-x2

OFFSETx = x- xref

σx'2=Ex-xref2

=Ex2-2xrefx+xref2

=Ex2-2xrefEx+xref2

=Ex2-2x-OFFSETx Ex+x-OFFSETx 2

= Ex2-2x2+2OFFSETx x+x2-2OFFSETx x+OFFSETx 2

=Ex2-x2+OFFSETx 2

= σx2+OFFSETx 2

If we assume that “ sx2 ≈σx2” is approximately true, we can derive the following equation.

sx'≈ sx2+OFFSETx 2

Here, sx stands for sample standard deviation and σx stands for population standard deviation.

sx= 1Nx-1i=1Nxx-x2

σx= 1Nxi=1Nxx-x2