3GPP TSG-RAN5 #85	R5-198444r2

# Reno, USA, 18th – 26th Nov 2019

Title:						Proposal of MU elements for FR2 spurious emission

Source:			Anritsu

Agenda Item:			5.3.5.17

Document for:			Discussion and Endorsement

# 1.	Introduction

Some of MU terms for spurious emission test cases are TBD though the original target date was over. In this paper we propose some remaining MU elements for spurious measurement uncertainties. Note that quality of quiet zone for spurious emission test case is treated in the separated paper [7].

# 2.	Discussion

# 2.1	Uncertainty of the RF power measurement equipment

For this term, as the main factor of this term is spectrum analyzer’s absolute power uncertainty, then we can use the same number as for in-band when the frequency is similar.  Measurement uncertainty data from actual SAs like [1], [2] and [3] are shown below, quoting the data-sheet available in their web sites.

[1]

[2]

(Represented with baseline uncertainty for 50MHz + frequency response)

[3]

For 23.45GHz <= f <= 40.8GHz, it can be straightforward to apply the same values as for in-band measurement. After the discussion in the RAN5#85 meeting among TE vendors, it is proposed to apply 2.73dB for 23.45GHz <= f <= 40.8GHz because of the difference of measurement device used for measuring spurious emission (signal analyzer for in-band signal vs spectrum analyzer for spurious emission).

Proposal 2.1.1: For “Uncertainty of the RF power measurement equipment “ for spurious emission for IFF, apply [2.73] dB for 23.45GHz <= f <= 40.8GHz

For 6Ghz <= f <= 12.75GHz, our proposal is still to apply the specification values. Seeing the balance of each spec, we propose 2.0dB for this frequency range.

Proposal 2.1.2: For “Uncertainty of the RF power measurement equipment “ for spurious emission for IFF, apply [2.0] dB for 6GHz <= f <= 12.75GHz

For 12.75Ghz <= f <= 23.45GHz, our proposal is still to apply the specification values. However, some values from [2] and [3] are already bigger than 2.16dB. So, we propose to apply the same level as in-band frequency, which is 2.16dB for 12.75GHz to 23.45GHz.

Proposal 2.1.3: For “Uncertainty of the RF power measurement equipment “ for spurious emission for IFF, apply [2.16] dB for 12.75GHz <= f <= 23.45 GHz

For frequencies above 40.8GHz <= f <= 66GHz, [1] and [2] show specification values, but [3](Anritsu) does not have specs(only norm value up to 44.5Ghz).  With our internal study specific to conformance test system, our view is 2.9dB is the estimated uncertainty, which is close to the value in [1]. After the discussion in the RAN5#85 meeting among TE vendors, as the values in [1] is for Preamp=OFF and is not applicable for spurious measurement, it is proposed to apply 3.1dB.

Proposal 2.1.4: For “Uncertainty of the RF power measurement equipment “ for spurious emission for IFF, apply [3.1] dB for 40.8GHz <= f <= 66 GHz

# 2.2	Amplifier Uncertainty

# The values for in-band measurement was discussed in [4] and calculated with 0.0565 dB/ºC * (55-18)ºC = 2.1 dB based on the some investigation of commercially available amplifiers as quoted below.

Table 1: Examples of amplifier gain variation with temperature at mmWave frequencies

| Manufacturer | Model | Amplifier type | Frequency range [GHz] | Output P1dB [dBm] | Gain [dB] | Temperature operating range [ºC] | Gain variation with temperature in the operating range |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| RF-Lambda | RLNA26G40GB | Connectorized | 20-40 | (15, 22) | (37, 45) | (-45, +55) | ± 3 dB | ± 0.03 dB/ºC |
| RF-Lambda | RFLUPA20G47GA | Connectorized | 20-47 | (22, 24) | (31, 38) | (-45, +85) | ± 3 dB | ± 0.023 dB/ºC |
| Analog Devices | HMC5805ALS | SMD part | DC-40 | (19, 24.5) | (9, 12.5) | (-40, +85) | ± 3.125dB | ± 0.025 dB/ºC |
| Analog Devices | HMC943APM5E | SMD part | 24-34 | (29, 33) | (20.5, 23) | (-40, +85) | ± 5 dB | ± 0.04 dB/ºC |
| Keysight Technologies | 83050A | Connectorized | 2-50 | 13 | 21 | (0, 55) | ± 4.95 dB | ± 0.09 dB/ºC |

For spurious emission test case, the same reasoning can be applied. We propose to apply 2.1dB for  6GHz <= f <= 40.8GHz.

Proposal 2.2.1 : For “Amplifier uncertainty” for spurious emission for IFF, apply 2.1dB for 6GHz  <=  f <= 40.8GHz

TBD for 40.8GHz and above.

# 2.3	Random Uncertainty

# There was no special technical basis for this value for this term. Mainly this term includes some other MU factors like repeatability. In in-band measurement, 0.5dB was used. As we don’t have clear reason to change this value to different value for spurious emission test case, we propose the same value for spurious emission test for entire frequency range.

Proposal 2.3.1 : For “Random Uncertainty” for spurious emission for IFF , apply 0.5dB for 6GHz <= f <= 80GHz

# 2.4	Influence of XPD

# In in-band measurement, -25dB XPD is assumed. In spurious-emission test case, it can be assumed that wider frequency(6- 80GHz) range needs to be measured using one or more measurement antennas, depending on the implementation strategy considering the balance of e.g. TE cost, complexity, dynamic range, accuracy. As XPD depends on the relative frequency bandwidth of the antenna coverage, our view is that XPD can become worse than in-band antenna if we think about the balance of number of antennas vs cost/complexity etc. With our view, -15dB XPD is reasonable assumption for calculating influence of XPD for spurious emission, which can be calculated by 10log10(1+10-15/10)~ 0.14 [dB].

Proposal 2.4.1 : For “Influence of XPD” for spurious emission for IFF, apply [0.14] dB for 6GHz <= f <= 80GHz

# 2.5	Misalignment of positioning System

# This term is derived from the positioner stop accuracy, and the amount of uncertainty depends on the HPBW of the measurement antenna.  In in-band measurement, this term is considered as ignorable level hence 0.0dB was adopted. Selection of HPBW for spurious measurement antenna is up to TE implementation strategy, balancing mainly about dynamic range and accuracy (e.g. QoQZ) etc. With our view, the HPBW of measurement antennas for spurious emission in the practical implementation can still kept wide enough to regard this term keep ignorable.

Proposal 2.5.1 : For “Misalignment of positioning system “ for spurious emission for IFF, apply 0.0 dB

# 2.6	Uncertainty of the Network Analyser

# Only the value for frequencies same as in-band measurement (23.45~40.8GHz) is set as TBD. If the spurious and in-band measurement are measured using the common path, the same value as for in-band should be applied. If independent/dedicated path for spurious measurement from in-band measurement is used, as typically the path loss will be smaller in such dedicated path, then the network analyzer uncertainty could be decreased for some level. But still, the same value as for in-band can be used as worst case values.

Proposal 2.6.1 : For “Uncertainty of the Network Analyser” for spurious emission for IFF, apply 0.73 dB for 23.45 <= f <= 40.8GHz

# 2.7	Uncertainty of the absolute gain of the calibration antenna

# This term is agreed in the previous meeting with 1.7dB for above 40.8GHz, however it is still not determined for f <= 40.8GHz. As this term depends on the performance of antenna calibration labs, we need to see the actual achievable uncertainties from the various calibration labs.

# Some uncertainties are investigated and are shown in Table 1. There are several calibration methods such as SAE ARP958, ANSI C63.5, CISPR 16-1-6 and some minor variations and different uncertainties are listed even in the same calibration labs.

Table 1 Performance of antenna calibration labs

| Lab | Source (Scope document, certification document, etc) | method | Low[GHz] | High[GHz] | Uncer(k=2) |
| --- | --- | --- | --- | --- | --- |
| A | A2LA 2123.01 | SAE ARP 958, d=1m | 0.7 | 50 | 0.45 |
| A | A2LA 2123.01 | SAE ARP 958, d=3m | 0.7 | 50 | 0.48 |
| A | A2LA 2123.01 | ANSI C63.5, CISPR 16-1-6, standard site method, three antenna method, d=1m,3m | 0.7 | 18 | 0.37 |
| A | A2LA 2123.01 | ANSI C63.5, CISPR 16-1-6, standard site method, three antenna method, d=1m,3m | 18 | 26.5 | 0.42 |
| A | A2LA 2123.01 | ANSI C63.5, CISPR 16-1-6, standard site method, three antenna method, d=1m,3m | 26.5 | 50 | 0.47 |
| B | A2LA 2067.01 | SSM per ANSI C63.5, TAM per CISPR 16-1-16 | 0.8 | 18 | 1.1 |
| B | A2LA 2067.01 | SSM per ANSI C63.5, TAM per CISPR 16-1-16 | 18 | 40 | 1.1 |
| B | A2LA 2067.01 | SAE ARP 958 | 0.8 | 18 | 1.2 |
| B | A2LA 2067.01 | SAE ARP 958 | 18 | 40 | 1.4 |
| C | Calibration certification | ANSI C63.5, SSM d=3m | 1 | 10 | 1.5 |
| C | Calibration certification | ANSI C63.5, SSM d=3m | 10 | 18 | 1.8 |
| C | A2LA 2127.01 | ANSI C63.5, SSM, CSISPR 16-1-6(TAM), Network Analyzer | 1 | 20 | 0.37 |
| C | A2LA 2127.01 | ANSI C63.5, SSM, CSISPR 16-1-6(TAM), Spectrum Analyzer | 1 | 20 | 1.2 |
| D | NVLAP 100290-0 | ANSI C63.5, SSM, d=3m | 2.6 | 12 | 1.4 |
| D | NVLAP 100290-0 | ANSI C63.5, SSM, d=3m | 12 | 18 | 1.4 |
| D | NVLAP 100290-0 | ANSI C63.5, SSM, d=3m | 18 | 26.5 | 1.8 |
| D | NVLAP 100290-0 | ANSI C63.5, SSM, d=3m | 26.5 | 40 | 2.3 |
| E | A2LA 1266.02 | ANSI C63.5, SAE ARP958 Rev. D, CISPR 16-1-6, d=1m | 1 | 18 | 0.96 |
| F | UKAS 0452 | d=1m and 3m | 1 | 18 | 1.5 |
| F | UKAS 6742 | CSISPR 16-1-6, EN 55016-1-6, Reference Antenna Method, freespace 10m or 3m factors | 0.8 | 18 | 1.5 |
| G | A2LA 2941.01 | SAE ARP958, ANSI C63.5 | 1 | 18 | 1.3 |
| H | A2LA 1207.01 | SAE ARP958 | 2 | 8 | 0.93 |
| H | A2LA 1207.01 | SAE ARP958 | 8 | 18 | 0.97 |
| H | A2LA 1207.01 | SAE ARP958 | 18 | 40 | 1.3 |
| H | A2LA 1207.01 | ANSI C63.5, SSM | 1 | 40 | 0.5 |
| H | A2LA 1207.01 | CISPR 16-1-6, SSM | 1 | 18 | 0.97 |

# A2LA scope document can be referred in [5]. Also, UKAS documents can be referred in [6].

# It is noted that A2LA scope document lists the smallest uncertainty for the laboratory achievable with nearly ideal condition, and all the documents indicate that the actual MU will become bigger by following notes. Actually, we confirmed some laboratories provided the bigger values than A2LA scope document in the certification document(e.g. lab C in Table 1)

# -----

# -----

# We see different values ranged from 0.37dB to 2.3dB in Table 1. There are some freedoms of choice for TE vendors/Test Labs to which calibration labs to use. However, considering some operational aspects such as downtime, calibration cost, etc…, we don’t need to apply the world champion value which is possible only in specific calibration sites in specific location. Hence, seeing the distribution of uncertainties in the actual calibration labs, we propose to apply 1.0dB (mid value) as reference antenna gain uncertainty.

Proposal 2.7.1 : For “Uncertainty of the absolute gain of the calibration antenna” for spurious emission for IFF, apply 1.0 dB for 6GHz <= f  <= 40.8GHz

We now think that the currently applied value for in-band measurement, which is 0.6dB, is it too small to be used for determination of MTSU as it can be considered as a world champion value for actual uncertainty. Hence, we propose to update the in-band reference antenna gain to 1.0dB as well.

Proposal 2.7.2 : Update the “Uncertainty of the absolute gain of the calibration antenna” for in-band measurement (23.45 GHz <= f <= 40.8Ghz) to 1.0dB.

# 2.8	Positioning and pointing misalignment between the reference antenna and the measurement antenna

This MU term means how well the reference antenna and measurement antenna aligns in the calibration stage, and is expected to depend on the positioner or AUT fixture positioning accuracy, and also the HPBW of reference antenna. Especially, for the case of spurious emission measurement, offset antenna can be used, and in that case the difference of antenna gain of reference antenna from its calibrated gain is the matter as depicted in Figure 1.

Figure 1 Impact of positioning and pointing misalignment for offset antenna

With our analysis, if such gain difference of reference antenna is not considered in calibration stage then the MU from this term can become bigger because the antenna beam slope is steeper at the side of the peak. One example, assuming +/-0.5deg as positioner/fixture angular uncertainty, and 25dBi SGH(standard gain horn) is shown in following Figure 2.

Figure 2 Reference antenna gain (25dBi SGH)

The width of red box is +/- 0.5deg, and height of red box means the gain difference inside +/- 0.5deg angular error. From the above example, when offset antenna is located at 5deg offset, +/- 0.83dB error for the reference antenna gain is possible. Our view is that when to utilize the offset antenna, it is necessary to change the direction of AUT to align its peak to offset measurement antenna so that this kind of uncertainty is eliminated. If such manipulation is adopted in calibration stage, then the residual uncertainty can be considered to be same as that for measurement antenna at the focal point. Assuming, 25dBi standard gain horn, and +/- 0.5deg angular accuracy, we propose +/- 0.05dB for spurious emission.

Proposal 2.8.1: For “Positioning and pointing misalignment between the reference antenna and the measurement antenna” for spurious emission for IFF, apply +/- 0.05dB

# 2.2	Assumption

| #1 | Assumption | Description |
| --- | --- | --- |
| #2 | Frequency ranges under consideration | 6 GHz to 80 GHz |
| #3 | Size of QZ for IFF | 30 cm |
| #5 | Power range for EIRP measurements considered at the conducted reference plane | N/A |
| #6 | Temperature variation impact | +18 to +28 degrees C or +15 to +25 degrees C |
| #7 | UE power class | N/A |
| #9 | Characterization for QoQZ for spurious measurements | N/A |

(Ref. R5-185986 for each assumption.)

- Conclusion

In this contribution, we give the analysis and proposals for some remaining MU terms for spurious emission test case.

RAN5 is asked to endorse following proposals.

Proposal 2.1.1: For “Uncertainty of the RF power measurement equipment “ for spurious emission for IFF, apply 2.73 dB for 23.45GHz <= f <= 40.8GHz

Proposal 2.1.2: For “Uncertainty of the RF power measurement equipment “ for spurious emission for IFF, apply 2.0 dB for 6GHz <= f <= 12.75GHz

Proposal 2.1.3: For “Uncertainty of the RF power measurement equipment “ for spurious emission for IFF, apply 2.16 dB for 12.75GHz <= f <= 23.45 GHz

Proposal 2.1.4: For “Uncertainty of the RF power measurement equipment “ for spurious emission for IFF, apply [3.1] dB for 40.8GHz <= f <= 66 GHz

Proposal 2.2.1 : For “Amplifier uncertainty” for spurious emission for IFF, apply 2.1dB for 6GHz  <=  f <= 40.8GHz

Proposal 2.3.1 : For “Random Uncertainty” for spurious emission for IFF , apply 0.5dB for 6GHz <= f <= 80GHz

Proposal 2.4.1 : For “Influence of XPD” for spurious emission for IFF, apply [0.14] dB for 6GHz <= f <= 80GHz

Proposal 2.5.1 : For “Misalignment of positioning system “ for spurious emission for IFF, apply 0.0 dB

Proposal 2.6.1 : For “Uncertainty of the Network Analyser” for spurious emission for IFF, apply 0.73 dB for 23.45 <= f <= 40.8GHz

Proposal 2.7.1 : For “Uncertainty of the absolute gain of the calibration antenna” for spurious emission for IFF, apply 1.0 dB for 6GHz <= f  <= 40.8GHz

Proposal 2.7.2 : Update the “Uncertainty of the absolute gain of the calibration antenna” for in-band measurement (23.45 GHz <= f <= 40.8Ghz) to 1.0dB.

Proposal 2.8.1: For “Positioning and pointing misalignment between the reference antenna and the measurement antenna” for spurious emission for IFF, apply +/- 0.05dB

# 4.	References

[1]  Rohde& Shwarz FSW data sheet, https://www.rohde-schwarz.com/us/product/fsw-productstartpage_63493-11793.htm

[2] Keysight N9041B data sheet, https://www.keysight.com/en/pdx-2743415-pn-N9041B/uxa-signal-analyzer-multi-touch-3-hz-to-110-ghz

[3] Anritsu MS2850A data sheet, https://www.anritsu.com/en-US/test-measurement/products/ms2850a

[4] R5-191615, “On amplifiers uncertainty”, Keysight, RAN5#82

[5] A2LA, “Search A2LA Directory of Accredited Organizations”, https://cabportal.touchstone.a2la.org/index.cfm?event=directory.index

[6] UKAS, “Search Accredited Organisations”, https://www.ukas.com/search-accredited-organisations/