Source:	HEAD acoustics GmbH

Title:	Comparison of ambient noise playback systems

Document for:	Discussion

Agenda Item:	10.6

# Introduction

During discussions of the work item SPAN [1], it became obvious that the inter-lab variability of background noise measurements according to TS 26.132 [2] in handheld hands-free mode is smaller than in handset mode. Even though the positioning of a device in handheld hands-free mode is more cumbersome than in handset mode, results obtained by the speech quality metrics were more reproducible here. The reason for this unexpected observation might be that two different background noise playback systems are used for the two modes. For handheld hands-free testing, the state-of-the-art system according to ETSI TS 103 224 is used, while for handset mode the predecessor according to ETSI ES 202 396-1 is used.

It was discussed during SA4#99, if the more advanced background noise playback system according to ETSI TS 103 224 may be used as either a replacement or an alternative solution for the ambient noise evaluations according to TS 26.132 in handset mode as well. It was noted that legacy issues should be considered, i.e. any modification of the setup should not have impacts on the requirements as defined in TS 26.131 [3]. This might be difficult to achieve due to the different structure of the noise scenarios. The binaural signals of ETSI ES 202 396-1 are not available in ETSI TS 103 224 where eight-channel recordings of complex noise fields are considered.

This contribution investigates the comparability between the two systems in narrowband (NB), wideband (WB) and super-wideband (SWB) mode for several terminals in handset mode. The applicability of the noise scenarios of TS 103 224 is evaluated. Similarities and differences of the overall results for noise suppression performance are reported.

# Noise Types

According to the descriptions of clauses 7.12.1 (NB), 8.12.1 (WB) and 9.12.1 (SWB/FB), the noise types as shown in Table 1 are used for the evaluation of noise suppression performance.

These noise scenarios are not available in the background noise system according to ETSI TS 103 224. However, the noise database here provides similar / related scenarios, which may be used instead. Table 2 shows the eight noise types which are proposed and used for the present study.

Note 1: Since the noise recordings are in 8-channel-format, only level information for channel 1 (close to left ear) and 7 (close to right ear) are provided. More information on the specific recordings can be found in clause 8 of ETSI TS 103 224 [4].

Note 2: In the noise database of TS 103 224, there is no equivalent noise condition for “mensa”. The noise “sales counter” is used instead.

Table 1: Currently used noise types acc. to ETSI ES 202 396-1

| Filename | Duration | Level |
| --- | --- | --- |
| Pub_Noise_binaural_V2 | 30 s | L: 75,0 dB(A) R: 73,0 dB(A) |
| Outside_Traffic_Road_binaural | 30 s | L: 74,9 dB(A) R: 73,9 dB(A) |
| Outside_Traffic_Crossroads_binaural | 20 s | L: 69,1 dB(A) R: 69,6 dB(A) |
| Train_Station_binaural | 30 s | L: 68,2 dB(A) R: 69,8 dB(A) |
| Fullsize_Car1_130Kmh_binaural | 30 s | L: 69,1 dB(A) R: 68,1 dB(A) |
| Cafeteria_Noise_binaural | 30 s | L: 68,4 dB(A) R: 67,3 dB(A) |
| Mensa_binaural | 22 s | L: 63,4 dB(A) R: 61,9 dB(A) |
| Work_Noise_Office_Callcenter_binaural | 30 s | L: 56,6 dB(A) R: 57,8 dB(A) |

Table 2: Investigated noise types acc. to ETSI TS 103 224

| Filename | Duration | Level Channel 1 & 7 |
| --- | --- | --- |
| Pub_handset | 30 s | 1: 77,2 dB(A) 7: 76,0 dB(A) |
| Roadnoise_handset | 30 s | 1: 72,8 dB(A) 7: 73,0 dB(A) |
| Crossroadnoise_handset | 30 s | 1: 70,6 dB(A) 7: 71,2 dB(A) |
| TrainStation_handset | 30 s | 1: 78,9 dB(A) 7: 78,8 dB(A) |
| FullSizeCar_130_handset | 30 s | 1: 68,5 dB(A) 7: 70,8 dB(A) |
| SalesCounter_handset | 30 s | 1: 66,6 dB(A) 7: 66,6 dB(A) |
| Cafeteria_handset | 30 s | 1: 70,0 dB(A) 7: 70,6 dB(A) |
| Callcenter2_handset | 30 s | 1: 60,2 dB(A) 7: 60,2 dB(A) |

# Speech quality test methods

As source speech material, 16 American English speech samples (fullband) according to ETSI TS 103 106 [5] were used for all bandwidth modes.

The test setup as described in clause 7.12.1 of TS 26.132 was used for terminals evaluated in NB mode. The analysis according to ETSI TS 103 106 [5] in NB-mode was carried out for each of the 16 samples and the results were averaged per background noise condition.

The test setup as described in clause 8.12.1 of TS 26.132 was used for terminals evaluated in WB mode. The analysis according to ETSI TS 103 106 [5] in WB-mode was carried out for each of the 16 samples and the results were averaged per background noise condition.

The test setup as described in clause 9.12.1 of TS 26.132 was used for terminals evaluated in SWB mode. Prior to the analysis of the noisy recordings, the calibration procedure described in clause 9.5 of ETSI TS 103 281 [6] was applied. The analysis according to ETSI TS 103 281 (model A) [6] in SWB-mode was then carried out for each of the 16 samples and the results were averaged per background noise condition.

The analyses in the three considered bandwidth modes lead to S-MOS, N-MOS and G-MOS result values per condition. The performance requirements and objectives of TS 26.131 are checked against the average across all eight noise types.

Note: Since only S- and N-MOS are evaluated in TS 26.131, G-MOS is not reported in the following analyses.

# Measurement rooms

The parameters of the measurement rooms evaluated in this study are described in Table 3. The equalization procedure according to ES 202 396-1 [7] for handset mode and according to TS 103 224 [4] for handheld hands-free mode was passed for all rooms. All rooms are equipped with sub-woofer setup for the playback according to ES 202 396-1 [7].

Table 3: Measurement rooms

| Name | Length [m] | Width [m] | Height [m] | Comment |
| --- | --- | --- | --- | --- |
| Room 1 | 2.40 | 3.40 | 2.05 | Same chamber type and design as Room 2 |
| Room 2 | 2.40 | 3.40 | 2.05 | Same chamber type and design as Room 1 |
| Room 3 | 2.90 | 3.10 | 2.05 | Different manufacturer than Room 1,2,4 |
| Room 4 | 1.80 | 2.40 | 2.05 | Smallest chamber |

Note: The measurement rooms shown in Table 3 were already investigated in the SPAN work item [8]. For the present investigation, only room 1 and 4 were considered, i.e. room 2 and 3 were not evaluated. In order to keep the naming / numbering convention consistent, the same names as in previous data are used in this contribution.

# Test devices

Three commercially available device under tests (DUTs), which are able to operate in NB, WB and SWB mode. Codecs and bitrates were configured according to the specified settings of the corresponding clauses in TS 26.132/131.

# Results

Table 4 shows the results for S- and N-MOS of the DUTs in all modes. The values are reported separately for the two background noise systems ES 202 396-1 (ES202) and TS 103 224 (TS103) and are averaged across the eight background noises according to Table 1 and Table 2.

Table 4: Results averaged across eight background noises

|  |  |  | S-MOS |  | N-MOS |  |
| --- | --- | --- | --- | --- | --- | --- |
|  |  | System | TS103 | ES202 | TS103 | ES202 |
| Bandwidth | DUT | Room |  |  |  |  |
| NB | DUT 1 | Room 1 | 3.90 | 4.04 | 2.64 | 2.76 |
|  |  | Room 4 | 3.81 | 3.95 | 2.70 | 2.66 |
|  | DUT 2 | Room 1 | 3.70 | 3.94 | 4.33 | 4.35 |
|  |  | Room 4 | 3.56 | 3.72 | 4.23 | 4.04 |
|  | DUT 3 | Room 1 | 3.72 | 3.78 | 2.86 | 2.77 |
|  |  | Room 4 | 3.96 | 4.07 | 2.99 | 2.69 |
| WB | DUT 1 | Room 1 | 3.99 | 4.10 | 2.95 | 3.07 |
|  |  | Room 4 | 3.94 | 4.00 | 3.05 | 3.02 |
|  | DUT 2 | Room 1 | 3.81 | 3.91 | 4.34 | 4.31 |
|  |  | Room 4 | 3.77 | 3.81 | 4.28 | 3.98 |
|  | DUT 3 | Room 1 | 3.83 | 3.93 | 3.27 | 3.46 |
|  |  | Room 4 | 3.96 | 4.01 | 3.40 | 3.35 |
| SWB | DUT 1 | Room 1 | 3.80 | 3.87 | 3.22 | 3.31 |
|  |  | Room 4 | 3.77 | 3.80 | 3.05 | 3.05 |
|  | DUT 2 | Room 1 | 3.56 | 3.81 | 4.03 | 3.96 |
|  |  | Room 4 | 3.50 | 4.17 | 4.02 | 4.27 |
|  | DUT 3 | Room 1 | 3.94 | 3.99 | 4.15 | 4.06 |
|  |  | Room 4 | 3.91 | 4.07 | 4.00 | 4.13 |

Table 5 provides the differences between the two obtained overall result values for S- and N-MOS (TS103 – ES202). It can be noticed that the sign is negative in most cases, i.e. the proposed eight noises lead consistently to more pessimistic results.

Since most of the average result values for S-MOS are in the quite narrow range of range of 3.5-4.0 (even across all bandwidths), a simple mathematical transformation (“mapping function”) or offset seems not feasible to compensate for the difference – at least based on this small amount of data.

Table 5: Difference between the two noise systems

| Bandwidth | DUT | Room | S-MOS | N-MOS |
| --- | --- | --- | --- | --- |
| NB | DUT 1 | Room 1 | -0.14 | -0.12 |
|  |  | Room 4 | -0.13 | 0.04 |
|  | DUT 2 | Room 1 | -0.24 | -0.02 |
|  |  | Room 4 | -0.15 | 0.19 |
|  | DUT 3 | Room 1 | -0.06 | 0.09 |
|  |  | Room 4 | -0.11 | 0.31 |
| WB | DUT 1 | Room 1 | -0.11 | -0.12 |
|  |  | Room 4 | -0.06 | 0.02 |
|  | DUT 2 | Room 1 | -0.11 | 0.03 |
|  |  | Room 4 | -0.05 | 0.29 |
|  | DUT 3 | Room 1 | -0.09 | -0.19 |
|  |  | Room 4 | -0.05 | 0.05 |
| SWB | DUT 1 | Room 1 | -0.06 | -0.09 |
|  |  | Room 4 | -0.03 | 0.00 |
|  | DUT 2 | Room 1 | -0.24 | 0.07 |
|  |  | Room 4 | -0.66 | -0.26 |
|  | DUT 3 | Room 1 | -0.05 | 0.09 |
|  |  | Room 4 | -0.16 | -0.13 |

# Results per measurement room

The measured data may also be used to analyse the reproduction accuracy of the noise field reproduction systems. Table 6 shows the differences of S- and N-MOS between the two measurement rooms which were used in the investigation. In general, the difference is expected to be close to 0.0; however, differences may occur due to non-linear and time-variant behaviour of the UE and/or impact of the noise field equalization in the corresponding room.

The magnitude of the delta-values (d) of Table 6 is coded as follows:

- OK (good reproduction accuracy)			d < 0.1 MOS

- MEDIUM (difference should be investigated)		0.1 ≤ d ≤ 0.2 MOS

- NOT OK (outlier)					d > 0.2 MOS

Table 6: Differences of average results between Room 1 and 4

|  |  |  | S-MOS |  | N-MOS |  |
| --- | --- | --- | --- | --- | --- | --- |
|  |  |  | TS103 | ES202 | TS103 | ES202 |
| Bandwidth | DUT | Room |  |  |  |  |
| NB | DUT 1 | Room 1 | -0.09 | -0.10 | 0.06 | -0.10 |
|  |  | Room 4 |  |  |  |  |
|  | DUT 2 | Room 1 | -0.14 | -0.23 | -0.09 | -0.31 |
|  |  | Room 4 |  |  |  |  |
|  | DUT 3 | Room 1 | 0.24 | 0.29 | 0.14 | -0.08 |
|  |  | Room 4 |  |  |  |  |
| WB | DUT 1 | Room 1 | -0.04 | -0.10 | 0.10 | -0.04 |
|  |  | Room 4 |  |  |  |  |
|  | DUT 2 | Room 1 | -0.04 | -0.10 | -0.07 | -0.33 |
|  |  | Room 4 |  |  |  |  |
|  | DUT 3 | Room 1 | 0.12 | 0.08 | 0.13 | -0.12 |
|  |  | Room 4 |  |  |  |  |
| SWB | DUT 1 | Room 1 | -0.03 | -0.06 | -0.16 | -0.26 |
|  |  | Room 4 |  |  |  |  |
|  | DUT 2 | Room 1 | -0.06 | 0.36 | -0.01 | 0.32 |
|  |  | Room 4 |  |  |  |  |
|  | DUT 3 | Room 1 | -0.03 | 0.08 | -0.15 | 0.07 |
|  |  | Room 4 |  |  |  |  |

The data indicates that in most cases (green & yellow cases) the reproduction accuracy of the measurements in general is acceptable, but also shows several outliers. Only one outlier is seen for the noise field simulation according to ETSI TS 103 224, while seven outliers are observed for the system according to ETSI ES 202 396-1.

Due to the limited amount of data, it cannot be assessed with certainty, if the terminal or the noise simulation cause the outlier cases. However, the noise field reproduction according to ETSI TS 103 224 obviously reduces the variability.

# Results with selected noise types

Since the results for S- and N-MOS obtained by the noise simulation according to ES 202 396-1 are the basis for performance requirements and objectives in TS 26.131, it is desirable to achieve highly comparable results when modifying the measurement methodology.

For the alignment of the results obtained with the background noise scenarios of TS 103 224, results obtained by ES 202 396-1 can be considered as “reference result values”, which have to be met as close as possible. This may be achieved by:

- Selection of background noises (more or less than eight)

- Weighting of noise conditions instead of averaging (plus possible bias for better fitting)

An example for this approach is shown for SWB mode in the following. In Table 7, five out of eight background noises are selected. Weights per background noise are calculated via non-linear parameter optimization. The weights of the residual noise conditions are set to zero. In addition, a bias value for a linear combination is introduced.

Table 7: Weighting factors for results obtained by TS 103 224

| Nbr | Filename in TS 103 224 | Weights |  |
| --- | --- | --- | --- |
|  |  | SWB |  |
|  |  | S-MOS | N-MOS |
| 0 | Bias | 1.182 | -0.098 |
| 1 | Pub_handset | 0.100 | 0.110 |
| 2 | Roadnoise_handset | 0.100 | 0.236 |
| 3 | Crossroadnoise_handset | 0.000 | 0.000 |
| 4 | TrainStation_handset | 0.000 | 0.000 |
| 5 | FullSizeCar_130_handset | 0.000 | 0.000 |
| 6 | SalesCounter_handset | 0.100 | 0.300 |
| 7 | Cafeteria_handset | 0.100 | 0.300 |
| 8 | Callcenter2_handset | 0.300 | 0.100 |

Calculating S- and N-MOS with this weighted sum leads to the updated values according to Table 8. Values obtained by noise simulation according to EG 202 396-1 are not re-calculated and provided for sake of completeness (columns ES202 are identical to Table 4).

Finally, Table 9 provides the differences between all eight and the selected and weighted five noises compared to the results obtained by ES 202 396-1. It can be observed that the average/bias and root-mean-square (RMS) are significantly reduced.

Table 8: Results of reduced / weighted noise conditions

|  |  |  | S-MOS |  | N-MOS |  |
| --- | --- | --- | --- | --- | --- | --- |
|  |  | System | TS103 | ES202 | TS103 | ES202 |
| Bandwidth | DUT | Room |  |  |  |  |
| SWB | DUT 1 | Room 1 | 3.90 | 3.87 | 3.27 | 3.31 |
|  |  | Room 4 | 3.89 | 3.80 | 3.11 | 3.05 |
|  | DUT 2 | Room 1 | 3.90 | 3.81 | 4.10 | 3.96 |
|  |  | Room 4 | 3.86 | 4.17 | 4.08 | 4.27 |
|  | DUT 3 | Room 1 | 4.10 | 3.99 | 4.16 | 4.06 |
|  |  | Room 4 | 4.06 | 4.07 | 4.05 | 4.13 |

Table 9: Differences between

|  |  |  | 8 Noises |  | 5 Noises (w) |  |
| --- | --- | --- | --- | --- | --- | --- |
| Bandwidth | DUT | Room | S-MOS | N-MOS | S-MOS | N-MOS |
| SWB | DUT 1 | Room 1 | -0.06 | -0.09 | 0.03 | -0.04 |
|  |  | Room 4 | -0.03 | 0.00 | 0.09 | 0.06 |
|  | DUT 2 | Room 1 | -0.24 | 0.07 | 0.09 | 0.14 |
|  |  | Room 4 | -0.66 | -0.26 | -0.31 | -0.19 |
|  | DUT 3 | Room 1 | -0.05 | 0.09 | 0.11 | 0.11 |
|  |  | Room 4 | -0.16 | -0.13 | -0.01 | -0.07 |
|  |  | Avg. | -0.20 | -0.05 | 0.00 | 0.00 |
|  |  | RMS | 0.20 | 0.11 | 0.11 | 0.10 |

# Conclusions

This contribution presented measurement results for the ambient noise test cases of TS 26.131/132. Three mobile phones in handset mode were evaluated in NB, WB and SWB mode. For the noise reproduction, two playback systems were investigated. First analyses show that it is in general possible to reach similar, but still different average values with the state-of-the-art noise field simulation according to ETSI TS 103 224. In addition, the inter-lab/-room variability can be reduced with this system.

The initial and preliminary data indicates that further and similar investigations seem to be meaningful. Thus, the source proposed to initiate a new work item in 3GPP SA4 to investigate the differences and possibilities for the recent background noise playback system for handset mode.

# References

| [1] | 3GPP S4-171071, „New WID on Speech quality in the presence of ambient noise for super-wideband and fullband modes (SPAN),“ HEAD acoustics GmbH, ORANGE, Intel, Qualcomm Incorporated, Sony Mobile Communications, Samsung Electronics Co., Ltd., Belgrade, Serbia, 09-13 October 2017. |
| --- | --- |
| [2] | 3GPP TS 26.132, „Speech and video telephony terminal acoustic test specification,“ Release 15.0.0. |
| [3] | 3GPP TS 26.131, „Terminal acoustic characteristics for telephony; Requirements,“ Release 15.0.0. |
| [4] | ETSI TS 103 224 v1.3.1, „Speech and multimedia Transmission Quality (STQ);. A sound field reproduction method for terminal testing including a background noise database,“ 07/2017. |
| [5] | ETSI TS 103 106 v1.5.1, „Speech Quality performance in the presence of background noise: Background noise transmission for mobile terminals - Objective test methods,“ 04/2018. |
| [6] | ETSI TS 103 281 v1.1.1, „Speech quality in the presence of background noise: Objective test methods for super-wideband and fullband terminals,“ 04/2017. |
| [7] | ETSI ES 202 396-1 v1.7.1, „Speech Processing, Transmission and Quality Aspects (STQ); Speech quality performance,“ 10/2017. |
| [8] | 3GPP S4-180326, „Updated results for SPAN work item,“ HEAD acoustics GmbH, Kista, Sweden, 09-13 April 2017. |