3GPP TSG-RAN5 Meeting #105	R5-246537

Orlando, FL, USA, 18th – 22nd November 2024

Agenda item:	5.4.20

Source:	Keysight Technologies UK Ltd

Title:	PC7 MUs based on antenna array assumptions

Document for:	Discussion and Endorsement

# Introduction

This contribution intends to provide an overview of the MU elements based on the PC7 antenna array assumptions.

# Discussion

In [1], the following key proposals were endorsed for antenna array assumptions for PC7 Devices (RedCap).

Proposal 1: Use HPBW of 90/90 as default antenna assumption for PC7 RedCap. Proposal 2: Use antenna array assumptions as outlined in Table 1 for PC7 RedCap => 4x2 array as default array size for PC7 measurement grids and MU analysis.

This array assumption matches the alternate array assumption for PC3 UEs based on a vendor declaration [2], [3]. The measurement grids for the 4x2 array assumptions were already provided in [4] and incorporated in [2]. It is proposed to leverage the beam peak search, spherical coverage, and TRP MUs and measurement grids of the PC3 alternate array assumption for PC7.

Proposal 1: Leverage the beam peak search, spherical coverage, and TRP measurement grids and MU from PC3 (4x2 alternate array assumption) for PC7

This contribution intends to provide MUs for the elements that are dependent on the array assumptions, similar to what was previously presented in [6], [7]. The antenna array assumptions are the same as those presented in [5] and will not be repeated here other than presenting the 4x2 antenna array pattern plot in Figure 1.

Figure 1: 4x2 Antenna Pattern with 90°/90° HPBW.

# MU Elements: “Positioning misalignment” and “DUT repositioning”

The MU elements, other than those related to the measurement grid, depending on the antenna pattern have been identified in [8]:

Positioning misalignment

DUT repositioning (EIRP/EIS)

The approach to determine the positioning misalignment MU is based on the approach outlined in [9] with a maximum deviation from the beam peak direction due to positioning misalignment to be 0.25o. The simulation results shown in Figure 2 show that the error based on the positioning misalignment is negligible with maximum error between the theoretical beam peak and the pattern at the misaligned beam peak direction of less than 0.002dB.

Figure 2: Histogram of maximum beam peak errors for a maximum misalignment error of 0.25o.

It is therefore proposed to define the standard uncertainty of the “Positioning misalignment” MU element for PC7 to be 0dB.

Proposal 2: Define standard uncertainty of the “Positioning misalignment” MU element for PC7 to be 0dB

The approach to determine the DUT repositioning for EIRP/EIS metrics is based on the approach outlined in [10] with a maximum deviation from the beam peak direction due to DUT re-positioning to be 1o. The simulation results shown in Figure 3 show that the error based on the positioning misalignment is very small with maximum error between the theoretical beam peak and the pattern at the misaligned beam peak direction is ~0.02dB. For simplicity and to equalize MTSU of PC7 with the MTSU of PC3 (as this is the only MU element with different values between PC7 and PC3 [11]), it is proposed to define the DUT repositioning MU value for PC7 the same as the PC3 value, i.e., 0.08dB. As discussed in [10], this MU element should be considered to be 0dB for TRP and spherical coverage test cases.

Figure 3: Histogram of maximum beam peak errors for a maximum misalignment error of 1o.

Proposal 3: Define the standard uncertainty of the “DUT repositioning” MU element for PC7 to be 0.08dB for EIRP and EIS metrics (same as for PC3) and 0dB for TRP and spherical coverage.

# Conclusion

The following observations and proposals were made in this contribution.

Proposal 1: Leverage the beam peak search, spherical coverage, and TRP measurement grids and MU from PC3 (4x2 alternate array assumption) for PC7

Proposal 2: Define standard uncertainty of the “Positioning misalignment” MU element for PC7 to be 0dB

Proposal 3: Define the standard uncertainty of the “DUT repositioning” MU element for PC7 to be 0.08dB for EIRP and EIS metrics (same as for PC3) and 0dB for TRP and spherical coverage.

# References

- R5-234096 , Discussion on PC7 RedCap Antenna Assumptions for measurement grids, Qualcomm Incorporated, 3GPP TSG-RAN5 Meeting #100, August 2023

- TS 38.521-2, User Equipment (UE) conformance specification; Radio transmission and reception; Part 2: Range 2 Standalone

- TS 38.508-2, User Equipment (UE) conformance specification; Part 2: Common Implementation Conformance Statement (ICS)

- R5-213812, Measurement Grids for Optional 4x2 PC3 Antenna Array Configuration, Keysight Technologies, 3GPP TSG-RAN5 Meeting #91-e, May 2021

- R5-231774, PC5 measurement grids including the alternate antenna array assumptions, Keysight Technologies, 3GPP TSG-RAN5 Meeting #98, March 2023

- R5-211189, PC1 MUs based on the revised antenna array assumptions, Keysight Technologies, 3GPP TSG-RAN5 Meeting #90-e, March 2021

- R5-244448, PC6 MUs based on antenna array assumptions, Keysight Technologies, 3GPP TSG-RAN5 Meeting #104, August 2024

- R5-191628, MU Elements: documentation of justification and interdependence, Keysight Technologies, 3GPP TSG-RAN5 Meeting #82, February 2019

- R5-190389, On Positioning Misalignment MU, Keysight Technologies, 3GPP TSG-RAN WG5 5G NR AH #4, January 2019

- R5-190387, On MU Element for DUT Repositioning, Keysight Technologies, 3GPP TSG-RAN WG5 5G NR AH #4, January 2019

- R5-246230, PC7 FR2 MU - Discussion on non-antenna related MU factors, Keysight Technologies, 3GPP TSG-RAN5 Meeting #105, November 2024