3GPP TSG-RAN5 Meeting #105	R5-247166

# Orlando, USA, 18th Nov. – 22nd Nov. 2024

Title:							QoQZ for QZ size 40 cm and ETC

Source:			Anritsu

Agenda Item:			5.4.20

Document for:			Endorsement

# 1.	Introduction

In the previous meeting, the action point AP#103.21 was created to define QoQZ MU for QZ size = 40 cm and ETC. This paper repeats our measurement results and proposal of [1] to proceed the discussion.

| Action ID | sWG | Action | Responsible | Relevant Tdoc | Deadline | Status |
| --- | --- | --- | --- | --- | --- | --- |
| AP#103.22 | RF | Define QoQZ MU for 40cm QZ and ETC (up to FR2d) | Anritsu, Keysight, R&S | R5-243002 | RAN5#105 | Open |

# 2.	Discussion

The following table shows our measurement results of QoQZ MU for QZ size = 40 cm and ETC.

Table 1: Standard Deviations for EIRP/EIS metric (QZ size = 40 cm, ETC)

| Positions | QoQZ MU [dB] |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
|  | 23.450 GHz | 32.125 GHz | 40.800 GHz | 44.300 GHz | 49.000 GHz |
| P1-P7 (Stage 2) | 0.87 | 0.85 | 0.53 | 0.73 | 0.68 |
| P1 only (Stage 1) | 0.58 | 0.38 | 0.15 | 0.16 | 0.41 |

Table 2: Standard Deviations for TRP metric (QZ size = 40 cm, ETC)

| Positions | QoQZ MU [dB] |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
|  | 23.450 GHz | 32.125 GHz | 40.800 GHz | 44.300 GHz | 49.000 GHz |
| P1-P7 (Stage 2) | 0.33 | 0.31 | 0.35 | 0.50 | 0.56 |
| P1 only (Stage 1) | 0.25 | 0.13 | 0.20 | 0.49 | 0.22 |

Considering the above measurement results and commonality with the already endorsed MU values for QZ size ≤ 30 cm, we propose the following QoQZ MU for QZ size = 40 cm and ETC.

Table 3: QoQZ MU for EIRP/EIS and TRP metric for P1-P7 (Stage 2)

| Condition | QoQZ MU [dB] |  |  |  |
| --- | --- | --- | --- | --- |
|  | FR2a | FR2b | FR2c | FR2d |
| QZ size ≤ 30 cm and NTC | 0.6 | 0.6 | 0.7 | 0.7 |
| QZ size ≤ 30 cm and ETC | 0.9 | 0.9 | 0.9 | 0.9 |
| QZ size = 40 cm and NTC | 0.7 | 0.7 | 0.8 | 0.8 |
| QZ size = 40 cm and ETC | 0.9 | 0.9 | 0.9 | 0.9 |
| Note 1: Black text indicates endorsed values, and green text indicates proposed values in this paper. Note 2: In MTSU calculation for NTC, QoQZ value for QZ size ≤ 30 cm is used for QZ size = 40 cm and FR2a through FR2c. Note 3: In MTSU calculation for NTC, QoQZ value for QZ size = 40 cm is used for QZ size ≤ 30 cm and FR2d. |  |  |  |  |

Table 4: QoQZ MU for EIRP/EIS and TRP metric for P1 only (Stage 1)

| Condition | QoQZ MU [dB] |  |  |  |
| --- | --- | --- | --- | --- |
|  | FR2a | FR2b | FR2c | FR2d |
| QZ size ≤ 40 cm and NTC | 0.4 | 0.4 | 0.5 | 0.5 |
| QZ size ≤ 30 cm and ETC | 0.6 | 0.6 | 0.6 | 0.6 |
| QZ size = 40 cm and ETC | 0.6 | 0.6 | 0.6 | 0.6 |
| Note 1: Black text indicates endorsed values, and green text indicates proposed values in this paper. |  |  |  |  |

Proposal 1: Adopt the QoQZ MU values in Table 3 and Table 4.

# 3.	Conclusion

RAN5 is asked to endorse the following proposals

Proposal 1: Adopt the QoQZ MU values in Table 3 and Table 4.

# 4.	References

[1] R5-243002, “QoQZ for QZ size 40 cm”, Anritsu, RAN5#103