3GPP TSG-RAN WG4 Meeting #80bis	R4-168876

# Ljubljana, Slovenia, 10 - 14 Oct, 2016

| CR-Form-v11.1 |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CHANGE REQUEST |  |  |  |  |  |  |  |  |
|  | 37.842 | CR | 0001 | rev | 1 | Current version: | 13.0.0 |  |
| For HELP on using this form: comprehensive instructions can be found at http://www.3gpp.org/Change-Requests. |  |  |  |  |  |  |  |  |

| Proposed change affects: | UICC apps |  | ME |  | Radio Access Network | X | Core Network |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |

| Title: | Uncertainty budget format and uncertainty assessment for the radiated transmit power OTA test for the one dimensional compact range chamber method |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Source to WG: | Kathrein |  |  |  |  |  |  |  |  |  |
| Source to TSG: | RAN WG4 |  |  |  |  |  |  |  |  |  |
| Work item code: | AAS_BS_LTE_UTRA-Perf |  |  |  |  |  |  | Date: |  | 2016-10-13 |
| Category: | F |  |  |  |  |  |  | Release: |  | Rel-13 |
|  | Use one of the following categories: F (correction) A (mirror corresponding to a change in an earlier release) B (addition of feature), C (functional modification of feature) D (editorial modification) Detailed explanations of the above categories can be found in 3GPP TR 21.900. |  |  |  |  |  |  |  | Use one of the following releases: Rel-8 (Release 8) Rel-9 (Release 9) Rel-10 (Release 10) Rel-11 (Release 11) Rel-12 (Release 12) Rel-13 (Release 13) Rel-14 (Release 14) |  |
| Reason for change: |  | For one dimensional compact range chamber the uncertainty budget and assessment tables for EIRP measurement are still open. |  |  |  |  |  |  |  |  |
| Summary of change: |  | - Change of Figure 10.3.1.1.4.1-1, editorial change for alignment to the figures of the other measurement methods - Insertion of Table 10.3.1.1.4.4 -1: uncertainty budget for EIRP - Insertion of Table 10.3.1.1.4.5 -1: uncertainty assessment for EIRP - Supplement to Table 10.3.1.2-1: expanded uncertainty values of one dimensional compact range chamber - Inclusion of Appendix B.3: description of the uncertainty budget terms for EIRP - Small editorial corrections |  |  |  |  |  |  |  |  |
| Consequences if not approved: |  | Description of one dimensional compact range chamber chapter is not complete |  |  |  |  |  |  |  |  |
| Clauses affected: |  | 10.3.1.1.4, 10.3.1.2 and B.3 |  |  |  |  |  |  |  |  |
|  |  | Y | N |  |  |  |  |  |  |  |
| Other specs |  |  | X | Other core specifications |  |  |  |  |  |  |
| affected: |  |  | X | Test specifications |  |  |  |  |  |  |
| (show related CRs) |  |  | X | O&M Specifications |  |  |  |  |  |  |
| Other comments: |  |  |  |  |  |  |  |  |  |  |

## --------------Start of text proposal------------

##### 10.3.1.1.4 	One Dimensional Compact Range

###### 10.3.1.1.4.1	Description

The principle of the One Dimensional Compact Range Chamber measurement setup is shown in Figure 10.3.1.1.4.1-1. It closely resembles the standard compact range chamber. But for the generation of the plane wave in the quiet zone the proposed method uses a special one dimensional probe instead of the standard set up comprising a concave mirror and a feed horn. The main advantage is a reduced chamber size and a simplified probe system replacing the standard mirror feed horn set up.

Figure 10.3.1.1.4.1-1: One Dimensional Compact Range Chamber measurement system setup for EIRP

The probe antenna is set up of a one dimensional array of antenna modules, connected by a special receive network. Optimized amplitude and phase settings in the receive network result in a plane wave in vertical direction. As no near field to far field transformation has to be carried out, it is not required to measure a whole set of azimuth and elevation angles.

###### 10.3.1.1.4.2	Test Method limitations and scope

One Dimensional Compact Range Chamber test method is only suitable for AAS BS whose antenna array consists of a single column antenna array. For an AAS BS equipped with multi-column antenna array, this method would not be suitable due to the high amplitude uncertainty of edge column elements. High amplitude uncertainty makes testing of the declared steering angles not possible to meet EIRP accuracy requirements.

###### 10.3.1.1.4.3 	Procedure

- Connect the receive network of the compact probe to the measurement equipment.

- Calibrate the test range, using a reference antenna with standard gain installed in the quiet zone of the probe and measure the path loss between reference antenna and the measurement equipment.

- Install the DUT in the quiet zone of the probe with its manufacturer declared coordinate system reference point in the same place as the phase centre of the reference antenna. The manufacturer declared coordinate system orientation of the DUT is set to be aligned with testing system.

- Align with the required conformance steering directions.

- Set the DUT to transmit at maximum EIRP appropriate to conformance steering directions according to the manufacturer declaration.

- Measure the received power at the probe and thus the EIRP of the DUT.

- Repeat test steps 2 to 6 for all declared beams and corresponding conformance steering directions.

###### 10.3.1.1.4.4 	Uncertainty budget format

Table 10.3.1.1.4.4-1: One Dimensional Compact Range Chamber uncertainty contributions

for AAS BS EIRP measurement

| UID | Description of uncertainty contribution | Details in annex |
| --- | --- | --- |
| Stage 2: DUT measurement |  |  |
| 1 | Misalignment DUT and pointing error | B3-1 |
| 2 | Standing wave between DUT and test range antenna | B3-2 |
| 3 | Quiet zone ripple DUT | B3-3 |
| 4 | Phase curvature | B3-4 |
| 5 | Polarization mismatch between DUT and receiving antenna | B3-5 |
| 6 | Mutual coupling between DUT and receiving antenna | B3-6 |
| 7 | Measurement equipment | B3-7, E |
| 8 | Impedance mismatch in receiving chain | B3-8 |
| 9 | RF leakage (DUT connector terminated and test range antenna connector cable terminated) | B3-9 |
| Stage 1-: Calibration measurement |  |  |
| 10 | Misalignment positioning system | B3-10 |
| 11 | Pointing error between reference antenna and test range antenna | B3-11 |
| 12 | Impedance mismatch in path to calibration antenna | B3-12 |
| 13 | Impedance mismatch in path to compact probe | B3-13 |
| 14 | Standing wave between reference antenna and test range antenna | B3-2 |
| 15 | Quiet zone ripple reference antenna | B3-3 |
| 16 | Phase curvature | B3-4 |
| 17 | Polarization mismatch between reference antenna and receiving antenna | B3-5 |
| 18 | Mutual coupling between reference antenna and receiving antenna | B3-6 |
| 19 | Measurement equipment | B3-7, E |
| 20 | Influence of reference antenna feed cable (flexing cables, adapters, attenuators, connector repeatability) | B3-14 |
| 21 | Mismatch of receiver chain | B3-15 |
| 22 | Insertion loss of receiver chain | B3-16 |
| 23 | Uncertainty of absolute gain of reference antenna | B3-17, E |
| 24 | RF leakage (SGH connector terminated and test range antenna connector cable terminated) | B3-9 |

###### 10.3.1.1.4.5 	Uncertainty assessment

Table 10.3.1.1.4.5-1: One Dimensional Compact Range Chamber uncertainty assessment for EIRP measurement

| EIRP uncertainty budget |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| UID | Uncertainty Source | Uncertainty value f ≦ 3GHz | Uncertainty value 3GHz < f ≦ 4.2 GHz | Distribution of the probability | Divisor based on distribution shape | Ci* | Standard uncertainty (σ) [dB] f ≦ 3GHz | Standard uncertainty (σ) [dB] 3GHz < f ≦ 4.2 GHz |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |
| 1 | Misalignment DUT and pointing error | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 |
| 2 | Standing wave between DUT and test range antenna | 0.18 | 0.18 | U-shaped | √2 | 1 | 0.13 | 0.13 |
| 3 | Quiet zone ripple DUT | 0.0325 | 0.0325 | Gaussian | 1 | 1 | 0.03 | 0.03 |
| 4 | Phase curvature | 0.01 | 0.01 | Gaussian | 1 | 1 | 0.01 | 0.01 |
| 5 | Polarization mismatch between DUT and receiving antenna | 0.05 | 0.05 | Rectangular | √3 | 1 | 0.03 | 0.03 |
| 6 | Mutual coupling between DUT and receiving antenna | 0 | 0 | Rectangular | √3 | 1 | 0 | 0 |
| 7 | Measurement equipment | 0.14 | 0.26 | Gaussian | 1 | 1 | 0.14 | 0.26 |
| 8 | Impedance mismatch in receiving chain | 0.0056 | 0.01 | U-shaped | √2 | 1 | 0 | 0.01 |
| 9 | RF leakage (DUT connector terminated and test range antenna connector cable terminated) | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |  |
| 10 | Misalignment positioning system | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 |
| 11 | Pointing error between reference antenna and test range antenna | 0 | 0 | Rectangular | √3 | 1 | 0 | 0 |
| 12 | Impedance mismatch in path to reference antenna | 0.05 | 0.05 | U-shaped | √2 | 1 | 0.04 | 0.04 |
| 13 | Impedance mismatch in path to compact probe | 0.03 | 0.03 | U-shaped | √2 | 1 | 0.02 | 0.02 |
| 14 | Standing wave between reference antenna and receiving antenna | 0.09 | 0.09 | U-shaped | √2 | 1 | 0.06 | 0.06 |
| 15 | Quiet zone ripple reference antenna | 0.178 | 0.178 | Gaussian | 1 | 1 | 0.18 | 0.18 |
| 16 | Phase curvature | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 |
| 17 | Polarization mismatch between reference antenna and receiving antenna | 0.05 | 0.05 | Rectangular | √3 | 1 | 0.03 | 0.03 |
| 18 | Mutual coupling between reference antenna and receiving antenna | 0 | 0 | Rectangular | √3 | 1 | 0 | 0 |
| 19 | Measurement equipment | 0.14 | 0.26 | Gaussian | 1 | 1 | 0.14 | 0.26 |
| 20 | Influence of the reference antenna feed cable (flexing cables, adapters, attenuators, connector repeatability) | 0.082 | 0.082 | Rectangular | √3 | 1 | 0.05 | 0.05 |
| 21 | Mismatch of receiver chain | 0.2 | 0.3 | U-shaped | √2 | 1 | 0.14 | 0.21 |
| 22 | Insertion loss of receiver chain | 0.18 | 0.18 | Rectangular | √3 | 1 | 0.1 | 0.1 |
| 23 | Uncertainty of absolute gain of reference antenna | 0.5 | 0.43 | Rectangular | √3 | 1 | 0.29 | 0.25 |
| 24 | RF leakage (SGH connector terminated and test range antenna connector cable terminated. | 0 | 0 | Gaussian | 1 | 1 | 0 | 0 |
| Combined standard uncertainty (1 σ) |  |  |  |  |  |  | 0.46 | 0.56 |
| Expanded uncertainty (2σ - confidence interval of 95 %) |  |  |  |  |  |  | 0.9 | 1.09 |

==== Unchanged sections omitted===========

#### 10.3.1.2	Maximum accepted test system uncertainty

Maximum test system uncertainties derivation methodology was described in subclause 10.1.1. The maximum accepted test system uncertainty values was derived based on test system specific values.

According to the methodology referred above, the common maximum accepted test system uncertainty values for the EIRP test can be derived from values captured in table 10.3.1.2-1, separately for each of the defined frequency ranges. The common maximum values are applicable for all test methods addressing EIRP test requirement. Based on the input values in table 10.3.2.2-1, the expanded uncertainty (1.96σ - confidence interval of 95 %) values are derived for two frequency ranges (i.e. f ≦ 3GHz and 3GHz < f ≦ 4.2 GHz).

Table 10.3.1.2-1: Test system specific measurement uncertainty values for the EIRP test

|  | Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |
| --- | --- | --- |
|  | f ≦ 3GHz | 3GHz < f ≦ 4.2 GHz |
| Indoor Anechoic Chamber | 0.87 | 1.06 |
| Compact Antenna Test Range | 1.00 | 1.17 |
| One Dimensional Compact Range Chamber | 0.9 | 1.09 |
| Near Field Test Range | 0.89 | 0.99 |
| Common maximum accepted test system uncertainty | 1.0 | 1.2 |

==== Unchanged sections omitted===========

## B.3	One Dimensional Compact Range

B3-1 	Misalignment DUT and pointing error

This contribution denotes uncertainty in DUT alignment and DUT pointing error. In this measurement the DUT is aligned to maximum, also allowing for a zero contribution for polarization mismatch uncertainty. By adjusting for maximums to align, this contribution can be a small contribution. The reference antenna´s phase centre and polarization purity changes slightly according to the frequency. Therefore, there should be some uncertainty reserved for this. To ensure that the pointing error is at a minimal, this contribution should be captured using the antenna pattern cut which is broadest (in the case of the DUT this would most likely be in the azimuth domain).

B3-2 	Standing wave between DUT and test range antenna

This value is extracting the uncertainty value and standard deviation of gain ripple coming from standing waves between DUT and test range antenna. This value can be captured by moving the DUT towards the test range antenna as the standing waves go in and out of phase causing a ripple in measured gain.

B3-3 	Quiet zone ripple DUT/reference antenna

This is the quiet zone (QZ) ripple experienced by the DUT/reference antenna during the measurement phase. The purpose of this component is to capture the contributions that the reflections from the walls, roof and floor that will add to the EIRP measurement. The sum of all these reflections from the walls, range reflector (if applicable), roof and floor will give the overall value for the QZ ripple. In other words, the uncertainty component from the wall will not be separated from the roof or the floor. The purpose of this uncertainty component is to capture the overall reflections from the chamber walls experienced by the DUT/reference antenna. To capture the full effect of the QZ ripple a distance of  1λ must be measured from each of the DUT/reference antenna physical aperture edges, i.e. total QZ distance =  physical aperture length + 2 λ, to ensure the full volume of the QZ is captured in the uncertainty measurement.

B3-4 	Phase curvature

This contribution originates from the finite far field measurement distance, which causes phase curvature across the antenna of AAS BS/reference antenna.

B3-5 	Polarization mismatch between DUT/reference antenna and receiving antenna

This contribution originates from the misaligned polarization between the DUT/reference antenna and the receiving antenna.

B3-6 	Mutual coupling between DUT/reference antenna and receiving antenna

This contribution originates from mutual coupling between the DUT/reference antenna and the receiving antenna. Mutual coupling degrades not just the antenna efficiency, i. e. the EIRP value, but it can alter the antenna’s radiation pattern as well. For compact range chamber, usually the spacing between the DUT/reference antenna and the receiving antennas is large enough so that the level of mutual coupling might be negligible.

B3-7 	Measurement equipment

This contribution originates from limited absolute level accuracy and non-linearity of the measurement equipment. The measurement equipment such as a BS simulator, spectrum analyser, or power meter measures the received signal level in EIRP tests either as an absolute level or as a relative level. The uncertainty value will be indicated in the manufacturer’s data sheet in logs.

B3-8 	Impedance mismatch in receiving chain

This contribution originates from multiple reflections between the receiving antenna and the measurement equipment. The multiple reflections can produce an overall reflection that depends not only upon the individual reflections of each part but their reflective interactions as well. The combination loss by the overall reflection can be higher or lower than individual loss by multiple reflections. The combination loss is called the mismatch error and leads to the measurement uncertainty.

B3-9 	RF leakage (DUT/SGH connector terminated and test range antenna connector cable terminated)

This contribution denotes noise leaking into connector and cable(s) between test range antenna and receiving equipment.  The contribution also includes the noise leakage between the connector and cable(s) between SGH/reference antenna and transmitting equipment.

B3-10 	Misalignment positioning system

This contribution originates from uncertainty in sliding position and turn table angle accuracy. If the calibration antenna is aligned to maximum this contribution can be considered negligible and therefore set to zero.

B3-11 	Pointing error between reference antenna and test range antenna

This contribution originates from the misalignment of the testing direction and the beam peak direction of the receiving antenna due to imperfect rotation operation. The pointing misalignment may happen in both azimuth and vertical directions and the effect of the misalignment depends highly on the beamwidth of the beam under test. The same level of misalignment results in a larger measurement error for a narrower beam.

B3-12 	Impedance mismatch in path to reference antenna

This contribution originates from multiple reflections between the reference antenna and the measurement equipment. After appropriate calibration, the measurement equipment may not introduce impedance mismatch error, but the error still happens between the reference antenna feed cable and the reference antenna.

B3-13 	Impedance mismatch in path to compact probe

This contribution originates from multiple reflections between the receiving antenna and the measurement equipment. After appropriate calibration, the measurement equipment may not introduce impedance mismatch error, but the error still happens between the receiving antenna feed cable and the receiving antenna.

B3-14 	Influence of reference antenna feed cable (flexing cables, adapters, attenuators, connector repeatability)

During the calibration phase this cable is used to feed the reference antenna and any influence it may have upon the measurements is captured. This is assessed by repeated measurements while flexing the cables and rotary joints. The largest difference between the results is recorded as the uncertainty.

B3-15 	Mismatch of receiver chain (i.e. between receiving antenna and measurement equipment)

This uncertainty is the residual uncertainty contribution coming from multiple reflections between the receiving antenna and the test receiver equipment. This value can be captured through measurement by measuring the S11 towards the receive antenna and also towards the test receiver. The mismatch between the antenna reflection and the receiver reflection can also be calculated. If the same cable is used for calibration stage, this can be considered systematic and negligible.

B3-16 	Insertion loss of receiver chain

This uncertainty is the residual uncertainty contribution coming from introducing an antenna at the end of the cable. If this cable does not change/move between the calibration and the measurement stage, the uncertainty is assumed to be systematic and negligible during the measurement stage. Alternatively, the insertion loss can also be calculated by taking the measurement of the cable where port 2 is the end of the cable connected to the DUT or reference antenna.

IL = -20log10|S21| dB

B3-17 	Uncertainty of absolute gain of reference antenna

This uncertainty consists of the uncertainty of the gain value associated with the gain value denoted from the antenna calibration.