3GPP TSG-RAN5 Meeting #104	R5-244441

Maastricht, Netherlands, 19th – 23rd Aug 2024

Agenda item:	5.4.20

Source:	Keysight Technologies

Title:	SE QoQZ MU for 40cm QZ Size

Document for:	Discussion and Endorsement

# Introduction

This contribution outlines our proposal for Spurious Emissions (SE) QoQZ MU for 40 cm QZ based on our recent measurements.

# Discussion

In [1], SE QoQZ MU measurement results were presented and reference values were proposed. An action point was created for system vendors to assess these MUs and to conclude the SE QoQZ MUs for 40 cm QZ.

Action ID; sWG; Action; Responsible; Relevant Tdoc; Deadline; Status. AP#103.21; RF; Define Spurious Emissions QoQZ MU for 40cm QZ and NTC; Anritsu, Keysight, R&S; R5-243002; RAN5#104; Open.

This contribution is presenting our measurement results for the SE QoQZ MU for the test frequencies inside, i.e., 23.45 GHz, 40.8 GHz, 49.0 GHz, and outside the FR2-1 range, i.e., 6 GHz, 12.75 GHz, 66 GHz, 80 GHz, and 87 GHz.

It should be noted that the test frequency for SE QoQZ of 87 GHz is currently not defined in Clause O.3.2 of [2], i.e.,

The frequencies to characterize the quality of the quiet zone shall be 6, 12.75, 23.45, 40.8, 49.0, 66, and 80 GHz. The quiet zone validation analysis is performed for each frequency individually.

It is proposed to add the 87 GHz frequency in Clause O.3.2 of [2].

Proposal 1: Add the 87 GHz frequency in Clause O.3.2 of TS 38.521-2.

The CR to implement Proposal 1 is in [4].

The measurements were performed according to the QoQZ for SE procedures in [2], specifically Clause O.3, with the re-positioning concept applied and the results are tabulated in Table 1.

Table 1: Analysis of TRP QoQZ std. deviations for spurious emissions

| Frequency [GHz] | TRP Standard Deviation [dB] | Number of Measurements | Angles  [o] | Angles  [o] | Angles Pol [o] |
| --- | --- | --- | --- | --- | --- |
| 6 | 0.42 | 14 | 0 | 0 | 0, 90 |
| 12.75 | 0.58 | 14 | 0 | 0 | 0, 90 |
| 23.45 | 0.50 | 14 | 0 | 0 | 0, 90 |
| 40.8 | 0.47 | 14 | 0 | 0 | 0, 90 |
| 43.5 | 0.47 | 14 | 0 | 0 | 0, 90 |
| 49 | 0.44 | 14 | 0 | 0 | 0, 90 |
| 66 | 0.38 | 14 | 0 | 0 | 0, 90 |
| 80 | 0.54 | 14 | 0 | 0 | 0, 90 |
| 87 | 0.58 | 14 | 0 | 0 | 0, 90 |

In [1], the following proposed QoQZ MUs for MTSU and TT definitions were provided

Table 6 [from [1]]: Endorsed/Proposed QoQZ MU for Spurious measurement for Stage 1 and Stage 2 Condition; QoQZ MU [dB]. 6 GHz to 12.75 GHz; 12.75 GHz to 23.45 GHz; 23.45 GHz to 40.8 GHz; 40.8 GHz to 66 GHz; 66 GHz to 87 GHz. QZ size ≤ 30 cm; 0.7; 0.6; 0.6; 0.6; 0.6. QZ size = 40 cm; 1.0; 0.6; 0.6; 0.6; 0.6. Note 1: Black text indicates endorsed values, and green text indicates proposed values in this paper..

It can be observed that the SE QoQZ measurements for 40 cm provided in Table 1 are within the proposed limits from [1] for 40 cm QZ and even the previously agreed 30 cm QZ limits.

It is proposed to define the 40 GHz SE QoQZ uncertainty values as shown in Table 2, i.e., to set the SE QoQZ MUs for 40 cm to those for 30 cm QZ. Alternatively, we could accept the proposal from [1] with the higher value at 6 GHz while applying Note 2 of Table B.2.2.3-1, i.e.,

NOTE 2: For Uncertainty assessment in tables across clause B and MTSU calculation, QoQZ uncertainty value for QZ ≤ 30cm is used also for 40cm for FR2a through FR2c. For uncertainty assessment in tables across clause B and MTSU calculation for FR2d, QoQZ uncertainty value for QZ size of 40cm is used also for QZ size ≤ 30cm.

to the SE QoQZ values for 40 cm QZ as well.

Table 2: SE QoQZ MUs

| Condition | SE QoQZ MU [dB] |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  | 6 GHz to 12.75 GHz | 12.75 GHz to 23.45 GHz | 23.45 GHz to 40.8 GHz | 40.8 GHz to 66 GHz | 66 GHz to 80 GHz (Note 1) | 80 GHz to 87 GHz (Note 1) |
| QZ size ≤ 30 cm | 0.7 | 0.6 | 0.6 | 0.6 | 0.6 |  |
| 30 cm < QZ size ≤ 40 cm | 0.7 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 |
| Note 1: It is suggested to split the range from 66 GHz to 87 GHz into two ranges: 66 GHz to 80 GHz and 80 GHz to 87 GHz |  |  |  |  |  |  |

Proposal 2: Define the 40 cm SE QoQZ uncertainty values as shown in Table 2.

As stated in Note 1 of Table 2, it is proposed to define the SE QoQZ MUs for two separate ranges from 66 GHz to 87 GHz, i.e., 66 GHz to 80 GHz and 80 GHz to 87 GHz, in [3], specifically Tables B.2.2.3-1 and B.2.2.19-1, to align the frequency ranges of SE QoQZ MU with the ranges MTSU and TT were defined, Clauses F.1.2, F.1.3, F.3.2, and F.3.3 of [2].

Proposal 3: Define the SE QoQZ MUs for two separate ranges from 66 GHz to 87 GHz, i.e., 66 GHz to 80 GHz and 80 GHz to 87 GHz, in TR 38.903, specifically Tables B.2.2.3-1 and B.2.2.19-1.

The CR to implement Proposal 2 and Proposal 3 is in [5].

# Conclusion

The following observations and proposals were made in this contribution.

Proposal 1: Add the 87 GHz frequency in Clause O.3.2 of TS 38.521-2.

Proposal 2: Define the 40 GHz SE QoQZ uncertainty values as shown in Table 2.

Proposal 3: Define the SE QoQZ MUs for two separate ranges from 66 GHz to 87 GHz, i.e., 66 GHz to 80 GHz and 80 GHz to 87 GHz, in TR 38.903, specifically Tables B.2.2.3-1 and B.2.2.19-1.

# References

- R5-243002, QoQZ for QZ size 40 cm, Anritsu, 3GPP TSG-RAN5 Meeting #103, May 2024

- TS 38.521-2, User Equipment (UE) conformance specification; Radio transmission and reception; Part 2: Range 2 Standalone

- TR 38.903, Derivation of test tolerances and measurement uncertainty for User Equipment (UE) conformance test cases

- R5-244442, Correction of SE QoQZ Validation Frequencies, Keysight Technologies, 3GPP TSG-RAN5 Meeting #104, August 2024

- R5-244443, Definition of 40cm SE QoQZ MU, Keysight Technologies, 3GPP TSG-RAN5 Meeting #104, August 2024