3GPP TSG-RAN WG5 #104		R5-245251
Maastricht, Netherlands, 19th - 23rd August 2024

Agenda item:	5.4.18

Source:	Rohde & Schwarz

Title:	On the MU for RRM FR2 PC1

Document for:	Endorsement

# 1.	Introduction

At RAN5#97, it was agreed to leverage RRM FR2 PC1 MU for IFF from RF. At the time, some parameters were in square brackets. Since RAN5#97, RF has completed the definition of the MU values for PC1.

This paper presents the results of our MU assessment and proposes to update RRM PC1 MU tables based on RF decisions.

# 2.	Discussion

The uncertainty assessment for EIS measurement for PC1 UEs for f ≤ 40.8 GHz for RF RefSens test case has been completed in B.19.2 of TR 38.903 [1].

Since DL AWGN absolute power or wanted DL signal absolute power are based on RF RefSens, the square brackets can be removed for RRM IFF, PC1.

Observation 1: In Table E.3.1.3-3 of TR 38.903 [3], values of “Amplifier uncertainties” and “DL AWGN absolute power expanded uncertainty” are in square brackets.

Table E.3.1.3-3: Uncertainty assessment for DL AWGN absolute power or wanted DL signal absolute power (f=23.45GHz, 32.125GHz, Quiet Zone size ≤ 30 cm), PC1

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment | 0.02 | Normal | 2.00 | 0.01 |
| 2 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 3 | Quality of Quiet Zone (NOTE 4) | 0.6 | Actual | 1.00 | 0.6 |
| 4 | Mismatch | 1.30 | Actual | 1.00 | 1.30 |
| 5 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | gNB uncertainty on absolute level | 2.9 | Normal | 2.00 | 1.45 |
| 7 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 8 | Amplifier uncertainties | [2.1] | Normal | 2.00 | [1.05] |
| 9 | Random uncertainty | 0.50 | Normal | 2.00 | 0.25 |
| 10 | Influence of the XPD | 0.01 | U-shaped | 1.41 | 0.00 |
| 11 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 13 | Multiple measurement antenna uncertainty (NOTE 3) | 0.15 | Actual | 1.00 | 0.15 |
| 14 | DUT repositioning | 0.35 | Rectangular | 1.73 | 0.20 |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 15 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 16 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 17 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 18 | Uncertainty of the Network Analyzer | 1.50 | Normal | 2.00 | 0.75 |
| 19 | Uncertainty of the absolute gain of the calibration antenna | 0.60 | Normal | 2.00 | 0.30 |
| 20 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.01 | Rectangular | 1.73 | 0.00 |
| 21 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 22 | Quality of quiet zone for calibration process (NOTE 4) | 0.4 | Actual | 1.00 | 0.4 |
| 23 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 24 | Influence of the calibration antenna feed cable | 0.14 | Normal | 2.00 | 0.07 |
| 25 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
|  | Systematic uncertainties (NOTE 2) |  |  |  | Value |
| 26 | Systematic error related to beam peak search |  |  |  | 0.7 |
| Total measurement uncertainty |  |  |  |  | Value |
| DL AWGN absolute power expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | [5.58] |
| NOTE 1: The analysis was done only for the case of operating in-band, non-CA. NOTE 2: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 3: Applies to the system which has a structure of mechanical feed antenna positioning. NOTE 4: Value based on procedure defined in Annex D.2 of TR 38.810 [13] for Quiet Zone size less or equal to 30 cm. NOTE 5: The values in this table have been derived for DL powers above and equal to REFSENS. The values might need to be revisited for power levels below REFSENS |  |  |  |  |  |

Observation 2: In RF Table B.19.2-3 of TR 38.903 [2], values were added for up to 32.125GHz, which was used as reference in [1]

Observation 3: In RF Table B.19.2-3 of TR 38.903 [3], all values have been agreed, up to 40.8 GHz

Table B.19.2-3: Uncertainty assessment for EIS measurement (f=23.45GHz, 32.125GHz, 40.8GHz, Quiet Zone size ≤ 30 cm) for PC1 and PC5 UEs and normal temperature condition

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment | 0.02 | Normal | 2.00 | 0.01 |
| 2 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 3 | Quality of Quiet Zone (NOTE 7) | 0.6 | Actual | 1.00 | 0.6 |
| 4 | Mismatch | 1.30 | Actual | 1.00 | 1.30 |
| 5 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | gNB uncertainty on absolute level | 2.9 | Normal | 2.00 | 1.45 |
| 7 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 8 | Amplifier uncertainties | 2.1 | Normal | 2.00 | 1.05 |
| 9 | Random uncertainty | 0.50 | Normal | 2.00 | 0.25 |
| 10 | Influence of the XPD | 0.01 | U-shaped | 1.41 | 0.00 |
| 11 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 13 | Multiple measurement antenna uncertainty (NOTE 6) | 0.15 | Actual | 1.00 | 0.15 |
| 14 | DUT repositioning | 0.00 (NOTE 4) 0.35 (NOTE 5) | Rectangular | 1.73 | 0.00 (NOTE 4) 0.20 (NOTE 5) |
| 15 | Influence of spherical coverage grid (NOTE 4) | 0.13 | Actual | 1 | 0.13 |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 16 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 17 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 18 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 19 | Uncertainty of the Network Analyzer | 1.50 | Normal | 2.00 | 0.75 |
| 20 | Uncertainty of the absolute gain of the calibration antenna | 0.60 | Normal | 2.00 | 0.30 |
| 21 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.01 | Rectangular | 1.73 | 0.00 |
| 22 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 23 | Quality of quiet zone for calibration process (NOTE 7) | 0.4 | Actual | 1.00 | 0.4 |
| 24 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 25 | Influence of the calibration antenna feed cable | 0.14 | Normal | 2.00 | 0.07 |
| 26 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
|  | Systematic uncertainties (NOTE 3) |  |  |  | Value |
| 27 | Systematic error related to beam peak search (NOTE 5) |  |  |  | 0.7 |
| 28 | Systematic error related to EIS spherical coverage (NOTE 4) |  |  |  | DL power step size, 0.2 |
| Total measurement uncertainty |  |  |  |  | Value |
| EIS Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 5.58 |
| EIS Spherical coverage Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 5.07 |
| NOTE 1: The analysis was done only for the case of operating at max output power, in-band, non-CA. NOTE 2: Void. NOTE 3: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 4: This contributor shall only be considered for spherical EIS measurements. NOTE 5: This contributor shall only be considered for EIS measurements. NOTE 6: Applies to the system which has a structure of mechanical feed antenna positioning. NOTE 7: Value based on procedure defined in clause D.2 of TR 38.810 for Quiet Zone size less or equal to 30 cm. |  |  |  |  |  |

Proposal 1: Remove the square brackets in Table E.3.1.3-3 [3] as RF has already agreed on the related values.

Table E.3.1.3-3: Uncertainty assessment for DL AWGN absolute power or wanted DL signal absolute power (f=23.45GHz, 32.125GHz, 40.8GHz, Quiet Zone size ≤ 30 cm), PC1

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment | 0.02 | Normal | 2.00 | 0.01 |
| 2 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 3 | Quality of Quiet Zone (NOTE 4) | 0.6 | Actual | 1.00 | 0.6 |
| 4 | Mismatch | 1.30 | Actual | 1.00 | 1.30 |
| 5 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | gNB uncertainty on absolute level | 2.9 | Normal | 2.00 | 1.45 |
| 7 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 8 | Amplifier uncertainties | 2.1 | Normal | 2.00 | 1.05 |
| 9 | Random uncertainty | 0.50 | Normal | 2.00 | 0.25 |
| 10 | Influence of the XPD | 0.01 | U-shaped | 1.41 | 0.00 |
| 11 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 13 | Multiple measurement antenna uncertainty (NOTE 3) | 0.15 | Actual | 1.00 | 0.15 |
| 14 | DUT repositioning | 0.35 | Rectangular | 1.73 | 0.20 |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 15 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 16 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 17 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 18 | Uncertainty of the Network Analyzer | 1.50 | Normal | 2.00 | 0.75 |
| 19 | Uncertainty of the absolute gain of the calibration antenna | 0.60 | Normal | 2.00 | 0.30 |
| 20 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.01 | Rectangular | 1.73 | 0.00 |
| 21 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 22 | Quality of quiet zone for calibration process (NOTE 4) | 0.4 | Actual | 1.00 | 0.4 |
| 23 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 24 | Influence of the calibration antenna feed cable | 0.14 | Normal | 2.00 | 0.07 |
| 25 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
|  | Systematic uncertainties (NOTE 2) |  |  |  | Value |
| 26 | Systematic error related to beam peak search |  |  |  | 0.7 |
| Total measurement uncertainty |  |  |  |  | Value |
| DL AWGN absolute power expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 5.58 |
| NOTE 1: The analysis was done only for the case of operating in-band, non-CA. NOTE 2: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 3: Applies to the system which has a structure of mechanical feed antenna positioning. NOTE 4: Value based on procedure defined in Annex D.2 of TR 38.810 [13] for Quiet Zone size less or equal to 30 cm. NOTE 5: The values in this table have been derived for DL powers above and equal to REFSENS. The values might need to be revisited for power levels below REFSENS |  |  |  |  |  |

Regarding the 2AoA TCs, the relative offsets between active probes has been specified in Table A.9-3 38.533 [6]. This enables the reuse of existing enhanced IFF test systems for PC1 and to complete the MU assessment for PC1.

Observation 4: The AoA offsets for UE power class 1 have already been defined in Annex A.9-3 of TS 38.533 [4]

Table A.9.3-1: Set of relative angular offsets between active probes for each power class

| UE Power class | Relative angular offset between active probes |
| --- | --- |
| 1 | 30°, 60°, 90° and 120° |
| 2 | FFS |
| 3 | 30°, 60°, 90°, 120° and 150° |
| 4 | FFS |

Proposal 2: Remove the square brackets in Table E.3.1.4-3 [3] for RRM Enhanced IFF and apply the updates as below.

Table E.3.1.4-3: Uncertainty assessment for fDL AWGN absolute power or wanted DL signal absolute power (f=23.45GHz, 32.125GHz, 40.8GHz, Quiet Zone size ≤ 30 cm) for PC1

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment | 0.02 | Normal | 2.00 | 0.020.01 |
| 2 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 3 | Quality of Quiet Zone (NOTE 4) | 0.7 | Actual | 1.00 | 0.7 |
| 4 | Mismatch | 1.30 | Actual | 1.00 | 1.30 |
| 5 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | gNB uncertainty on absolute level | 2.9 | Normal | 2.00 | 1.45 |
| 7 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 8 | Amplifier uncertainties | 2.1 | Normal | 2.00 | 1.05 |
| 9 | Random uncertainty | 0.50 | Normal | 2.00 | 0.25 |
| 10 | Influence of the XPD | 0.01 | U-shaped | 1.41 | 0.00 |
| 11 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 13 | Multiple measurement antenna uncertainty (NOTE 3) | 0.15 | Actual | 1.00 | 0.15 |
| 14 | DUT repositioning | 0.35 | Rectangular | 1.73 | 0.20 |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 17 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 18 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 19 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 20 | Uncertainty of the Network Analyzer | 1.50 | Normal | 2.00 | 0.75 |
| 21 | Uncertainty of the absolute gain of the calibration antenna | 0.60 | Normal | 2.00 | 0.30 |
| 22 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.01 | Rectangular | 1.73 | 0.00 |
| 23 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 24 | Quality of quiet zone for calibration process (NOTE 4) | 0.4 | Actual | 1.00 | 0.4 |
| 25 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 26 | Influence of the calibration antenna feed cable | 0.14 | Normal | 2.00 | 0.07 |
| 27 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
|  | Systematic uncertainties (NOTE 2) |  |  |  | Value |
| 28 | Systematic error related to beam peak search |  |  |  | 0.7 |
| Total measurement uncertainty |  |  |  |  | Value |
| DL AWGN absolute power expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 5.655.63 |
| NOTE 1: The analysis was done only for the case of operating in-band, non-CA. NOTE 2: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 3: Applies to the system which has a structure of mechanical feed antenna positioning. NOTE 4: Value based on procedure defined in Annex M of TR 38.521-2 [7] for Quiet Zone size less or equal to 30 cm. NOTE 5: The values in this table have been derived for DL powers above and equal to REFSENS. The values might need to be revisited for power levels below REFSENS |  |  |  |  |  |

For the MTSU, we propose to reuse the value of 5.65 dB, since the difference to 5.63 dB is not significant. This allows for a simpler test requirement definition and MTSU definition in TS 38.533.

Proposal 3: Apply a value of 5.65 dB for the MU threshold for DL AWGN absolute power for RRM FR2 for PC1.

Table E.3.1-1: MU threshold for DL AWGN absolute power for RRM FR2

| Power Class | Frequency | MBW | Power | Threshold MU value (NOTE 1) |
| --- | --- | --- | --- | --- |
| PC3 | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | As configured in the test case | 5.65 dB2 |
|  | 32.125GHz < f <= 40.8GHz |  |  | 5.65 dB2 |
| PC1 | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | As configured in the test case | 5.65 dB2 |
|  | 32.125GHz < f <= 40.8GHz |  |  | 5.65 dB2 |
| NOTE 1: Total Expanded MU for IFF for Quiet Zone size ≤ 30cm in Table E.3.1.3-2 for PC3 UEs and Table E.3.1.4-3 for PC1 UEs NOTE 2: If the TT analysis for a specific test case based on this MU value results in an unsolvable conflict, making the test case untestable, even after the alternative solutions listed in clause A.4 have been considered for the test case in TS 38.133 [6] Annex A, the TT analysis shall be repeated using a lower MU value, taking into account the lower values defined in this clause. The test case will be applicable for the subset of the test systems meeting this reduced MU Threshold. |  |  |  |  |

The accompanying CR for the update of TR 38.903 [3] is contained in Ref. [5].

The accompanying CR for the update of TS 38.533 [4] is contained in Ref. [6].

# 3.	Conclusion

Observation 1: In Table E.3.1.3-3 of TR 38.903 [3], values of “Amplifier uncertainties” and “DL AWGN absolute power expanded uncertainty” are in square brackets.

Observation 2: In RF Table B.19.2-3 of TR 38.903 [2], values were added for up to 32.125GHz, which was used as reference in [1]

Observation 3: In RF Table B.19.2-3 of TR 38.903 [3], all values have been agreed, up to 40.8 GHz

Observation 4: The AoA offsets for UE power class 1 have already been defined in Annex A.9-3 of TS 38.533 [4]

Proposal 1: Remove the square brackets in Table E.3.1.3-3 [3] as RF has already agreed on the related values.

Proposal 2: Remove the square brackets in Table E.3.1.4-3 [3] for RRM Enhanced IFF.

Proposal 3: Apply a value of 5.65 dB for the MU threshold for DL AWGN absolute power for RRM FR2 for PC1.

# 4.	References

- R5-227821, “Update of MU for PC1 RRM”, 3GPP TSG-RAN WG5 Meeting #97

- TR 38.903 “Derivation of test tolerances and measurement uncertainty for User Equipment (UE) conformance test cases”, V16.13.0 (2022-09)

- TR 38.903 “Derivation of test tolerances and measurement uncertainty for User Equipment (UE) conformance test cases”, V18.3.0 (2024-06)

- TS 38.533 “User Equipment (UE) conformance specification; Radio Resource Management (RRM)”, V18.3.0 (2024-06)

- R5-245253, “MU update for RRM FR2 PC1”, 3GPP TSG-RAN WG5 Meeting #104

- R5-245252, “Correction of MU values for RRM FR2 PC1”, 3GPP TSG-RAN WG5 Meeting #104