3GPP TSG-RAN WG4 Meeting #101-e	R4-2119558
Electronic Meeting, 01 – 12 November, 2021

Agenda item:	8.1.3.3

Source:	CAICT

Title:	FR1 MIMO OTA channel model validation results

Document for:	Discussion

1	Introduction

The FR1 MIMO OTA lab alignment activity has started, and channel mode validation is an important step [1,2].

CAICT has shared some initial FR1 channel model validation results in 2019 and 2020 [3,4]. In this paper, we present some new FR1 channel model validation results based on the latest adopted FR1 CDL-C UMa channel model. Specifically, measurement results of Temporal correlation and Spatial correlation are shown. Other channel model validation measurement results will be shared later.

2	Discussion

  - 2.1	General

Based on the parameters in TS 38.151 V0.5.0, we did some FR1 channel model validation measurements in a 16-probe MIMO OTA chamber. The key equipment utilized in the test system are shown as below:

- ETS-Lindgren 16-probe MIMO OTA chamber

- Spirent Vertex channel emulator

The detailed configuration is aligned with TS 38.151 V0.5.0. General parameters of the measurements are listed in Table 1.

Table 1. General parameters

| Channel model | CDL-C UMa |
| --- | --- |
| NR FR1 band (Test frequency) | n41 (2450MHz), n78 (3600MHz) |
| Measurement | Doppler/Temporal correlation Spatial correlation |
| Beam | Beam 1, Beam 2, Combined beams |
| Polarization | X2V, H11, H12 |

The reference curves in the following figures are adopted from [5,6].

2.2  FR1 channel model validation results

    - 2.2.1	Doppler/Temporal correlation

Measurement results:

(a) Beam 1, H11                                  (b) Beam 2, H11

(c) Beam 1, H12                                 (d) Beam 2, H12

(e) Beam 1, X2V                                  (f) Beam 2, X2V

Figure 1. Temporal correlation measurement results and simulated reference for CDL-C UMa, fc ≤ 2.5 GHz.

Table 2(a). Temporal correlation verification results for CDL-C UMa with Beam 1 and ≤ 2.5 GHz

| CDL-C UMa, Beam 1, X2V, 2450 MHz |  |  |  |
| --- | --- | --- | --- |
| lambdas | Reference TCF | Measured TCF | delta |
| 0 | 1 | 1 | 0 |
| 0.1 | 0.986 | 0.9858 | 0.0002 |
| 0.2 | 0.945 | 0.9447 | 0.0003 |
| 0.3 | 0.882 | 0.8808 | 0.0012 |
| 0.4 | 0.801 | 0.7997 | 0.0013 |
| 0.5 | 0.709 | 0.7083 | 0.0007 |
| 0.6 | 0.613 | 0.6131 | -0.0001 |
| 0.7 | 0.518 | 0.52 | -0.002 |
| 0.8 | 0.43 | 0.4339 | -0.0039 |
| 0.9 | 0.353 | 0.3581 | -0.0051 |
| 1 | 0.289 | 0.2947 | -0.0057 |
| 1.1 | 0.24 | 0.2446 | -0.0046 |
| 1.2 | 0.204 | 0.2074 | -0.0034 |
| 1.3 | 0.181 | 0.1817 | -0.0007 |
| 1.4 | 0.167 | 0.1654 | 0.0016 |
| 1.5 | 0.159 | 0.1557 | 0.0033 |
| 1.6 | 0.155 | 0.1501 | 0.0049 |
| 1.7 | 0.153 | 0.1459 | 0.0071 |
| 1.8 | 0.15 | 0.1413 | 0.0087 |
| 1.9 | 0.144 | 0.1347 | 0.0093 |
| 2 | 0.135 | 0.1256 | 0.0094 |
| 2.1 | 0.121 | 0.1138 | 0.0072 |
| 2.2 | 0.105 | 0.0998 | 0.0052 |
| 2.3 | 0.085 | 0.0844 | 0.0006 |
| 2.4 | 0.065 | 0.0689 | -0.0039 |
| 2.5 | 0.048 | 0.0545 | -0.0065 |
| 2.6 | 0.039 | 0.0429 | -0.0039 |
| 2.7 | 0.038 | 0.0349 | 0.0031 |
| 2.8 | 0.042 | 0.0306 | 0.0114 |
| 2.9 | 0.043 | 0.0297 | 0.0133 |
| 3 | 0.041 | 0.0327 | 0.0083 |
| 3.1 | 0.037 | 0.0403 | -0.0033 |
| 3.2 | 0.036 | 0.0515 | -0.0155 |
| 3.3 | 0.044 | 0.0642 | -0.0202 |
| 3.4 | 0.056 | 0.0758 | -0.0198 |
| 3.5 | 0.068 | 0.084 | -0.016 |
| 3.6 | 0.075 | 0.087 | -0.012 |
| 3.7 | 0.076 | 0.0837 | -0.0077 |
| 3.8 | 0.068 | 0.0734 | -0.0054 |
| 3.9 | 0.051 | 0.0567 | -0.0057 |
| 4 | 0.027 | 0.0352 | -0.0082 |
| 4.1 | 0.007 | 0.0167 | -0.0097 |
| 4.2 | 0.036 | 0.029 | 0.007 |
| 4.3 | 0.067 | 0.0537 | 0.0133 |
| 4.4 | 0.093 | 0.076 | 0.017 |
| 4.5 | 0.111 | 0.0926 | 0.0184 |
| 4.6 | 0.119 | 0.1019 | 0.0171 |
| 4.7 | 0.116 | 0.1039 | 0.0121 |
| 4.8 | 0.101 | 0.0996 | 0.0014 |
| 4.9 | 0.078 | 0.0909 | -0.0129 |
| 5 | 0.051 | 0.0809 | -0.0299 |

Table 2(b). Temporal correlation verification results for CDL-C UMa with Beam 2 and ≤ 2.5 GHz

| CDL-C UMa, Beam 2, X2V, 2450 MHz |  |  |  |
| --- | --- | --- | --- |
| lambdas | Reference TCF | Measured TCF | delta |
| 0 | 1 | 1 | 0 |
| 0.1 | 0.974 | 0.9763 | -0.0023 |
| 0.2 | 0.907 | 0.9165 | -0.0095 |
| 0.3 | 0.832 | 0.8464 | -0.0144 |
| 0.4 | 0.776 | 0.7892 | -0.0132 |
| 0.5 | 0.738 | 0.7461 | -0.0081 |
| 0.6 | 0.695 | 0.6982 | -0.0032 |
| 0.7 | 0.623 | 0.6283 | -0.0053 |
| 0.8 | 0.525 | 0.5373 | -0.0123 |
| 0.9 | 0.426 | 0.4457 | -0.0197 |
| 1 | 0.36 | 0.38 | -0.02 |
| 1.1 | 0.335 | 0.3476 | -0.0126 |
| 1.2 | 0.32 | 0.3272 | -0.0072 |
| 1.3 | 0.287 | 0.2956 | -0.0086 |
| 1.4 | 0.233 | 0.2478 | -0.0148 |
| 1.5 | 0.176 | 0.1958 | -0.0198 |
| 1.6 | 0.141 | 0.159 | -0.018 |
| 1.7 | 0.135 | 0.1465 | -0.0115 |
| 1.8 | 0.137 | 0.1459 | -0.0089 |
| 1.9 | 0.132 | 0.1422 | -0.0102 |
| 2 | 0.117 | 0.1296 | -0.0126 |
| 2.1 | 0.097 | 0.1091 | -0.0121 |
| 2.2 | 0.076 | 0.086 | -0.01 |
| 2.3 | 0.062 | 0.0721 | -0.0101 |
| 2.4 | 0.071 | 0.0775 | -0.0065 |
| 2.5 | 0.09 | 0.091 | -0.001 |
| 2.6 | 0.099 | 0.095 | 0.004 |
| 2.7 | 0.088 | 0.0819 | 0.0061 |
| 2.8 | 0.058 | 0.057 | 0.001 |
| 2.9 | 0.037 | 0.0481 | -0.0111 |
| 3 | 0.067 | 0.074 | -0.007 |
| 3.1 | 0.103 | 0.102 | 0.001 |
| 3.2 | 0.12 | 0.1141 | 0.0059 |
| 3.3 | 0.115 | 0.1079 | 0.0071 |
| 3.4 | 0.097 | 0.0908 | 0.0062 |
| 3.5 | 0.082 | 0.0764 | 0.0056 |
| 3.6 | 0.083 | 0.0737 | 0.0093 |
| 3.7 | 0.09 | 0.075 | 0.015 |
| 3.8 | 0.089 | 0.07 | 0.019 |
| 3.9 | 0.079 | 0.0582 | 0.0208 |
| 4 | 0.068 | 0.0482 | 0.0198 |
| 4.1 | 0.063 | 0.0482 | 0.0148 |
| 4.2 | 0.062 | 0.0522 | 0.0098 |
| 4.3 | 0.057 | 0.0527 | 0.0043 |
| 4.4 | 0.052 | 0.0512 | 0.0008 |
| 4.5 | 0.055 | 0.0545 | 0.0005 |
| 4.6 | 0.063 | 0.0632 | -0.0002 |
| 4.7 | 0.066 | 0.0699 | -0.0039 |
| 4.8 | 0.058 | 0.0684 | -0.0104 |
| 4.9 | 0.047 | 0.0574 | -0.0104 |
| 5 | 0.048 | 0.0393 | 0.0087 |

(a) Beam 1, H11                                  (b) Beam 2, H11

(c) Beam 1, X2V                                  (d) Beam 2, X2V

Figure 2. Temporal correlation measurement results and simulated reference for CDL-C UMa, fc > 2.5 GHz.

Table 3(a). Temporal correlation verification results for CDL-C UMa with Beam 1 and > 2.5 GHz

| CDL-C UMa, Beam 1, X2V, 3600 MHz |  |  |  |
| --- | --- | --- | --- |
| lambdas | Reference TCF | Measured TCF | delta |
| 0 | 1 | 1 | 0 |
| 0.1 | 0.985 | 0.9833 | 0.0017 |
| 0.2 | 0.942 | 0.9355 | 0.0065 |
| 0.3 | 0.874 | 0.8612 | 0.0128 |
| 0.4 | 0.787 | 0.7675 | 0.0195 |
| 0.5 | 0.689 | 0.6622 | 0.0268 |
| 0.6 | 0.586 | 0.5535 | 0.0325 |
| 0.7 | 0.486 | 0.4482 | 0.0378 |
| 0.8 | 0.394 | 0.3519 | 0.0421 |
| 0.9 | 0.315 | 0.2686 | 0.0464 |
| 1 | 0.252 | 0.2008 | 0.0512 |
| 1.1 | 0.206 | 0.1497 | 0.0563 |
| 1.2 | 0.174 | 0.115 | 0.059 |
| 1.3 | 0.154 | 0.0957 | 0.0583 |
| 1.4 | 0.143 | 0.0895 | 0.0535 |
| 1.5 | 0.137 | 0.0926 | 0.0444 |
| 1.6 | 0.135 | 0.1002 | 0.0348 |
| 1.7 | 0.134 | 0.1077 | 0.0263 |
| 1.8 | 0.134 | 0.1116 | 0.0224 |
| 1.9 | 0.13 | 0.1096 | 0.0204 |
| 2 | 0.122 | 0.1008 | 0.0212 |
| 2.1 | 0.109 | 0.0853 | 0.0237 |
| 2.2 | 0.09 | 0.0644 | 0.0256 |
| 2.3 | 0.069 | 0.0406 | 0.0284 |
| 2.4 | 0.047 | 0.0211 | 0.0259 |
| 2.5 | 0.031 | 0.0277 | 0.0033 |
| 2.6 | 0.033 | 0.0483 | -0.0153 |
| 2.7 | 0.046 | 0.0664 | -0.0204 |
| 2.8 | 0.057 | 0.078 | -0.021 |
| 2.9 | 0.062 | 0.0816 | -0.0196 |
| 3 | 0.06 | 0.0773 | -0.0173 |
| 3.1 | 0.05 | 0.0664 | -0.0164 |
| 3.2 | 0.036 | 0.0519 | -0.0159 |
| 3.3 | 0.019 | 0.0384 | -0.0194 |
| 3.4 | 0.01 | 0.0315 | -0.0215 |
| 3.5 | 0.019 | 0.0328 | -0.0138 |
| 3.6 | 0.029 | 0.0365 | -0.0075 |
| 3.7 | 0.034 | 0.0393 | -0.0053 |
| 3.8 | 0.036 | 0.044 | -0.008 |
| 3.9 | 0.044 | 0.0573 | -0.0133 |
| 4 | 0.062 | 0.0817 | -0.0197 |
| 4.1 | 0.09 | 0.114 | -0.024 |
| 4.2 | 0.123 | 0.1498 | -0.0268 |
| 4.3 | 0.155 | 0.1846 | -0.0296 |
| 4.4 | 0.182 | 0.2142 | -0.0322 |
| 4.5 | 0.2 | 0.2353 | -0.0353 |
| 4.6 | 0.207 | 0.2455 | -0.0385 |
| 4.7 | 0.2 | 0.244 | -0.044 |
| 4.8 | 0.18 | 0.2311 | -0.0511 |
| 4.9 | 0.149 | 0.2082 | -0.0592 |
| 5 | 0.11 | 0.1775 | -0.0675 |

Table 3(b). Temporal correlation verification results for CDL-C UMa with Beam 2 and > 2.5 GHz

| CDL-C UMa, Beam 2, X2V, 3600 MHz |  |  |  |
| --- | --- | --- | --- |
| lambdas | Reference TCF | Measured TCF | delta |
| 0 | 1 | 1 | 0 |
| 0.1 | 0.973 | 0.9745 | -0.0015 |
| 0.2 | 0.904 | 0.9096 | -0.0056 |
| 0.3 | 0.825 | 0.8308 | -0.0058 |
| 0.4 | 0.765 | 0.7617 | 0.0033 |
| 0.5 | 0.723 | 0.7056 | 0.0174 |
| 0.6 | 0.675 | 0.6462 | 0.0288 |
| 0.7 | 0.599 | 0.5681 | 0.0309 |
| 0.8 | 0.496 | 0.473 | 0.023 |
| 0.9 | 0.391 | 0.3813 | 0.0097 |
| 1 | 0.319 | 0.3178 | 0.0012 |
| 1.1 | 0.29 | 0.2858 | 0.0042 |
| 1.2 | 0.273 | 0.2617 | 0.0113 |
| 1.3 | 0.239 | 0.225 | 0.014 |
| 1.4 | 0.185 | 0.1749 | 0.0101 |
| 1.5 | 0.129 | 0.1257 | 0.0033 |
| 1.6 | 0.096 | 0.0942 | 0.0018 |
| 1.7 | 0.092 | 0.0823 | 0.0097 |
| 1.8 | 0.095 | 0.0785 | 0.0165 |
| 1.9 | 0.093 | 0.0797 | 0.0133 |
| 2 | 0.089 | 0.0869 | 0.0021 |
| 2.1 | 0.086 | 0.0925 | -0.0065 |
| 2.2 | 0.076 | 0.089 | -0.013 |
| 2.3 | 0.064 | 0.08 | -0.016 |
| 2.4 | 0.067 | 0.0793 | -0.0123 |
| 2.5 | 0.088 | 0.0907 | -0.0027 |
| 2.6 | 0.103 | 0.099 | 0.004 |
| 2.7 | 0.099 | 0.0911 | 0.0079 |
| 2.8 | 0.073 | 0.0653 | 0.0077 |
| 2.9 | 0.038 | 0.0358 | 0.0022 |
| 3 | 0.045 | 0.0446 | 0.0004 |
| 3.1 | 0.08 | 0.0725 | 0.0075 |
| 3.2 | 0.1 | 0.0873 | 0.0127 |
| 3.3 | 0.099 | 0.0841 | 0.0149 |
| 3.4 | 0.081 | 0.0686 | 0.0124 |
| 3.5 | 0.061 | 0.052 | 0.009 |
| 3.6 | 0.053 | 0.0445 | 0.0085 |
| 3.7 | 0.06 | 0.0459 | 0.0141 |
| 3.8 | 0.073 | 0.0546 | 0.0184 |
| 3.9 | 0.091 | 0.073 | 0.018 |
| 4 | 0.111 | 0.0965 | 0.0145 |
| 4.1 | 0.127 | 0.116 | 0.011 |
| 4.2 | 0.133 | 0.1257 | 0.0073 |
| 4.3 | 0.129 | 0.1266 | 0.0024 |
| 4.4 | 0.126 | 0.1254 | 0.0006 |
| 4.5 | 0.131 | 0.1277 | 0.0033 |
| 4.6 | 0.139 | 0.1306 | 0.0084 |
| 4.7 | 0.138 | 0.1264 | 0.0116 |
| 4.8 | 0.117 | 0.1103 | 0.0067 |
| 4.9 | 0.079 | 0.0832 | -0.0042 |
| 5 | 0.034 | 0.0518 | -0.0178 |

Observation 1: The measured temporal correlation of FR1 CDL-C UMa channel model matches well with the simulated reference values.

    - 2.2.2 Spatial correlation

Measurement results:

(a) Beam 1                                   (b) Beam 2

(c) Combined beams

Figure 3. Spatial correlation measurement results and simulated reference for CDL-C UMa, fc ≤ 2.5 GHz.

Table 4. Spatial correlation verification results for CDL-C UMa with Combined beams and ≤ 2.5 GHz

| CDL-C UMa, Combined beams, 2450 MHz |  |  |  |
| --- | --- | --- | --- |
| Azimuth [o] | Reference SCF | Measured SCF | delta |
| 270 | 1 | 1 | 0 |
| 263 | 0.99 | 0.983 | 0.007 |
| 256 | 0.95 | 0.9545 | -0.0045 |
| 249 | 0.91 | 0.9225 | -0.0125 |
| 242 | 0.87 | 0.8688 | 0.0012 |
| 234.9 | 0.85 | 0.8399 | 0.0101 |
| 227.9 | 0.83 | 0.8297 | 0.0003 |
| 220.9 | 0.82 | 0.8261 | -0.0061 |
| 213.9 | 0.82 | 0.8116 | 0.0084 |
| 206.9 | 0.8 | 0.7994 | 0.0006 |
| 199.9 | 0.77 | 0.7711 | -0.0011 |
| 192.9 | 0.73 | 0.7227 | 0.0073 |
| 185.9 | 0.66 | 0.6449 | 0.0151 |
| 144.9 | 0.26 | 0.2689 | -0.0089 |
| 109.9 | 0.23 | 0.2459 | -0.0159 |
| 74.8 | 0.19 | 0.1798 | 0.0102 |
| 39.8 | 0.13 | 0.11 | 0.02 |
| 4.7 | 0.15 | 0.1585 | -0.0085 |
| 329.7 | 0.38 | 0.3513 | 0.0287 |
| 294.6 | 0.74 | 0.7559 | -0.0159 |

Figure 4. Spatial correlation measurement results and simulated reference for CDL-C UMa, fc > 2.5 GHz.

Table 5. Spatial correlation verification results for CDL-C UMa with Combined beams and > 2.5 GHz

| CDL-C UMa, Combined beams, 3600 MHz |  |  |  |
| --- | --- | --- | --- |
| Azimuth [o] | Reference SCF | Measured result | delta |
| 270 | 1 | 1 | 0 |
| 265.2 | 0.98 | 0.99 | -0.01 |
| 260.5 | 0.95 | 0.9633 | -0.0133 |
| 255.7 | 0.9 | 0.8989 | 0.0011 |
| 250.9 | 0.84 | 0.8607 | -0.0207 |
| 246.1 | 0.8 | 0.7943 | 0.0057 |
| 241.4 | 0.77 | 0.7778 | -0.0078 |
| 236.6 | 0.75 | 0.7457 | 0.0043 |
| 231.8 | 0.73 | 0.7362 | -0.0062 |
| 227.1 | 0.72 | 0.7185 | 0.0015 |
| 222.3 | 0.71 | 0.7018 | 0.0082 |
| 217.5 | 0.7 | 0.7078 | -0.0078 |
| 212.7 | 0.69 | 0.693 | -0.003 |
| 208 | 0.67 | 0.6759 | -0.0059 |
| 203.2 | 0.64 | 0.6524 | -0.0124 |
| 198.4 | 0.61 | 0.6168 | -0.0068 |
| 193.7 | 0.56 | 0.5655 | -0.0055 |
| 188.9 | 0.49 | 0.4875 | 0.0025 |
| 184.1 | 0.41 | 0.4055 | 0.0045 |
| 156.1 | 0.42 | 0.3893 | 0.0307 |
| 132.3 | 0.19 | 0.2919 | -0.1019 |
| 108.4 | 0.64 | 0.6766 | -0.0366 |
| 84.6 | 0.47 | 0.5184 | -0.0484 |
| 60.7 | 0.44 | 0.4908 | -0.0508 |
| 36.9 | 0.28 | 0.3259 | -0.0459 |
| 13 | 0.16 | 0.2143 | -0.0543 |
| 349.1 | 0.16 | 0.1409 | 0.0191 |
| 325.3 | 0.41 | 0.4387 | -0.0287 |
| 301.4 | 0.4 | 0.4464 | -0.0464 |
| 277.6 | 0.95 | 0.9562 | -0.0062 |

Observation 2: The measured spatial correlation of FR1 CDL-C UMa channel model matches well with the simulated reference values.

3	Conclusion

Observation 1: The measured temporal correlation of FR1 CDL-C UMa channel model matches well with the simulated reference values.

Observation 2: The measured spatial correlation of FR1 CDL-C UMa channel model matches well with the simulated reference values.

References

[1] R4-2115758, “Time plan for FR1 lab alignment and requirement development”, CAICT, OPPO, vivo, RAN4#100-e, Aug. 2021.

[2] R4-2115756, “WF on NR MIMO OTA”, vivo, CAICT, OPPO, RAN4#100-e, Aug. 2021.

[3] R4-1914996, “FR1 MIMO OTA channel model validation results”, CAICT, Keysight, RAN4#93, Nov. 2019.

[4] R4-2016561, “FR1 MIMO OTA channel model validation results”, CAICT, Keysight, vivo, RAN4#97, Nov. 2020.

[5] R4-2115759, “Reference channel emulation curves for validation purposes”, Keysight Technologies, Spirent Communications, CMCC, CAICT, RAN4 #100-e, Aug. 2021.

[6] R4-2119379, “Beam specific reference channel emulation curves for validation purposes for FR1 CDL-C UMa”, Keysight Technologies, Spirent Communications, CMCC, CAICT, RAN4 #100-e, Nov. 2021.