3GPP TSG-RAN4 Meeting AH#2	R4- 1808652

Montreal, Canada, 2nd – 6th July 2018

Source:	Ericsson

Title:	TP to TR: MU framework and values for OTA Base station output power for CATR

Agenda item:	4.2.1.1.4.1

Document for:	Discussion

# Introduction

This contribution presents an example measurement uncertainty framework and assessment for OTA Base station output power measurement in a compact antenna test range which is one possible method for conformance testing of AAS base stations.

# Discussion

Based on eAAS OTA radiated transmit power in [1], framework on uncertainty budget have been reused.

Compact antenna test range uncertainty contributions for OTA Base station output power measurement

| UID | Description of uncertainty contribution | Details in annex of [1] |
| --- | --- | --- |
| Stage 2: DUT measurement |  |  |
| 1 | Misalignment DUT & pointing error | B2-1 |
| 2 | RF power measurement equipment | E |
| 3 | Standing wave between DUT and test range antenna | B2-3 |
| 4 | RF leakage (SGH connector terminated & test range antenna connector cable terminated) | B2-4 |
| 5 | QZ ripple DUT | B2-5 |
| 19 | Miscellaneous uncertainty | B2-14 |
| Stage 1: Calibration measurement |  |  |
| 6 | Uncertainty of network analyser | E |
| 7 | Mismatch of receiver chain | B2-7 |
| 8 | Insertion loss variation of receiver chain | B2-8 |
| 9 | RF leakage, (SGH connector terminated & test range antenna connector cable terminated) | B2-4 |
| 10 | Influence of the calibration antenna feed cable: a) Flexing cables, adapters, attenuators, connector repeatability | B2-9 |
| 11 | Uncertainty of the absolute gain of the calibration antenna | E |
| 12 | Misalignment positioning system | B2-11 |
| 13 | Misalignment of calibration antenna and test range antenna | B2-1 |
| 14 | Rotary Joints | B2-12 |
| 15 | Standing wave between reference calibration antenna and test range antenna | B2-3 |
| 16 | Quality of quiet zone | B2-5 |
| 20 | Switching uncertainty | B2-15 |

# Compact antenna test range uncertainty assessment for OTA Base station output power measurement

| UID | Uncertainty Source | Uncertainty value f ≦ 3GHz | Uncertainty value 3GHz ≦ f < 4.2 GHz | Distribution of the probability | Divisor based on distribution shape | ci | Standard uncertainty ui [dB] f ≦ 3GHz | Standard uncertainty ui [dB] 3GHz < f ≦ 4.2 GHz |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |
| 1 | Misalignment DUT & pointing error | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 |
| 2 | RF power measurement equipment (e.g. spectrum analyzer, power meter) | 0.14 | 0.26 | Gaussian | 1 | 1 | 0.14 | 0.26 |
| 3 | Standing wave between DUT and test range antenna | 0.21 | 0.21 | U-shaped | √2 | 1 | 0.15 | 0.15 |
| 4 | RF leakage, test range antenna cable connector terminated. | 0.0012 | 0.0012 | Normal | 1 | 1 | 0.0012 | 0.0012 |
| 5 | QZ ripple with DUT | 0.0928 | 0.0928 | Normal | 1 | 1 | 0.0928 | 0.0928 |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |  |
| 6 | Network Analyzer | 0.13 | 0.20 | Normal | 1 | 1 | 0.13 | 0.20 |
| 7 | Uncertainty of return loss (S11) measurement of SGH and test receiver (VNA) ports | 0.127 | 0.325 | U-shaped | √2 | 1 | 0.09 | 0.23 |
| 8 | Insertion loss variation in receiver chain | 0.18 | 0.18 | Rectangular | √3 | 1 | 0.10 | 0.10 |
| 9 | RF leakage, test range antenna cable connector terminated. | 0.0012 | 0.0012 | Normal | 1 | 1 | 0.0012 | 0.0012 |
| 10 | Influence of the calibration antenna feed cable | 0.022 | 0.022 | U-shaped | √2 | 1 | 0.015 | 0.015 |
| 11 | SGH Calibration uncertainty | 0.50 | 0.433 | Rectangular | √3 | 1 | 0.29 | 0.25 |
| 12 | Misalignment positioning system | 0 | 0 | Exp. normal | 2 | 1 | 0 | 0 |
| 13 | Misalignment SGH and pointing error | 0.5 | 0.5 | Exp. normal | 2 | 1 | 0.25 | 0.25 |
| 14 | Rotary joints | 0.048 | 0.048 | U-shaped | √2 | 1 | 0.034 | 0.034 |
| 15 | Standing wave between SGH and test range antenna | 0.09 | 0.09 | U-shaped | √2 | 1 | 0.06 | 0.06 |
| 16 | QZ ripple with SGH | 0.009 | 0.009 | Normal | 1 | 1 | 0.009 | 0.009 |
| 17 | Switching uncertainty | 0.26 | 0.26 | Rectangular | √3 | 1 | 0.15 | 0.15 |
| Combined standard uncertainty (1σ) [dB] |  |  |  |  |  |  | 0.51 | 0.60 |
| Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  |  | 0.99 | 1.17 |

# Conclusion

This contribution presents an example of the measurement uncertainty framework and budget for OTA Base station output power measurements.

# References

TS 37.842, Evolved Universal Terrestrial Radio Access (E-UTRA) and Universal Terrestrial Radio Access (UTRA; Radio Frequency (RF) requirement background for Active Antenna System (AAS) Base Station (BS)

Huawei, R4-1808500, “Draft CR for 37.843”, RAN4#87, May 2018