3GPP TSG-RAN5 Meeting # 91-e	R5-213217r1

# Electronic Meeting,  17th May - 28th May 2021

| CR-Form-v12.1 |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CHANGE REQUEST |  |  |  |  |  |  |  |  |
|  | 38.903 | CR | 0234 | rev | - | Current version: | 16.7.0 |  |
| For HELP on using this form: comprehensive instructions can be found at http://www.3gpp.org/Change-Requests. |  |  |  |  |  |  |  |  |

| Proposed change affects: | UICC apps |  | ME |  | Radio Access Network |  | Core Network |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |

| Title: | Measurement Uncertainties updates for FR2 Extreme Testing Conditions |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Source to WG: | Keysight Technologies |  |  |  |  |  |  |  |  |  |
| Source to TSG: | R5 |  |  |  |  |  |  |  |  |  |
| Work item code: | 5GS_NR_LTE-UEConTest |  |  |  |  |  | Date: |  |  | 2021-05-06 |
| Category: | F |  |  |  |  |  | Release: |  |  | Rel-16 |
|  | Use one of the following categories: F (correction) A (mirror corresponding to a change in an earlier release) B (addition of feature), C (functional modification of feature) D (editorial modification) Detailed explanations of the above categories can be found in 3GPP TR 21.900. |  |  |  |  |  |  |  | Use one of the following releases: Rel-8 (Release 8) Rel-9 (Release 9) Rel-10 (Release 10) Rel-11 (Release 11) … Rel-15 (Release 15) Rel-16 (Release 16) Rel-17 (Release 17) Rel-18 (Release 18) |  |
| Reason for change: |  | As per R5-213210 and R5-213212, some additional measurements should be run under extreme temperature condition. Measurement uncertainties for those test cases under ETC are not defined |  |  |  |  |  |  |  |  |
| Summary of change: |  | Define measurement uncertainties for extreme temperature condition for the following measurements: - UE maximum output power – TRP - UE maximum output power for modulation / channel bandwidth - Minimum output power - EIRP OFF power - Frequency error - Adjacent Channel Leakage Ratio |  |  |  |  |  |  |  |  |
| Consequences if not approved: |  | Test specification will remain incomplete |  |  |  |  |  |  |  |  |
| Clauses affected: |  | B.3, B.4, B.7, B.8, B.10, B.17 |  |  |  |  |  |  |  |  |
|  |  | Y | N |  |  |  |  |  |  |  |
| Other specs |  |  | X | Other core specifications |  |  |  | TS/TR ... CR ... |  |  |
| affected: |  | X |  | Test specifications |  |  |  | TS/TR 38.521-2 CR 0534 |  |  |
| (show related CRs) |  |  | X | O&M Specifications |  |  |  | TS/TR ... CR ... |  |  |
| Other comments: |  | This CR agreement requires R5-213210 proposals endorsement and R5-213212 agreement. |  |  |  |  |  |  |  |  |
| This CR's revision history: |  | Revision 1: Year added in the coverpage date. |  |  |  |  |  |  |  |  |

<< Start of change 1>>

# B.3	UE maximum output power

Following tables summarize the MU threshold for EIRP and TRP measurements for UE maximum output power. The origin MU values for different test setups with varies parameters can be found in following clauses.

Table B.3-1: MU threshold for EIRP measurement for UE maximum output power

| Power Class | Frequency | MBW | Power (NOTE2) | Threshold MU value for NTC [dB] (NOTE1) | Threshold MU value for ETC [dB] (NOTE1) |
| --- | --- | --- | --- | --- | --- |
| PC3 | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = Max Output Power | 4.89 | 5.17 |
|  | 32.125GHz < f <= 40.8GHz |  |  | 5.09 | 5.37 |
| PC1 | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = Max Output Power | FFS | FFS |
|  | 32.125GHz < f <= 40.8GHz |  |  | FFS | FFS |
| NOTE 1: Total EIRP Expanded MU for IFF for Quiet Zone size ≤30cm in Table B.3.2-2 for PC3 UEs (NTC), in Table B.3.2-8 for PC3 UEs (ETC) and B.3.2-6 for PC1 UEs. NOTE 2: Max output power level for device with corresponding power class. |  |  |  |  |  |

Table B.3-2: MU threshold for TRP measurement for UE maximum output power

| Power Class | Frequency |  | MBW |  | Power (NOTE2) |  | Threshold MU value for NTC [dB] (NOTE 1) |  | Threshold MU value for ETC [dB] (NOTE 1) |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| PC3 | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = Max Output Power | 4.42 | 4.70 |
|  | 32.125GHz < f <= 40.8GHz |  |  | 4.62 | 4.90 |
| PC1 | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = Max Output Power | FFS | FFS |
|  | 32.125GHz < f <= 40.8GHz |  |  | FFS | FFS |
| NOTE 1: Total TRP Expanded MU for IFF for Quiet Zone size ≤ 30cm in Table B.3.2-2 for PC3 UEs and B.3.2-6 for PC1 UEs NOTE 2: Max output power level for device with corresponding power class. |  |  |  |  |  |  |  |  |  |  |

Table B.3-3: MU threshold for Spherical coverage measurement for UE maximum output power

| Power Class | Frequency | MBW | Power | Threshold MU value (NOTE 1) |
| --- | --- | --- | --- | --- |
| PC3 | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = TBD | 4.60 |
|  | 32.125GHz < f <= 40.8GHz |  |  | 5.20 |
| PC1 | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = TBD | FFS |
|  | 32.125GHz < f <= 40.8GHz |  |  | FFS |
| NOTE 1: Total Spherical coverage Expanded MU for IFF for Quiet Zone size ≤ 30cm in Tables B.3.2-4 for PC3 UEs and B.3.2-7 for PC1 UEs |  |  |  |  |

## B.3.1	Uncertainty budget format and assessment for DFF

The uncertainty contributions that may impact the overall MU value are listed in Table B.3.1-1.

Table B.3.1-1: Uncertainty contributions for EIRP and TRP measurement

| UID | Description of uncertainty contribution | Details in annex |
| --- | --- | --- |
| Stage 2: DUT measurement |  |  |
| 1 | Positioning misalignment | B.2.1.1 |
| 2 | Measure distance uncertainty | B.2.1.2 |
| 3 | Quality of quiet zone | B.2.1.3 |
| 4 | Mismatch | B.2.1.4 |
| 5 | Standing Wave Between the DUT and measurement antenna | B.2.1.5 |
| 6 | Uncertainty of the RF power measurement equipment | B.2.1.6 |
| 7 | Phase curvature | B.2.1.7 |
| 8 | Amplifier uncertainties | B.2.1.8 |
| 9 | Random uncertainty | B.2.1.9 |
| 10 | Influence of the XPD | B.2.1.10 |
| 11 | Insertion Loss Variation | B.2.1.11 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | B.2.1.12 |
| 13 | Influence of TRP measurement grid | B.2.1.22 |
| 14 | Influence of beam peak search grid | B.2.1.23 |
| 15 | Multiple measurement antenna uncertainty | B.2.1.25 |
| 16 | DUT repositioning | B.2.1.26 |
| 17 | Influence of spherical coverage grid | B.2.1.29 |
| Stage 1: Calibration measurement |  |  |
| 18 | Mismatch | B.2.1.4 |
| 19 | Amplifier uncertainties | B.2.1.8 |
| 20 | Misalignment of positioning System | B.2.1.13 |
| 21 | Uncertainty of the Network Analyzer | B.2.1.14 |
| 22 | Uncertainty of the absolute gain of the calibration antenna | B.2.1.15 |
| 23 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | B.2.1.16 |
| 24 | Phase centre offset of calibration antenna | B.2.1.18 |
| 25 | Quality of quiet zone for calibration process | B.2.1.19 |
| 26 | Standing wave between reference calibration antenna and measurement antenna | B.2.1.20 |
| 27 | Influence of the calibration antenna feed cable | B.2.1.21 |
| 28 | Insertion Loss Variation | B.2.1.11 |
| Systematic uncertainties |  |  |
| 29 | Systematic error due to TRP calculation/quadrature | B.2.1.24 |
| 30 | Influence of noise | B.2.1.27 |
| 31 | Systematic error related to beam peak search | B.2.1.28 |

The uncertainty assessment tables are organized as follows:

-	For the purpose of uncertainty assessment, the radiating antenna aperture of the DUT is denoted as D

-	The uncertainty assessment has been derived for the case of D = [5 cm], f = {22.65GHz, 31.1GHz, 45.1GHz}, P = [maximum output power].

-	The uncertainty assessment for EIRP and TRP is provided in Table B.3.1-2.

Table B.3.1-2: Uncertainty assessment for EIRP and TRP measurement (f=TBD, D=TBD)

| UID |  | Uncertainty source |  | Uncertainty value |  | Distribution of the probability |  | Divisor |  | Standard uncertainty (σ) [dB] |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |  |  |  |
| 1 |  | Positioning misalignment |  |  |  |  |  |  |  |  |  |
| 2 |  | Measure distance uncertainty |  |  |  |  |  |  |  |  |  |
| 3 |  | Quality of quiet zone (NOTE 2) |  |  |  |  |  |  |  |  |  |
| 4 |  | Mismatch (NOTE 3) |  |  |  |  |  |  |  |  |  |
| 5 |  | Standing Wave Between the DUT and measurement antenna |  |  |  |  |  |  |  |  |  |
| 6 |  | Uncertainty of the RF power measurement equipment (NOTE 4) |  |  |  |  |  |  |  |  |  |
| 7 |  | Phase curvature |  |  |  |  |  |  |  |  |  |
| 8 |  | Amplifier uncertainties |  |  |  |  |  |  |  |  |  |
| 9 |  | Random uncertainty |  |  |  |  |  |  |  |  |  |
| 10 |  | Influence of the XPD |  |  |  |  |  |  |  |  |  |
| 11 |  | Insertion Loss Variation |  |  |  |  |  |  |  |  |  |
| 12 |  | RF leakage (from measurement antenna to the receiver/transmitter) |  |  |  |  |  |  |  |  |  |
| 13 |  | Influence of TRP measurement grid (NOTE 5) |  | 0.25 |  | Actual |  | 1 |  | 0.25 |  |
| 14 |  | Influence of beam peak search grid (NOTE 6) |  | 0.0 |  | Actual |  | 1 |  | 0.0 |  |
| 15 |  | Multiple measurement antenna uncertainty |  |  |  |  |  |  |  |  |  |
| 16 |  | DUT repositioning |  |  |  |  |  |  |  |  |  |
| 17 |  | Influence of spherical coverage grid (NOTE 8) |  | 0.12 |  | Actual |  | 1 |  | 0.12 |  |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |  |  |  |  |
| 18 |  | Mismatch |  |  |  |  |  |  |  |  |  |
| 19 |  | Amplifier uncertainties |  |  |  |  |  |  |  |  |  |
| 20 |  | Misalignment of positioning System |  |  |  |  |  |  |  |  |  |
| 21 |  | Uncertainty of the Network Analyzer |  |  |  |  |  |  |  |  |  |
| 22 |  | Uncertainty of the absolute gain of the calibration antenna |  |  |  |  |  |  |  |  |  |
| 23 |  | Positioning and pointing misalignment between the reference antenna and the measurement antenna |  |  |  |  |  |  |  |  |  |
| 24 |  | Phase centre offset of calibration antenna |  |  |  |  |  |  |  |  |  |
| 25 |  | Quality of quiet zone for calibration process (NOTE 2) |  |  |  |  |  |  |  |  |  |
| 26 |  | Standing wave between reference calibration antenna and measurement antenna |  |  |  |  |  |  |  |  |  |
| 27 |  | Influence of the calibration antenna feed cable |  |  |  |  |  |  |  |  |  |
| 28 |  | Insertion Loss Variation |  |  |  |  |  |  |  |  |  |
|  |  | Systematic uncertainties (NOTE 7) |  |  |  |  |  |  |  | Value |  |
| 29 |  | Systematic error due to TRP calculation/quadrature (NOTE 5) |  |  |  |  |  |  |  | 0.00 |  |
| 30 |  | Influence of noise |  |  |  |  |  |  |  |  |  |
| 31 |  | Systematic error related to beam peak search (NOTE 6) |  |  |  |  |  |  |  | 0.5 |  |
| Total measurement uncertainty |  |  |  |  |  |  |  |  |  | Value |  |
| EIRP Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  |  |  |  |  | TBD |  |
| TRP Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  |  |  |  |  | TBD |  |
| NOTE 1: The impact of phase variation on EIRP shall be taken into account during final MU definition for the test method.. NOTE 2: The quality of quiet zone is different for EIRP and TRP. For TRP, the standard uncertainty is FFS; for EIRP, the standard uncertainty of quiet zone is FFS. NOTE 3: The analysis was done only for the case of operating at max output power, in-band, non-CA. NOTE 4: The assessment assumes maximum DUT output power. NOTE 5: This contributor shall only be considered for TRP measurements. NOTE 6: This contributor shall only be considered for EIRP measurements. NOTE 7: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 8: This contributor shall only be considered for spherical EIRP measurements |  |  |  |  |  |  |  |  |  |  |  |

## B.3.2	Uncertainty budget format and assessment for IFF

The uncertainty contributions that may impact the overall MU value are listed in Table B.3.2-1.

Table B.3.2-1: Uncertainty contributions for EIRP and TRP measurement

| UID |  | Description of uncertainty contribution |  | Details in clause |  |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 |  | Positioning misalignment |  | B.2.2.1 |  |
| 2 |  | Measure distance uncertainty |  | B.2.2.2 |  |
| 3 |  | Quality of Quiet Zone |  | B.2.2.3 |  |
| 4 |  | Mismatch |  | B.2.2.4 |  |
| 5 |  | Standing wave between the DUT and measurement antenna |  | B.2.2.5 |  |
| 6 |  | Uncertainty of the RF power measurement equipment |  | B.2.2.6 |  |
| 7 |  | Phase curvature |  | B.2.2.7 |  |
| 8 |  | Amplifier uncertainties |  | B.2.2.8 |  |
| 9 |  | Random uncertainty |  | B.2.2.9 |  |
| 10 |  | Influence of the XPD |  | B.2.2.10 |  |
| 11 |  | Insertion Loss Variation |  | B.2.2.11 |  |
| 12 |  | RF leakage (from measurement antenna to the receiver/transmitter) |  | B.2.2.12 |  |
| 13 |  | Influence of TRP measurement grid |  | B.2.2.22 |  |
| 14 |  | Influence of beam peak search grid |  | B.2.2.23 |  |
| 15 |  | Multiple measurement antenna uncertainty |  | B.2.2.25 |  |
| 16 |  | DUT repositioning |  | B.2.2.26 |  |
| 17 |  | Influence of spherical coverage grid |  | B.2.2.29 |  |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 18 |  | Mismatch |  | B.2.2.4 |  |
| 19 |  | Amplifier Uncertainties |  | B.2.2.8 |  |
| 20 |  | Misalignment of positioning System |  | B.2.2.13 |  |
| 21 |  | Uncertainty of the Network Analyzer |  | B.2.2.14 |  |
| 22 |  | Uncertainty of the absolute gain of the calibration antenna |  | B.2.2.15 |  |
| 23 |  | Positioning and pointing misalignment between the reference antenna and the measurement antenna |  | B.2.2.16 |  |
| 24 |  | Phase centre offset of calibration antenna |  | B.2.2.18 |  |
| 25 |  | Quality of quiet zone for calibration process |  | B.2.2.19 |  |
| 26 |  | Standing wave between reference calibration antenna and measurement antenna |  | B.2.2.20 |  |
| 27 |  | Influence of the calibration antenna feed cable |  | B.2.2.21 |  |
| 28 |  | Insertion Loss Variation |  | B.2.1.11 |  |
| Systematic uncertainties |  |  |  |  |  |
| 29 |  | Systematic error due to TRP calculation/quadrature |  | B.2.2.24 |  |
| 30 |  | Influence of noise |  | B.2.1.27 |  |
| 31 |  | Systematic error related to beam peak search |  | B.2.2.28 |  |

The uncertainty assessment tables are organized as follows:

-	For the purpose of uncertainty assessment, the radiating antenna aperture of the DUT is denoted as D

-	The uncertainty assessment has been derived for the case of Quiet Zone size ≤ [30 cm], f = {23.45GHz, 32.125GHz, 40.8GHz}, [P = maximum output power].

-	The uncertainty assessment for EIRP and TRP is provided in Table B.3.2-2 for PC3 UEs and in Table B.3.2-6 for PC1 UEs.

-	The uncertainty assessment for Spherical coverage is provided in Table B.3.2-4 for PC3 UEs in Table B.3.2-7 for PC1 UEs.

Table B.3.2-2: Uncertainty assessment for EIRP and TRP measurement (f=23.45GHz, 32.125GHz, 40.8GHz, Quiet Zone size ≤ 30 cm) for PC3 UEs and normal temperature condition

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment | 0.00 | Normal | 2.00 | 0.00 |
| 2 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 3 | Quality of Quiet Zone (NOTE 1) | 0.6 | Actual | 1.00 | 0.6 |
| 4 | Mismatch | 1.30 | Actual | 1.00 | 1.30 |
| 5 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | Uncertainty of the RF power measurement equipment (NOTE 3) | 2.16 | Normal | 2.00 | 1.08 |
| 7 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 8 | Amplifier uncertainties | 2.10 | Normal | 2.00 | 1.05 |
| 9 | Random uncertainty | 0.50 | Normal | 2.00 | 0.25 |
| 10 | Influence of the XPD | 0.01 | U-shaped | 1.41 | 0.00 |
| 11 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 13 | Influence of TRP measurement grid (NOTE 4) | 0.25 | Actual | 1 | 0.25 |
| 14 | Influence of beam peak search grid (NOTE 5) | 0.00 | Actual | 1 | 0.00 |
| 15 | Multiple measurement antenna uncertainty (NOTE 9) | 0.15 | Actual | 1 | 0.15 |
| 16 | DUT repositioning | 0.00 (NOTE 4) 0.08 (NOTE 5) | Rectangular | 1.73 | 0.00 (NOTE 4) 0.05 (NOTE 5) |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 17 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 18 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 19 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 20 | Uncertainty of the Network Analyzer | 0.73 | Normal | 2.00 | 0.37 |
| 21 | Uncertainty of the absolute gain of the calibration antenna | 0.60 | Normal | 2.00 | 0.30 |
| 22 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.01 | Rectangular | 1.73 | 0.00 |
| 23 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 24 | Quality of quiet zone for calibration process (NOTE 1) | 0.4 | Actual | 1.00 | 0.4 |
| 25 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 26 | Influence of the calibration antenna feed cable | 0.14 | Normal | 2.00 | 0.07 |
| 27 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
|  | Systematic uncertainties (NOTE 6) |  |  |  | Value |
| 28 | Systematic error due to TRP calculation/quadrature (NOTE 4) |  |  |  | 0.00 |
| 29 | Influence of noise (23.45GHz <= f <= 32.125GHz) |  |  |  | 0.1 |
| 29 | Influence of noise (32.125GHz < f <= 40.8GHz) |  |  |  | 0.3 |
| 30 | Systematic error related to beam peak search (NOTE 5) |  |  |  | 0.5 |
| Total measurement uncertainty |  |  |  |  | Value |
| EIRP Expanded uncertainty (23.45GHz <= f <= 32.125GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 4.89 |
| EIRP Expanded uncertainty (32.125GHz < f <= 40.8GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 5.09 |
| TRP Expanded uncertainty (23.45GHz <= f <= 32.125GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 4.42 |
| TRP Expanded uncertainty (32.125GHz < f <= 40.8GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 4.62 |
| NOTE 1: The quality of quiet zone is the same for EIRP and TRP. Value based on procedure defined in clause D.2 of TR 38.810 for Quiet Zone size less or equal to 30 cm. NOTE 2: The analysis was done only for the case of operating at max output power, in-band, non-CA. NOTE 3: The assessment assumes maximum DUT output power. NOTE 4: This contributor shall only be considered for TRP measurements. NOTE 5: This contributor shall only be considered for EIRP measurements. NOTE 6: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 7: Void. NOTE 8: Void NOTE 9: Applies to the system which has a structure of mechanical feed antenna positioning. |  |  |  |  |  |

Table B.3.2-3: Void

Table B.3.2-4: Uncertainty assessment for Spherical coverage measurement (f=23.45GHz, 32.125GHz, 40.8GHz, Quiet Zone size ≤ 30 cm) for PC3 UEs and normal temperature condition

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment | 0.00 | Normal | 2.00 | 0.00 |
| 2 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 3 | Quality of Quiet Zone (NOTE 1) | 0.6 | Actual | 1.00 | 0.6 |
| 4 | Mismatch | 1.30 | Actual | 1.00 | 1.30 |
| 5 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | Uncertainty of the RF power measurement equipment (NOTE 3) | 2.16 | Normal | 2.00 | 1.08 |
| 7 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 8 | Amplifier uncertainties | 2.1 | Normal | 2.00 | 1.05 |
| 9 | Random uncertainty | 0.50 | Normal | 2.00 | 0.25 |
| 10 | Influence of the XPD | 0.01 | U-shaped | 1.41 | 0.00 |
| 11 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 13 | Multiple measurement antenna uncertainty (NOTE 5) | 0.15 | Actual | 1 | 0.15 |
| 14 | DUT repositioning | 0.00 | Rectangular | 1.73 | 0.00 |
| 15 | Influence of spherical coverage grid | 0.12 | Actual | 1 | 0.12 |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 16 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 17 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 18 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 19 | Uncertainty of the Network Analyzer | 0.73 | Normal | 2.00 | 0.37 |
| 20 | Uncertainty of the absolute gain of the calibration antenna | 0.60 | Normal | 2.00 | 0.30 |
| 21 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.01 | Rectangular | 1.73 | 0.00 |
| 22 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 23 | Quality of quiet zone for calibration process (NOTE 1) | 0.4 | Actual | 1.00 | 0.4 |
| 24 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 25 | Influence of the calibration antenna feed cable | 0.14 | Normal | 2.00 | 0.07 |
| 26 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
|  | Systematic uncertainties (NOTE 4) |  |  |  | Value |
| 27 | Influence of noise (23.45GHz <= f <= 32.125GHz) |  |  |  | 0.3 |
| 27 | Influence of noise (32.125GHz < f <= 40.8GHz) |  |  |  | 0.9 |
| Total measurement uncertainty |  |  |  |  | Value |
| Spherical coverage Expanded uncertainty (23.45GHz <= f <= 32.125GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 4.60 |
| Spherical coverage Expanded uncertainty (32.125GHz < f <= 40.8GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 5.19 |
| NOTE 1: The quality of quiet zone is the same for EIRP and TRP. Value based on procedure defined in clause D.2 of TR 38.810 for Quiet Zone size less or equal to 30 cm. NOTE 2: The analysis was done only for the case of operating at max output power, in-band, non-CA. NOTE 3: The assessment assumes maximum DUT output power. NOTE 4: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 5: Applies to the system which has a structure of mechanical feed antenna positioning. |  |  |  |  |  |

Table B.3.2-5: Void

Table B.3.2-6: Uncertainty assessment for EIRP and TRP measurement (f=23.45GHz, 32.125GHz, 40.8GHz, Quiet Zone size ≤ 30 cm) for PC1 UEs and normal temperature condition

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment | 0.02 | Normal | 2.00 | 0.01 |
| 2 | Measure distance uncertainty | FFS | Rectangular | 1.73 | FFS |
| 3 | Quality of Quiet Zone (NOTE 1) | FFS | Actual | 1.00 | FFS |
| 4 | Mismatch | FFS | Actual | 1.00 | FFS |
| 5 | Standing wave between the DUT and measurement antenna | FFS | U-shaped | 1.41 | FFS |
| 6 | Uncertainty of the RF power measurement equipment (NOTE 3) | FFS | Normal | 2.00 | FFS |
| 7 | Phase curvature | FFS | U-shaped | 1.41 | FFS |
| 8 | Amplifier uncertainties | FFS | Normal | 2.00 | FFS |
| 9 | Random uncertainty | FFS | Normal | 2.00 | FFS |
| 10 | Influence of the XPD | FFS | U-shaped | 1.41 | FFS |
| 11 | Insertion Loss Variation | FFS | Rectangular | 1.73 | FFS |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | FFS | Actual | 1.00 | FFS |
| 13 | Influence of TRP measurement grid (NOTE 4) | 0.25 | Actual | 1 | 0.25 |
| 14 | Influence of beam peak search grid (NOTE 5) | 0.00 | Actual | 1 | 0.00 |
| 15 | Multiple measurement antenna uncertainty (NOTE 9) | FFS | Actual | 1 | FFS |
| 16 | DUT repositioning | 0.00 (NOTE 4) 0.35 (NOTE 5) | Rectangular | 1.73 | 0.00 (NOTE 4) 0.20 (NOTE 5) |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 17 | Mismatch | FFS | U-shaped | 1.41 | FFS |
| 18 | Amplifier Uncertainties | FFS | Normal | 2.00 | FFS |
| 19 | Misalignment of positioning System | FFS | Normal | 2.00 | FFS |
| 20 | Uncertainty of the Network Analyzer | FFS | Normal | 2.00 | FFS |
| 21 | Uncertainty of the absolute gain of the calibration antenna | FFS | Normal | 2.00 | FFS |
| 22 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | FFS | Rectangular | 1.73 | FFS |
| 23 | Phase centre offset of calibration antenna | FFS | Rectangular | 1.73 | FFS |
| 24 | Quality of quiet zone for calibration process (NOTE 1) | FFS | Actual | 1.00 | FFS |
| 25 | Standing wave between reference calibration antenna and measurement antenna | FFS | U-shaped | 1.41 | FFS |
| 26 | Influence of the calibration antenna feed cable | FFS | Normal | 2.00 | FFS |
| 27 | Insertion Loss Variation | FFS | Rectangular | 1.73 | FFS |
|  | Systematic uncertainties (NOTE 6) |  |  |  | Value |
| 28 | Systematic error due to TRP calculation/quadrature (NOTE 4) |  |  |  | 0.00 |
| 29 | Influence of noise (23.45GHz <= f <= 32.125GHz) |  |  |  | FFS |
| 29 | Influence of noise (32.125GHz < f <= 40.8GHz) |  |  |  | FFS |
| 30 | Systematic error related to beam peak search (NOTE 5) |  |  |  | 0.7 |
| Total measurement uncertainty |  |  |  |  | Value |
| EIRP Expanded uncertainty (23.45GHz <= f <= 32.125GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | FFS |
| EIRP Expanded uncertainty (32.125GHz < f <= 40.8GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | FFS |
| TRP Expanded uncertainty (23.45GHz <= f <= 32.125GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | FFS |
| TRP Expanded uncertainty (32.125GHz < f <= 40.8GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | FFS |
| NOTE 1: The quality of quiet zone is the same for EIRP and TRP. Value based on procedure defined in clause D.2 of TR 38.810 for Quiet Zone size less or equal to 30 cm. NOTE 2: The analysis was done only for the case of operating at max output power, in-band, non-CA. NOTE 3: The assessment assumes maximum DUT output power. NOTE 4: This contributor shall only be considered for TRP measurements. NOTE 5: This contributor shall only be considered for EIRP measurements. NOTE 6: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 7: Void. NOTE 8: Void NOTE 9: Applies to the system which has a structure of mechanical feed antenna positioning. |  |  |  |  |  |

Table B.3.2-7: Uncertainty assessment for Spherical coverage measurement (f=23.45GHz, 32.125GHz, 40.8GHz, Quiet Zone size ≤ 30 cm) for PC1 UEs and normal temperature condition

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment | 0.02 | Normal | 2.00 | 0.01 |
| 2 | Measure distance uncertainty | FFS | Rectangular | 1.73 | FFS |
| 3 | Quality of Quiet Zone (NOTE 1) | FFS | Actual | 1.00 | FFS |
| 4 | Mismatch | FFS | Actual | 1.00 | FFS |
| 5 | Standing wave between the DUT and measurement antenna | FFS | U-shaped | 1.41 | FFS |
| 6 | Uncertainty of the RF power measurement equipment (NOTE 3) | FFS | Normal | 2.00 | FFS |
| 7 | Phase curvature | FFS | U-shaped | 1.41 | FFS |
| 8 | Amplifier uncertainties | FFS | Normal | 2.00 | FFS |
| 9 | Random uncertainty | FFS | Normal | 2.00 | FFS |
| 10 | Influence of the XPD | FFS | U-shaped | 1.41 | FFS |
| 11 | Insertion Loss Variation | FFS | Rectangular | 1.73 | FFS |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | FFS | Actual | 1.00 | FFS |
| 13 | Multiple measurement antenna uncertainty (NOTE 5) | FFS | Actual | 1 | FFS |
| 14 | DUT repositioning | 0.00 | Rectangular | 1.73 | 0.00 |
| 15 | Influence of spherical coverage grid | 0.13 | Actual | 1 | 0.13 |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 16 | Mismatch | FFS | U-shaped | 1.41 | FFS |
| 17 | Amplifier Uncertainties | FFS | Normal | 2.00 | FFS |
| 18 | Misalignment of positioning System | FFS | Normal | 2.00 | FFS |
| 19 | Uncertainty of the Network Analyzer | FFS | Normal | 2.00 | FFS |
| 20 | Uncertainty of the absolute gain of the calibration antenna | FFS | Normal | 2.00 | FFS |
| 21 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | FFS | Rectangular | 1.73 | FFS |
| 22 | Phase centre offset of calibration antenna | FFS | Rectangular | 1.73 | FFS |
| 23 | Quality of quiet zone for calibration process (NOTE 1) | FFS | Actual | 1.00 | FFS |
| 24 | Standing wave between reference calibration antenna and measurement antenna | FFS | U-shaped | 1.41 | FFS |
| 25 | Influence of the calibration antenna feed cable | FFS | Normal | 2.00 | FFS |
| 26 | Insertion Loss Variation | FFS | Rectangular | 1.73 | FFS |
|  | Systematic uncertainties (NOTE 4) |  |  |  | Value |
| 27 | Influence of noise (23.45GHz <= f <= 32.125GHz) |  |  |  | FFS |
| 27 | Influence of noise (32.125GHz < f <= 40.8GHz) |  |  |  | FFS |
| Total measurement uncertainty |  |  |  |  | Value |
| Spherical coverage Expanded uncertainty (23.45GHz <= f <= 32.125GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | FFS |
| Spherical coverage Expanded uncertainty (32.125GHz < f <= 40.8GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | FFS |
| NOTE 1: The quality of quiet zone is the same for EIRP and TRP. Value based on procedure defined in clause D.2 of TR 38.810 for Quiet Zone size less or equal to 30 cm. NOTE 2: The analysis was done only for the case of operating at max output power, in-band, non-CA. NOTE 3: The assessment assumes maximum DUT output power. NOTE 4: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 5: Applies to the system which has a structure of mechanical feed antenna positioning. |  |  |  |  |  |

Table B.3.2-8: Uncertainty assessment for EIRP and TRP measurement (f=23.45GHz, 32.125GHz, 40.8GHz, Quiet Zone size ≤ 30 cm) for PC3 UEs and extreme temperature condition

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment | 0.00 | Normal | 2.00 | 0.00 |
| 2 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 3 | Quality of Quiet Zone (NOTE 1) | 0.9 | Actual | 1.00 | 0.9 |
| 4 | Mismatch | 1.30 | Actual | 1.00 | 1.30 |
| 5 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | Uncertainty of the RF power measurement equipment (NOTE 3) | 2.16 | Normal | 2.00 | 1.08 |
| 7 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 8 | Amplifier uncertainties | 2.10 | Normal | 2.00 | 1.05 |
| 9 | Random uncertainty | 0.50 | Normal | 2.00 | 0.25 |
| 10 | Influence of the XPD | 0.01 | U-shaped | 1.41 | 0.00 |
| 11 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 13 | Influence of TRP measurement grid (NOTE 4) | 0.25 | Actual | 1 | 0.25 |
| 14 | Influence of beam peak search grid (NOTE 5) | 0.00 | Actual | 1 | 0.00 |
| 15 | Multiple measurement antenna uncertainty (NOTE 7) | 0.15 | Actual | 1 | 0.15 |
| 16 | DUT repositioning | 0.00 (NOTE 4) 0.08 (NOTE 5) | Rectangular | 1.73 | 0.00 (NOTE 4) 0.05 (NOTE 5) |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 17 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 18 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 19 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 20 | Uncertainty of the Network Analyzer | 0.73 | Normal | 2.00 | 0.37 |
| 21 | Uncertainty of the absolute gain of the calibration antenna | 0.60 | Normal | 2.00 | 0.30 |
| 22 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.01 | Rectangular | 1.73 | 0.00 |
| 23 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 24 | Quality of quiet zone for calibration process (NOTE 1) | 0.6 | Actual | 1.00 | 0.6 |
| 25 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 26 | Influence of the calibration antenna feed cable | 0.14 | Normal | 2.00 | 0.07 |
| 27 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
|  | Systematic uncertainties (NOTE 6) |  |  |  | Value |
| 28 | Systematic error due to TRP calculation/quadrature (NOTE 4) |  |  |  | 0.00 |
| 29 | Influence of noise (23.45GHz <= f <= 32.125GHz) |  |  |  | 0.1 |
| 29 | Influence of noise (32.125GHz < f <= 40.8GHz) |  |  |  | 0.3 |
| 30 | Systematic error related to beam peak search (NOTE 5) |  |  |  | 0.5 |
| Total measurement uncertainty |  |  |  |  | Value |
| EIRP Expanded uncertainty (23.45GHz <= f <= 32.125GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 5.17 |
| EIRP Expanded uncertainty (32.125GHz < f <= 40.8GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 5.37 |
| TRP Expanded uncertainty (23.45GHz <= f <= 32.125GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 4.70 |
| TRP Expanded uncertainty (32.125GHz < f <= 40.8GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 4.90 |
| NOTE 1: The quality of quiet zone is the same for EIRP and TRP. Value based on procedure defined in clause D.2 of TR 38.810 for Quiet Zone size less or equal to 30 cm. The ETC QoQZ MU and ETC calibration path losses shall be applied to the NTC test cases if the ETC environment is used for NTC test cases. NOTE 2: The analysis was done only for the case of operating at max output power, in-band, non-CA. NOTE 3: The assessment assumes maximum DUT output power. NOTE 4: This contributor shall only be considered for TRP measurements. NOTE 5: This contributor shall only be considered for EIRP measurements. NOTE 6: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 7: Applies to the system which has a structure of mechanical feed antenna positioning. |  |  |  |  |  |

## B.3.3	Uncertainty budget format and assessment for NFTF

The uncertainty contributions that may impact the overall MU value are listed in Table B.3.3-1.

Table B.3.3-1: Uncertainty contributions for EIRP and TRP measurement

| UID | Description of uncertainty contribution | Details in paragraph |
| --- | --- | --- |
|  | Stage 2: EIRP Near Field Radiation Pattern Measurement and EIRP Near Field DUT power measurement |  |
| 1 | Axis Alignment | # B.2.3.1 |
| 2 | Measurement Distance Uncertainty | B.2.3.2 |
| 3 | Quality of the Quiet Zone | B.2.3.3 |
| 4 | Mismatch | B.2.3.4 |
| 5 | Multiple Reflections: Coupling between Measurement Antenna and DUT | B.2.3.5 |
| 6 | Uncertainty of the RF power measurement equipment | B.2.3.6 |
| 7 | Phase curvature | B.2.3.7 |
| 8 | Amplifier uncertainties | B.2.3.8 |
| 9 | Random uncertainty | B.2.3.9 |
| 10 | Influence of the XPD | B.2.3.10 |
| 11 | NF to FF truncation | B.2.3.11 |
| 12 | Probe Polarization Amplitude and Phase | B.2.3.12 |
| 13 | Probe Array Uniformity (for multi-probe systems only) | B.2.3.13 |
| 14 | Phase Recovery Non-Linearity over signal bandwidth | B.2.3.16 |
| 15 | Probe Pattern Effect | B.2.3.17 |
| 16 | Phase Drift and Noise | B.2.3.20 |
| 17 | Leakage and Crosstalk | B.2.3.25 |
|  | Stage 1: Calibration measurement |  |
| 18 | Mismatch | B.2.3.4 |
| 19 | Amplifier uncertainties | B.2.3.8 |
| 20 | Uncertainty of the Network Analyzer | B.2.3.14 |
| 21 | Uncertainty of the absolute gain of the calibration antenna | B.2.3.15 |
| 22 | Phase centre offset of calibration | B.2.3.18 |
| 23 | Quality of the Quiet Zone for Calibration Process | B.2.3.19 |
| 24 | Mismatch in the connection of the calibration antenna | B.2.3.21 |

The uncertainty assessment table is organized as follows:

-	For the purpose of uncertainty assessment, the radiating antenna aperture of the DUT is denoted as D

-	The uncertainty assessment has been derived for the case of D = [5 cm], f = {22.65GHz, 31.1GHz, 45.1GHz}, P = [maximum output power].

-	The uncertainty assessment for EIRP and TRP is provided in Table B.3.1-2.

Table B.3.3-2: Uncertainty assessment for EIRP and TRP measurement (f=TBD, D=TBD)

| UID | Description of uncertainty contribution | Uncertainty Value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
|  | Stage 2: EIRP Near Field Radiation Pattern Measurement and EIRP Near Field DUT power measurement |  |  |  |  |
| 1 | Axis Alignment |  |  |  |  |
| 2 | Measurement Distance Uncertainty |  |  |  |  |
| 3 | Quality of the Quiet Zone |  |  |  |  |
| 4 | Mismatch |  |  |  |  |
| 5 | Multiple Reflections: Coupling between Measurement Antenna and DUT |  |  |  |  |
| 6 | Uncertainty of the RF power measurement equipment |  |  |  |  |
| 7 | Phase curvature |  |  |  |  |
| 8 | Amplifier uncertainties |  |  |  |  |
| 9 | Random uncertainty |  |  |  |  |
| 10 | Influence of the XPD |  |  |  |  |
| 11 | NF to FF truncation |  |  |  |  |
| 12 | Probe Polarization Amplitude and Phase |  |  |  |  |
| 13 | Probe Array Uniformity (for multi-probe systems only) |  |  |  |  |
| 14 | Phase Recovery Non-Linearity over signal bandwidth |  |  |  |  |
| 15 | Probe Pattern Effect |  |  |  |  |
| 16 | Phase Drift and Noise |  |  |  |  |
| 17 | Leakage and Crosstalk |  |  |  |  |
|  | Stage 1: Calibration measurement |  |  |  |  |
| 18 | Mismatch |  |  |  |  |
| 19 | Amplifier uncertainties |  |  |  |  |
| 20 | Uncertainty of the Network Analyzer |  |  |  |  |
| 21 | Uncertainty of the absolute gain of the calibration antenna |  |  |  |  |
| 22 | Phase centre offset of calibration |  |  |  |  |
| 23 | Quality of the Quiet Zone for Calibration Process |  |  |  |  |
| 24 | Mismatch in the connection of the calibration antenna |  |  |  |  |
| EIRP Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  |
| TRP Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  |
| NOTE 1: The impact of phase variation on EIRP shall be taken into account during final MU definition for the test method. NOTE 2: The quality of quiet zone is different for EIRP and TRP. For TRP, the standard uncertainty is FFS; for EIRP FFS. NOTE 3: The analysis was done only for the case of operating at max output power, in-band, non-CA NOTE 4: The assessment assumes maximum DUT output power. NOTE 5: The Phase Recovery Non-Linearity over signal bandwidth is shall be taken into account during final MU definition for the test method. |  |  |  |  |  |

<< End of change 1>>

<< Start of change 2>>

# B.4	UE maximum output power for modulation / channel bandwidth

Following tables summarize the MU threshold for EIRP measurements for UE maximum output power for modulation / channel bandwidth (a.k.a Maximum Power Reduction/MPR). The origin MU values for different test setups with varies parameters can be found in following clauses.

Table B.4-1: MU threshold for EIRP measurement for UE maximum output power for modulation / channel bandwidth

| Power Class |  | Frequency |  | MBW |  | Power (NOTE2) |  | Threshold MU value for NTC [dB] (NOTE1) | Threshold MU value for ETC [dB] (NOTE1) |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| PC3 | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = Max Output Power –MBR - MPR – T(MPR) | 4.92 | 5.20 |
|  | 32.125GHz < f <= 40.8GHz |  |  | 5.10 | 5.38 |
| PC1 | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = Max Output Power –MBR - MPR – T(MPR) | FFS | FFS |
|  | 32.125GHz < f <= 40.8GHz |  |  | FFS | FFS |
| NOTE 1: Total EIRP Expanded MU for IFF for Quiet Zone size ≤30cm in Table B.4.2-2 for PC2 UEs and B.4.2-x for PC1 UEs. NOTE 2: Max output power level for device with corresponding power class. |  |  |  |  |  |  |  |  |  |  |

## B.4.1	Uncertainty budget format and assessment for DFF

The uncertainty contributions that may impact the overall MU value are listed in Table B.4.1-1.

Table B.4.1-1: Uncertainty contributions for EIRP measurement

| UID | Description of uncertainty contribution | Details in annex |
| --- | --- | --- |
| Same as Table 3.1-1 for EIRP |  |  |

The uncertainty assessment tables are organized as follows:

-	For the purpose of uncertainty assessment, the radiating antenna aperture of the DUT is denoted as D

-	The uncertainty assessment has been derived for the case of D = 5 cm, f = {22.65GHz, 31.1GHz, 45.1GHz}, P = maximum output power – MBR - MPR – T(MPR).

-	The uncertainty assessment for EIRP is provided in Table B.4.1-2.

Table B.4.1-2: Uncertainty assessment for EIRP measurement (f=TBD, D=TBD)

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 to 17 | Same as Stage 2 of Table 3.1-2 for EIRP |  |  |  |  |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 18 to 28 | Same as Stage 1 of Table 3.1-2 for EIRP |  |  |  |  |
|  | Systematic uncertainties |  |  |  | Value |
| 29 | Systematic error due to TRP calculation/quadrature |  |  |  | N/A |
| 30 | Influence of noise |  |  |  |  |
| 31 | Systematic error related to beam peak search |  |  |  | Same as Table 3.1-2 |
| Total measurement uncertainty |  |  |  |  | Value |
| EIRP Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | TBD |
| TRP Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | TBD |
| NOTE 1: The assessment assumes maximum DUT output power – MBR - MPR – T(MPR) |  |  |  |  |  |

## B.4.2	Uncertainty budget format and assessment for IFF

The uncertainty contributions that may impact the overall MU value are listed in Table B.4.2-1.

Table B.4.2-1: Uncertainty contributions for EIRP and TRP measurement

| UID | Description of uncertainty contribution | Details in annex |
| --- | --- | --- |
| Same as Stage 2 of Table 3.2-1 for EIRP |  |  |

The uncertainty assessment tables are organized as follows:

-	For the purpose of uncertainty assessment, the radiating antenna aperture of the DUT is denoted as D

-	The uncertainty assessment has been derived for the case of Quiet Zone size ≤ 30 cm, f = {23.45GHz, 32.125GHz, 40.8GHz}, P = maximum output power – MPR – T(MPR).

-	The uncertainty assessment for EIRP and TRP is provided in Table B.3.2-2 for PC3 UEs and in Table B.4.2-x for PC1 UEs.

Table B.4.2-2: Uncertainty assessment for EIRP measurement (f=23.45GHz, 32.125GHz, 40.8GHz, Quiet Zone size ≤ 30 cm) for PC3 UEs and normal temperature condition

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 to 16 | Same as Stage 2 of Table 3.2-2 for EIRP |  |  |  |  |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 17 to 27 | Same as Stage 1 of Table 3.2-2 for EIRP |  |  |  |  |
|  | Systematic uncertainties |  |  |  | Value |
| 28 | Systematic error due to TRP calculation/quadrature |  |  |  | N/A |
| 29 | Influence of noise (23.45GHz <= f <= 32.125GHz) |  |  |  | 0.13 |
| 29 | Influence of noise (32.125GHz < f <= 40.8GHz) |  |  |  | 0.31 |
| 30 | Systematic error related to beam peak search |  |  |  | Same as Table 3.2-2 |
| Total measurement uncertainty |  |  |  |  | Value |
| EIRP Expanded uncertainty (23.45GHz <= f <= 32.125GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 4.92 |
| EIRP Expanded uncertainty (32.125GHz < f <= 40.8GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 5.10 |
| NOTE 1: The assessment assumes maximum DUT output power – MBR - MPR – T(MPR) |  |  |  |  |  |

Table B.4.2-3: Uncertainty assessment for EIRP measurement (f=23.45GHz, 32.125GHz, 40.8GHz, Quiet Zone size ≤ 30 cm) for PC3 UEs and extreme temperature condition

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 to 16 | Same as Stage 2 of Table 3.2-8 for EIRP |  |  |  |  |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 17 to 27 | Same as Stage 1 of Table 3.2-8 for EIRP |  |  |  |  |
|  | Systematic uncertainties |  |  |  | Value |
| 28 | Systematic error due to TRP calculation/quadrature |  |  |  | N/A |
| 29 | Influence of noise (23.45GHz <= f <= 32.125GHz) |  |  |  | 0.13 |
| 29 | Influence of noise (32.125GHz < f <= 40.8GHz) |  |  |  | 0.31 |
| 30 | Systematic error related to beam peak search |  |  |  | Same as Table 3.2-8 |
| Total measurement uncertainty |  |  |  |  | Value |
| EIRP Expanded uncertainty (23.45GHz <= f <= 32.125GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 5.20 |
| EIRP Expanded uncertainty (32.125GHz < f <= 40.8GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 5.38 |
| NOTE 1: The assessment assumes maximum DUT output power – MBR - MPR – T(MPR) |  |  |  |  |  |

## B.4.3	Uncertainty budget format and assessment for NFTF

FFS

<< End of change 2>>

<< Start of change 3>>

# B.7	Minimum Output power

Following tables summarize the MU threshold for EIRP measurements for Minimum Output Power. The origin MU values for different test setups can be found in following clauses.

Table B.7-1: MU threshold for EIRP measurement for Minimum output power

| Frequency | MBW |  | Power |  | Threshold MU value for NTC [dB] (NOTE1) |  | Threshold MU value For ETC [dB] (NOTE1) |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = Minimum Output Power | PC1: FFS PC2: FFS PC3: 6.15 PC4: FFS | PC1: FFS PC2: FFS PC3: 6.41 PC4: FFS |
| 32.125GHz < f <= 40.8GHz |  |  | PC1: FFS PC2: FFS PC3: 6.15 PC4: FFS | PC1: FFS PC2: FFS PC3: 6.41 PC4: FFS |
| NOTE 1: Total Expanded MU for IFF for Quiet Zone size ≤ 30cm in Table B.7.2-2 |  |  |  |  |  |  |  |  |

## B.7.1	Uncertainty budget format and assessment for DFF

The uncertainty contributions that may impact the overall MU value are listed in Table B.7.1-1.

Table B.7.1-1: Uncertainty contributions for EIRP measurement

| UID | Description of uncertainty contribution | Details in annex |
| --- | --- | --- |
| Stage 2: DUT measurement |  |  |
| 1 | Positioning misalignment | B.2.1.1 |
| 2 | Measure distance uncertainty | B.2.1.2 |
| 3 | Quality of quiet zone | B.2.1.3 |
| 4 | Mismatch | B.2.1.4 |
| 5 | Standing Wave Between the DUT and measurement antenna | B.2.1.5 |
| 6 | Uncertainty of the RF power measurement equipment | B.2.1.6 |
| 7 | Phase curvature | B.2.1.7 |
| 8 | Amplifier uncertainties | B.2.1.8 |
| 9 | Random uncertainty | B.2.1.9 |
| 10 | Influence of the XPD | B.2.1.10 |
| 11 | Insertion Loss Variation | B.2.1.11 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | B.2.1.12 |
| 13 | Influence of beam peak search grid | B.2.1.23 |
| 14 | Multiple measurement antenna uncertainty | B.2.1.25 |
| 15 | DUT repositioning | B.2.1.26 |
| Stage 1: Calibration measurement |  |  |
| 16 | Mismatch | B.2.1.4 |
| 17 | Amplifier uncertainties | B.2.1.8 |
| 18 | Misalignment of positioning System | B.2.1.13 |
| 19 | Uncertainty of the Network Analyzer | B.2.1.14 |
| 20 | Uncertainty of the absolute gain of the calibration antenna | B.2.1.15 |
| 21 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | B.2.1.16 |
| 22 | Phase centre offset of calibration antenna | B.2.1.18 |
| 23 | Quality of quiet zone for calibration process | B.2.1.19 |
| 24 | Standing wave between reference calibration antenna and measurement antenna | B.2.1.20 |
| 25 | Influence of the calibration antenna feed cable | B.2.1.21 |
| 26 | Insertion Loss Variation | B.2.1.11 |
| Systematic uncertainties |  |  |
| 27 | Systematic error related to beam peak search | B.2.1.28 |
| 28 | Influence of noise | B.2.1.27 |

The uncertainty assessment tables are organized as follows:

-	For the purpose of uncertainty assessment, the radiating antenna aperture of the DUT is denoted as D

-	The uncertainty assessment has been derived for the case of D = [5 cm], f = {23.45 GHz, 32.125 GHz, 40.8 GHz}, P = [Minimum output power].

-	The uncertainty assessment for EIRP is provided in Table B.7.1-2.

Table B.7.1-2: Uncertainty assessment for EIRP measurement (f=TBD, D=TBD)

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment |  |  |  |  |
| 2 | Measure distance uncertainty |  |  |  |  |
| 3 | Quality of quiet zone (NOTE 2) |  |  |  |  |
| 4 | Mismatch (NOTE 3) |  |  |  |  |
| 5 | Standing Wave Between the DUT and measurement antenna |  |  |  |  |
| 6 | Uncertainty of the RF power measurement equipment (NOTE 4) |  |  |  |  |
| 7 | Phase curvature |  |  |  |  |
| 8 | Amplifier uncertainties |  |  |  |  |
| 9 | Random uncertainty |  |  |  |  |
| 10 | Influence of the XPD |  |  |  |  |
| 11 | Insertion Loss Variation |  |  |  |  |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) |  |  |  |  |
| 13 | Influence of beam peak search grid (NOTE 6) |  |  |  |  |
| 14 | Multiple measurement antenna uncertainty |  |  |  |  |
| 15 | DUT repositioning |  | Actual | 1 |  |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 16 | Mismatch |  |  |  |  |
| 17 | Amplifier uncertainties |  |  |  |  |
| 18 | Misalignment of positioning System |  |  |  |  |
| 19 | Uncertainty of the Network Analyzer |  |  |  |  |
| 20 | Uncertainty of the absolute gain of the calibration antenna |  |  |  |  |
| 21 | Positioning and pointing misalignment between the reference antenna and the measurement antenna |  |  |  |  |
| 22 | Phase centre offset of calibration antenna |  |  |  |  |
| 23 | Quality of quiet zone for calibration process (NOTE 2) |  |  |  |  |
| 24 | Standing wave between reference calibration antenna and measurement antenna |  |  |  |  |
| 25 | Influence of the calibration antenna feed cable |  |  |  |  |
| 26 | Insertion Loss Variation |  |  |  |  |
| EIRP Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  |
| Systematic uncertainties (NOTE 7) |  |  |  |  | Value |
| 27 | Systematic error related to beam peak search |  |  |  |  |
| 28 | Influence of noise |  |  |  |  |
| Total measurement uncertainty |  |  |  |  |  |
| EIRP total measurement uncertainty [dB] |  |  |  |  |  |
| NOTE 1: The analysis was done only for the case of operating at Minimum output power, in-band, non-CA. NOTE 2: The assessment assumes DUT Minimum output power. NOTE 3: This contributor shall only be considered for EIRP measurements. NOTE 4: Void NOTE 5: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 6: Void. NOTE 7: Void NOTE 8: Value based on procedure defined in Annex D.2 of TR 38.810 for Quiet Zone size less or equal to 30 cm. NOTE 9: Applies to the system which has a structure of mechanical feed antenna positioning. |  |  |  |  |  |

## B.7.2	Uncertainty budget format and assessment for IFF

The uncertainty contributions that may impact the overall MU value are listed in Table B.7.2-1.

Table B.7.2-1: Uncertainty contributions for EIRP measurement

| UID | Description of uncertainty contribution | Details in annex |
| --- | --- | --- |
| Stage 2: DUT measurement |  |  |
| 1 | Positioning misalignment | # B.2.2.1 |
| 2 | Measure distance uncertainty | B.2.2.2 |
| 3 | Quality of Quiet Zone | B.2.2.3 |
| 4 | Mismatch | B.2.2.4 |
| 5 | Standing wave between the DUT and measurement antenna | B.2.2.5 |
| 6 | Uncertainty of the RF power measurement equipment | B.2.2.6 |
| 7 | Phase curvature | B.2.2.7 |
| 8 | Amplifier uncertainties | B.2.2.8 |
| 9 | Random uncertainty | B.2.2.9 |
| 10 | Influence of the XPD | B.2.2.10 |
| 11 | Insertion Loss Variation | B.2.2.11 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | B.2.2.12 |
| 13 | Influence of beam peak search grid | B.2.2.23 |
| 14 | Multiple measurement antenna uncertainty | B.2.2.25 |
| 15 | DUT repositioning | B.2.2.26 |
| Stage 1: Calibration measurement |  |  |
| 16 | Mismatch | B.2.2.4 |
| 17 | Amplifier Uncertainties | B.2.2.8 |
| 18 | Misalignment of positioning System | B.2.2.13 |
| 19 | Uncertainty of the Network Analyzer | B.2.2.14 |
| 20 | Uncertainty of the absolute gain of the calibration antenna | B.2.2.15 |
| 21 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | B.2.2.16 |
| 22 | Phase centre offset of calibration antenna | B.2.2.18 |
| 23 | Quality of quiet zone for calibration process | B.2.2.19 |
| 24 | Standing wave between reference calibration antenna and measurement antenna | B.2.2.20 |
| 25 | Influence of the calibration antenna feed cable | B.2.2.21 |
| 26 | Insertion Loss Variation | B.2.2.11 |
| Systematic uncertainties |  |  |
| 27 | Systematic error related to beam peak search | B.2.2.28 |
| 28 | Influence of noise | B.2.2.27 |

The uncertainty assessment tables are organized as follows:

-	For the purpose of uncertainty assessment, the radiating antenna aperture of the DUT is denoted as D

-	The uncertainty assessment has been derived for the case of Quiet Zone size ≤ 30 cm, f = {23.45GHz, 32.125GHz, 40.8GHz}, P = Minimum output power.

-	The uncertainty assessment for EIRP is provided in Table B.7.2-2 for PC3 UEs and in Table B.7.2-3 for PC1 UEs.

Table B.7.2-2: Uncertainty assessment for EIRP measurement (f=23.45GHz, 32.125GHz, 40.8GHz, Quiet Zone size ≤ 30 cm) for PC3 UEs and normal temperature condition

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment | 0.00 | Normal | 2.00 | 0.00 |
| 2 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 3 | Quality of Quiet Zone (NOTE 8) | 0.6 | Actual | 1.00 | 0.6 |
| 4 | Mismatch (NOTE 1) | 1.30 | Actual | 1.00 | 1.30 |
| 5 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | Uncertainty of the RF power measurement equipment (NOTE 2) | 2.50 | Normal | 2.00 | 1.25 |
| 7 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 8 | Amplifier uncertainties | 2.10 | Normal | 2.00 | 1.05 |
| 9 | Random uncertainty | 0.50 | Normal | 2.00 | 0.25 |
| 10 | Influence of the XPD | 0.01 | U-shaped | 1.41 | 0.00 |
| 11 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 13 | Influence of beam peak search grid (NOTE 3) | 0.00 | Actual | 1 | 0.00 |
| 14 | Multiple measurement antenna uncertainty (NOTE 9) | 0.15 | Actual | 1 | 0.15 |
| 15 | DUT repositioning (NOTE 3) | 0.08 | Rectangular | 1.73 | 0.05 |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 16 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 17 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 18 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 19 | Uncertainty of the Network Analyzer | 1.50 | Normal | 2.00 | 0.75 |
| 20 | Uncertainty of the absolute gain of the calibration antenna | 0.60 | Normal | 2.00 | 0.30 |
| 21 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.01 | Rectangular | 1.73 | 0.00 |
| 22 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 23 | Quality of quiet zone for calibration process (NOTE 8) | 0.4 | Actual | 1.00 | 0.4 |
| 24 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 25 | Influence of the calibration antenna feed cable | 0.14 | Normal | 2.00 | 0.07 |
| 26 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
|  | Systematic uncertainties (NOTE 5) |  |  |  | Value |
| 27 | Systematic error related to beam peak search |  |  |  | 0.5 |
| 28 | Influence of noise (23.45GHz <= f <= 32.125GHz) |  |  |  | 1.0 |
| 29 | Influence of noise (32.125GHz < f <= 40.8GHz) |  |  |  | 1.0 |
| Total measurement uncertainty |  |  |  |  | Value |
| EIRP Expanded uncertainty (23.45GHz <= f <= 32.125GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 6.15 |
| EIRP Expanded uncertainty (32.125GHz < f <= 40.8GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 6.15 |
| NOTE 1: The analysis was done only for the case of operating at Minimum output power, in-band, non-CA. NOTE 2: The assessment assumes DUT Minimum output power. NOTE 3: This contributor shall only be considered for EIRP measurements. NOTE 4: Void NOTE 5: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 6: Void. NOTE 7: Void NOTE 8: Value based on procedure defined in Annex D.2 of TR 38.810 for Quiet Zone size less or equal to 30 cm. NOTE 9: Applies to the system which has a structure of mechanical feed antenna positioning. |  |  |  |  |  |

Table B.7.2-3: Uncertainty assessment for EIRP measurement (f=23.45GHz, 32.125GHz, 40.8GHz, Quiet Zone size ≤ 30 cm) for PC1 UEs and normal temperature condition

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment | 0.02 | Normal | 2.00 | 0.01 |
| 2 | Measure distance uncertainty | FFS | Rectangular | 1.73 | FFS |
| 3 | Quality of Quiet Zone (NOTE 8) | FFS | Actual | 1.00 | FFS |
| 4 | Mismatch (NOTE 1) | FFS | Actual | 1.00 | FFS |
| 5 | Standing wave between the DUT and measurement antenna | FFS | U-shaped | 1.41 | FFS |
| 6 | Uncertainty of the RF power measurement equipment (NOTE 2) | FFS | Normal | 2.00 | FFS |
| 7 | Phase curvature | FFS | U-shaped | 1.41 | FFS |
| 8 | Amplifier uncertainties | FFS | Normal | 2.00 | FFS |
| 9 | Random uncertainty | FFS | Normal | 2.00 | FFS |
| 10 | Influence of the XPD | FFS | U-shaped | 1.41 | FFS |
| 11 | Insertion Loss Variation | FFS | Rectangular | 1.73 | FFS |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | FFS | Actual | 1.00 | FFS |
| 13 | Influence of beam peak search grid (NOTE 3) | 0.00 | Actual | 1 | 0.00 |
| 14 | Multiple measurement antenna uncertainty (NOTE 9) | FFS | Actual | 1 | FFS |
| 15 | DUT repositioning (NOTE 3) | 0.35 | Rectangular | 1.73 | 0.20 |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 16 | Mismatch | FFS | U-shaped | 1.41 | FFS |
| 17 | Amplifier Uncertainties | FFS | Normal | 2.00 | FFS |
| 18 | Misalignment of positioning System | FFS | Normal | 2.00 | FFS |
| 19 | Uncertainty of the Network Analyzer | FFS | Normal | 2.00 | FFS |
| 20 | Uncertainty of the absolute gain of the calibration antenna | FFS | Normal | 2.00 | FFS |
| 21 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | FFS | Rectangular | 1.73 | FFS |
| 22 | Phase centre offset of calibration antenna | FFS | Rectangular | 1.73 | FFS |
| 23 | Quality of quiet zone for calibration process (NOTE 8) | FFS | Actual | 1.00 | FFS |
| 24 | Standing wave between reference calibration antenna and measurement antenna | FFS | U-shaped | 1.41 | FFS |
| 25 | Influence of the calibration antenna feed cable | FFS | Normal | 2.00 | FFS |
| 26 | Insertion Loss Variation | FFS | Rectangular | 1.73 | FFS |
|  | Systematic uncertainties (NOTE 5) |  |  |  | Value |
| 27 | Systematic error related to beam peak search |  |  |  | 0.7 |
| 28 | Influence of noise (23.45GHz <= f <= 32.125GHz) |  |  |  | FFS |
| 29 | Influence of noise (32.125GHz < f <= 40.8GHz) |  |  |  | FFS |
| Total measurement uncertainty |  |  |  |  | Value |
| EIRP Expanded uncertainty (23.45GHz <= f <= 32.125GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | FFS |
| EIRP Expanded uncertainty (32.125GHz < f <= 40.8GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | FFS |
| NOTE 1: The analysis was done only for the case of operating at Minimum output power, in-band, non-CA. NOTE 2: The assessment assumes DUT Minimum output power. NOTE 3: This contributor shall only be considered for EIRP measurements. NOTE 4: Void NOTE 5: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 6: Void. NOTE 7: Void NOTE 8: Value based on procedure defined in Annex D.2 of TR 38.810 for Quiet Zone size less or equal to 30 cm. NOTE 9: Applies to the system which has a structure of mechanical feed antenna positioning. |  |  |  |  |  |

NOTE: MU assessment in Table B.7.2-2 and Table B7.2-3 is based on the following relaxations for 400MHz BW:

Table B.7.2-4: Minimum output power requirement relaxation considered in MU assessment for 400 MHz EIRP measurement (f=23.45GHz, 32.125GHz, 40.8GHz, Quiet Zone size ≤ 30 cm)

| Frequency | Power Class | Relaxation |
| --- | --- | --- |
| 23.45GHz <= f <= 32.125GHz | PC1 | FFS |
|  | PC2, PC3, PC4 | 8.4 dB |
| 32.125GHz < f <= 40.8GHz | PC1 | FFS |
|  | PC2, PC3, PC4 | 13.5 dB |  |

Table B.7.2-5: Uncertainty assessment for EIRP measurement (f=23.45GHz, 32.125GHz, 40.8GHz, Quiet Zone size ≤ 30 cm) for PC3 UEs and extreme temperature condition

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment | 0.00 | Normal | 2.00 | 0.00 |
| 2 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 3 | Quality of Quiet Zone (NOTE 8) | 0.9 | Actual | 1.00 | 0.6 |
| 4 | Mismatch (NOTE 1) | 1.30 | Actual | 1.00 | 1.30 |
| 5 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | Uncertainty of the RF power measurement equipment (NOTE 2) | 2.50 | Normal | 2.00 | 1.25 |
| 7 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 8 | Amplifier uncertainties | 2.10 | Normal | 2.00 | 1.05 |
| 9 | Random uncertainty | 0.50 | Normal | 2.00 | 0.25 |
| 10 | Influence of the XPD | 0.01 | U-shaped | 1.41 | 0.00 |
| 11 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 13 | Influence of beam peak search grid (NOTE 3) | 0.00 | Actual | 1 | 0.00 |
| 14 | Multiple measurement antenna uncertainty (NOTE 9) | 0.15 | Actual | 1 | 0.15 |
| 15 | DUT repositioning (NOTE 3) | 0.08 | Rectangular | 1.73 | 0.05 |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 16 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 17 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 18 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 19 | Uncertainty of the Network Analyzer | 1.50 | Normal | 2.00 | 0.75 |
| 20 | Uncertainty of the absolute gain of the calibration antenna | 0.60 | Normal | 2.00 | 0.30 |
| 21 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.01 | Rectangular | 1.73 | 0.00 |
| 22 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 23 | Quality of quiet zone for calibration process (NOTE 8) | 0.6 | Actual | 1.00 | 0.4 |
| 24 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 25 | Influence of the calibration antenna feed cable | 0.14 | Normal | 2.00 | 0.07 |
| 26 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
|  | Systematic uncertainties (NOTE 5) |  |  |  | Value |
| 27 | Systematic error related to beam peak search |  |  |  | 0.5 |
| 28 | Influence of noise (23.45GHz <= f <= 32.125GHz) |  |  |  | 1.0 |
| 29 | Influence of noise (32.125GHz < f <= 40.8GHz) |  |  |  | 1.0 |
| Total measurement uncertainty |  |  |  |  | Value |
| EIRP Expanded uncertainty (23.45GHz <= f <= 32.125GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 6.41 |
| EIRP Expanded uncertainty (32.125GHz < f <= 40.8GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 6.41 |
| NOTE 1: The analysis was done only for the case of operating at Minimum output power, in-band, non-CA. NOTE 2: The assessment assumes DUT Minimum output power. NOTE 3: This contributor shall only be considered for EIRP measurements. NOTE 4: Void NOTE 5: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 6: Void. NOTE 7: Void NOTE 8: Value based on procedure defined in Annex D.2 of TR 38.810 for Quiet Zone size less or equal to 30 cm. NOTE 9: Applies to the system which has a structure of mechanical feed antenna positioning. |  |  |  |  |  |

NOTE: MU assessment in Table B.7.2-5 is based on the relaxations for 400MHz BW in Table B.7.2-4:

<< End of change 3>>

<< Start of change 4>>

# B.8	Transmit OFF power

Following tables summarize the MU threshold for TRP and EIRP measurements for Transmit OFF power. The origin MU values for different test setups can be found in following clauses.

Table B.8-1: MU threshold for TRP measurement for Transmit OFF power

| Power Class | Frequency | MBW | Power | Threshold MU value (NOTE1) |
| --- | --- | --- | --- | --- |
| PC3 | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = Off Power | 5.49 |
|  | 32.125GHz < f <= 40.8GHz |  |  | N/A |
| NOTE 1: Total TRP Expanded MU for IFF for Quiet Zone size ≤ 30cm in Table B.8.2-2 for PC3 UEs |  |  |  |  |

Table B.8-2: MU threshold for EIRP measurement for Transmit OFF power

| Frequency |  | CBW |  | Power |  | Threshold MU value for NTC [dB] (NOTE1) | Threshold MU value for ETC [dB] (NOTE1) |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 23.45GHz <= f <= 32.125GHz | 50MHz | P = Off Power | 6.15 | 6.41 |
|  | 100MHz |  |  |  |
|  | 200MHz |  |  |  |
|  | 400MHz |  |  |  |
| 32.125GHz < f <= 40.8GHz | 50MHz | P = Off Power | 6.15 | 6.41 |
|  | 100MHz |  |  |  |
|  | 200MHz |  |  |  |
|  | 400MHz |  |  |  |
| NOTE 1: Total Expanded MU for IFF for Quiet Zone size ≤ 30cm in Table B.8.2-4 for PC3 UEs |  |  |  |  |  |  |  |  |

## B.8.1	Uncertainty budget format and assessment for DFF

The uncertainty contributions that may impact the overall MU value are listed in Table B.8.1-1.

Table B.8.1-1: Uncertainty contributions for TRP measurement

| UID | Description of uncertainty contribution | Details in annex |
| --- | --- | --- |
| Stage 2: DUT measurement |  |  |
| 1 | Positioning misalignment | B.2.1.1 |
| 2 | Measure distance uncertainty | B.2.1.2 |
| 3 | Quality of quiet zone | B.2.1.3 |
| 4 | Mismatch | B.2.1.4 |
| 5 | Standing Wave Between the DUT and measurement antenna | B.2.1.5 |
| 6 | Uncertainty of the RF power measurement equipment | B.2.1.6 |
| 7 | Phase curvature | B.2.1.7 |
| 8 | Amplifier uncertainties | B.2.1.8 |
| 9 | Random uncertainty | B.2.1.9 |
| 10 | Influence of the XPD | B.2.1.10 |
| 11 | Insertion Loss Variation | B.2.1.11 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | B.2.1.12 |
| 13 | Influence of TRP measurement grid | B.2.1.22 |
| 14 | Influence of beam peak search grid | B.2.1.23 |
| 15 | Multiple measurement antenna uncertainty | B.2.1.25 |
| 16 | DUT repositioning | B.2.1.26 |
| Stage 1: Calibration measurement |  |  |
| 17 | Mismatch | B.2.1.4 |
| 18 | Amplifier uncertainties | B.2.1.8 |
| 19 | Misalignment of positioning System | B.2.1.13 |
| 20 | Uncertainty of the Network Analyzer | B.2.1.14 |
| 21 | Uncertainty of the absolute gain of the calibration antenna | B.2.1.15 |
| 22 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | B.2.1.16 |
| 23 | Phase centre offset of calibration antenna | B.2.1.18 |
| 24 | Quality of quiet zone for calibration process | B.2.1.19 |
| 25 | Standing wave between reference calibration antenna and measurement antenna | B.2.1.20 |
| 26 | Influence of the calibration antenna feed cable | B.2.1.21 |
| 27 | Insertion Loss Variation | B.2.1.11 |
| Systematic uncertainties |  |  |
| 28 | Systematic error due to TRP calculation/quadrature | B.2.1.24 |
| 29 | Influence of noise | B.2.1.27 |

The uncertainty assessment tables are organized as follows:

-	For the purpose of uncertainty assessment, the radiating antenna aperture of the DUT is denoted as D

-	The uncertainty assessment has been derived for the case of D = [5 cm], f = {23.45 GHz, 32.125 GHz, 40.8 GHz}, P = [Off power].

-	The uncertainty assessment for TRP is provided in Table B.8.1-2.

Table B.8.1-2: Uncertainty assessment for TRP measurement (f=TBD, D=TBD)

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment |  |  |  |  |
| 2 | Measure distance uncertainty |  |  |  |  |
| 3 | Quality of quiet zone (NOTE 2) |  |  |  |  |
| 4 | Mismatch (NOTE 3) |  |  |  |  |
| 5 | Standing Wave Between the DUT and measurement antenna |  |  |  |  |
| 6 | Uncertainty of the RF power measurement equipment (NOTE 4) |  |  |  |  |
| 7 | Phase curvature |  |  |  |  |
| 8 | Amplifier uncertainties |  |  |  |  |
| 9 | Random uncertainty |  |  |  |  |
| 10 | Influence of the XPD |  |  |  |  |
| 11 | Insertion Loss Variation |  |  |  |  |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) |  |  |  |  |
| 13 | Influence of TRP measurement grid (NOTE 5) |  |  |  |  |
| 14 | Influence of beam peak search grid (NOTE 6) |  |  |  |  |
| 15 | Multiple measurement antenna uncertainty |  |  |  |  |
| 16 | DUT repositioning |  | Actual | 1 |  |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 17 | Mismatch |  |  |  |  |
| 18 | Amplifier uncertainties |  |  |  |  |
| 19 | Misalignment of positioning System |  |  |  |  |
| 20 | Uncertainty of the Network Analyzer |  |  |  |  |
| 21 | Uncertainty of the absolute gain of the calibration antenna |  |  |  |  |
| 22 | Positioning and pointing misalignment between the reference antenna and the measurement antenna |  |  |  |  |
| 23 | Phase centre offset of calibration antenna |  |  |  |  |
| 24 | Quality of quiet zone for calibration process (NOTE 2) |  |  |  |  |
| 25 | Standing wave between reference calibration antenna and measurement antenna |  |  |  |  |
| 26 | Influence of the calibration antenna feed cable |  |  |  |  |
| 27 | Insertion Loss Variation |  |  |  |  |
| TRP Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  |
| Systematic uncertainties (NOTE 7) |  |  |  |  | Value |
| 28 | Systematic error due to TRP calculation/quadrature (NOTE 5) |  |  |  |  |
| 29 | Influence of noise |  |  |  |  |
| Total measurement uncertainty |  |  |  |  |  |
| TRP total measurement uncertainty [dB] |  |  |  |  |  |
| NOTE 1: The impact of phase variation on EIRP is FFS. NOTE 2: The quality of quiet zone is different for EIRP and TRP. For TRP, the standard uncertainty is FFS; for EIRP, the standard uncertainty of quiet zone is FFS. NOTE 3: The analysis was done only for the case of operating at max output power, in-band, non-CA. NOTE 4: The assessment assumes maximum DUT output power. NOTE 5: This contributor shall only be considered for TRP measurements. NOTE 6: This contributor shall only be considered for EIRP measurements. NOTE 7: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. |  |  |  |  |  |

## B.8.2	Uncertainty budget format and assessment for IFF

The uncertainty contributions that may impact the overall MU value are listed in Table B.8.2-1.

Table B.8.2-1: Uncertainty contributions for TRP and EIRP measurement

| UID | Description of uncertainty contribution | Details in annex |
| --- | --- | --- |
| Stage 2: DUT measurement |  |  |
| 1 | Positioning misalignment | # B.2.2.1 |
| 2 | Measure distance uncertainty | B.2.2.2 |
| 3 | Quality of Quiet Zone | B.2.2.3 |
| 4 | Mismatch | B.2.2.4 |
| 5 | Standing wave between the DUT and measurement antenna | B.2.2.5 |
| 6 | Uncertainty of the RF power measurement equipment | B.2.2.6 |
| 7 | Phase curvature | B.2.2.7 |
| 8 | Amplifier uncertainties | B.2.2.8 |
| 9 | Random uncertainty | B.2.2.9 |
| 10 | Influence of the XPD | B.2.2.10 |
| 11 | Insertion Loss Variation | B.2.2.11 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | B.2.2.12 |
| 13 | Influence of TRP measurement grid | B.2.2.22 |
| 14 | Influence of beam peak search grid | B.2.2.23 |
| 15 | Multiple measurement antenna uncertainty | B.2.2.25 |
| 16 | DUT repositioning | B.2.2.26 |
| Stage 1: Calibration measurement |  |  |
| 17 | Mismatch | B.2.2.4 |
| 18 | Amplifier Uncertainties | B.2.2.8 |
| 19 | Misalignment of positioning System | B.2.2.13 |
| 20 | Uncertainty of the Network Analyzer | B.2.2.14 |
| 21 | Uncertainty of the absolute gain of the calibration antenna | B.2.2.15 |
| 22 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | B.2.2.16 |
| 23 | Phase centre offset of calibration antenna | B.2.2.18 |
| 24 | Quality of quiet zone for calibration process | B.2.2.19 |
| 25 | Standing wave between reference calibration antenna and measurement antenna | B.2.2.20 |
| 26 | Influence of the calibration antenna feed cable | B.2.2.21 |
| 27 | Insertion Loss Variation | B.2.2.11 |
| Systematic uncertainties |  |  |
| 28 | Systematic error due to TRP calculation/quadrature | B.2.2.24 |
| 29 | Influence of noise | B.2.2.27 |

The uncertainty assessment tables are organized as follows:

-	For the purpose of uncertainty assessment, the radiating antenna aperture of the DUT is denoted as D

-	The uncertainty assessment has been derived for the case of Quiet Zone size ≤ 30 cm, f = {23.45GHz, 32.125GHz, 40.8GHz}, P = Off power.

-	The uncertainty assessment for TRP is provided in Table B.8.2-2.

Table B.8.2-2: Uncertainty assessment for TRP measurement (f=23.45GHz, 32.125GHz, 40.8GHz, Quiet Zone size ≤ 30 cm) for PC3 UEs

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment | 0.00 | Normal | 2.00 | 0.00 |
| 2 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 3 | Quality of Quiet Zone (NOTE 8) | 0.6 | Actual | 1.00 | 0.6 |
| 4 | Mismatch (NOTE 1) | 1.30 | Actual | 1.00 | 1.30 |
| 5 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | Uncertainty of the RF power measurement equipment (NOTE 2) | 2.50 | Normal | 2.00 | 1.25 |
| 7 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 8 | Amplifier uncertainties | 2.10 | Normal | 2.00 | 1.05 |
| 9 | Random uncertainty | 0.50 | Normal | 2.00 | 0.25 |
| 10 | Influence of the XPD | 0.01 | U-shaped | 1.41 | 0.00 |
| 11 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 13 | Influence of TRP measurement grid (NOTE 3) | 0.25 | Actual | 1 | 0.25 |
| 14 | Influence of beam peak search grid | 0.00 | Actual | 1 | 0.00 |
| 15 | Multiple measurement antenna uncertainty (NOTE 9) | 0.15 | Actual | 1 | 0.15 |
| 16 | DUT repositioning | 0.00 | Rectangular | 1.73 | 0.00 |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 17 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 18 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 19 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 20 | Uncertainty of the Network Analyzer | 0.73 | Normal | 2.00 | 0.37 |
| 21 | Uncertainty of the absolute gain of the calibration antenna | 0.60 | Normal | 2.00 | 0.30 |
| 22 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.01 | Rectangular | 1.73 | 0.00 |
| 23 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 24 | Quality of quiet zone for calibration process (NOTE 8) | 0.4 | Actual | 1.00 | 0.4 |
| 25 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 26 | Influence of the calibration antenna feed cable | 0.14 | Normal | 2.00 | 0.07 |
| 27 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
|  | Systematic uncertainties (NOTE 5) |  |  |  | Value |
| 28 | Systematic error due to TRP calculation/quadrature (NOTE 3) |  |  |  | 0.0 |
| 29 | Influence of noise (23.45GHz <= f <= 32.125GHz) |  |  |  | 1.0 |
| 30 | Influence of noise (32.125GHz < f <= 40.8GHz) |  |  |  | N/A |
| Total measurement uncertainty |  |  |  |  | Value |
| TRP Expanded uncertainty (23.45GHz <= f <= 32.125GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 5.49 |
| TRP Expanded uncertainty (32.125GHz < f <= 40.8GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | N/A |
| NOTE 1: The analysis was done only for the case of operating at TX OFF power, in-band, non-CA. NOTE 2: The assessment assumes DUT Off power. NOTE 3: This contributor shall only be considered for TRP measurements. NOTE 4: Void NOTE 5: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 6: Void. NOTE 7: Void NOTE 8: Value based on procedure defined in Annex D.2 of TR 38.810 for Quiet Zone size less or equal to 30 cm. NOTE 9: Applies to the system which has a structure of mechanical feed antenna positioning. |  |  |  |  |  |

NOTE: MU assessment in Table B.8.2-2 for FR2a is based on the relaxation of 30.4dB for 400MHz BW.

Table B.8.2-3: Void

Table B.8.2-4: Uncertainty assessment for EIRP measurement (f=23.45GHz, 32.125GHz, 40.8GHz, Quiet Zone size ≤ 30 cm) for PC3 UEs and normal temperature condition

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment | 0.00 | Normal | 2.00 | 0.00 |
| 2 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 3 | Quality of Quiet Zone (NOTE 8) | 0.6 | Actual | 1.00 | 0.6 |
| 4 | Mismatch (NOTE 1) | 1.30 | Actual | 1.00 | 1.30 |
| 5 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | Uncertainty of the RF power measurement equipment (NOTE 2) | 2.50 | Normal | 2.00 | 1.25 |
| 7 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 8 | Amplifier uncertainties | 2.10 | Normal | 2.00 | 1.05 |
| 9 | Random uncertainty | 0.50 | Normal | 2.00 | 0.25 |
| 10 | Influence of the XPD | 0.01 | U-shaped | 1.41 | 0.00 |
| 11 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 13 | Influence of beam peak search grid (NOTE 3) | 0.00 | Actual | 1 | 0.00 |
| 14 | Multiple measurement antenna uncertainty (NOTE 9) | 0.15 | Actual | 1 | 0.15 |
| 15 | DUT repositioning (NOTE 3) | 0.08 | Rectangular | 1.73 | 0.05 |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 16 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 17 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 18 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 19 | Uncertainty of the Network Analyzer | 1.5 | Normal | 2.00 | 0.75 |
| 20 | Uncertainty of the absolute gain of the calibration antenna | 0.60 | Normal | 2.00 | 0.30 |
| 21 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.01 | Rectangular | 1.73 | 0.00 |
| 22 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 23 | Quality of quiet zone for calibration process (NOTE 8) | 0.4 | Actual | 1.00 | 0.4 |
| 24 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 25 | Influence of the calibration antenna feed cable | 0.14 | Normal | 2.00 | 0.07 |
| 26 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
|  | Systematic uncertainties (NOTE 5) |  |  |  | Value |
| 27 | Systematic error related to beam peak search |  |  |  | 0.5 |
| 28 | Influence of noise (23.45GHz <= f <= 32.125GHz) |  |  |  | 1 |
| 29 | Influence of noise (32.125GHz < f <= 40.8GHz) |  |  |  | 1 |
| Total measurement uncertainty |  |  |  |  | Value |
| EIRP Expanded uncertainty (23.45GHz <= f <= 32.125GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 6.15 |
| EIRP Expanded uncertainty (32.125GHz < f <= 40.8GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 6.15 |
| NOTE 1: The analysis was done only for the case of operating at TX OFF power, in-band, non-CA. NOTE 2: The assessment assumes DUT Off power. NOTE 3: This contributor shall only be considered for EIRP measurements. NOTE 4: Void NOTE 5: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 6: Void. NOTE 7: Void NOTE 8: Value based on procedure defined in Annex D.2 of TR 38.810 for Quiet Zone size less or equal to 30 cm. NOTE 9: Applies to the system which has a structure of mechanical feed antenna positioning. |  |  |  |  |  |

NOTE: MU assessment in Table B.8.2-4 is based on the relaxation of 28.4 dB for FR2a and 30.5 dB for FR2b for 400MHz BW.

Table B.8.2-5: Uncertainty assessment for EIRP measurement (f=23.45GHz, 32.125GHz, 40.8GHz, Quiet Zone size ≤ 30 cm) for PC3 UEs and extreme temperature condition

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment | 0.00 | Normal | 2.00 | 0.00 |
| 2 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 3 | Quality of Quiet Zone (NOTE 8) | 0.9 | Actual | 1.00 | 0.6 |
| 4 | Mismatch (NOTE 1) | 1.30 | Actual | 1.00 | 1.30 |
| 5 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | Uncertainty of the RF power measurement equipment (NOTE 2) | 2.50 | Normal | 2.00 | 1.25 |
| 7 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 8 | Amplifier uncertainties | 2.10 | Normal | 2.00 | 1.05 |
| 9 | Random uncertainty | 0.50 | Normal | 2.00 | 0.25 |
| 10 | Influence of the XPD | 0.01 | U-shaped | 1.41 | 0.00 |
| 11 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 13 | Influence of beam peak search grid (NOTE 3) | 0.00 | Actual | 1 | 0.00 |
| 14 | Multiple measurement antenna uncertainty (NOTE 9) | 0.15 | Actual | 1 | 0.15 |
| 15 | DUT repositioning (NOTE 3) | 0.08 | Rectangular | 1.73 | 0.05 |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 16 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 17 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 18 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 19 | Uncertainty of the Network Analyzer | 1.5 | Normal | 2.00 | 0.75 |
| 20 | Uncertainty of the absolute gain of the calibration antenna | 0.60 | Normal | 2.00 | 0.30 |
| 21 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.01 | Rectangular | 1.73 | 0.00 |
| 22 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 23 | Quality of quiet zone for calibration process (NOTE 8) | 0.6 | Actual | 1.00 | 0.4 |
| 24 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 25 | Influence of the calibration antenna feed cable | 0.14 | Normal | 2.00 | 0.07 |
| 26 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
|  | Systematic uncertainties (NOTE 5) |  |  |  | Value |
| 27 | Systematic error related to beam peak search |  |  |  | 0.5 |
| 28 | Influence of noise (23.45GHz <= f <= 32.125GHz) |  |  |  | 1 |
| 29 | Influence of noise (32.125GHz < f <= 40.8GHz) |  |  |  | 1 |
| Total measurement uncertainty |  |  |  |  | Value |
| EIRP Expanded uncertainty (23.45GHz <= f <= 32.125GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 6.41 |
| EIRP Expanded uncertainty (32.125GHz < f <= 40.8GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 6.41 |
| NOTE 1: The analysis was done only for the case of operating at TX OFF power, in-band, non-CA. NOTE 2: The assessment assumes DUT Off power. NOTE 3: This contributor shall only be considered for EIRP measurements. NOTE 4: Void NOTE 5: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 6: Void. NOTE 7: Void NOTE 8: Value based on procedure defined in Annex D.2 of TR 38.810 for Quiet Zone size less or equal to 30 cm. NOTE 9: Applies to the system which has a structure of mechanical feed antenna positioning. |  |  |  |  |  |

NOTE: MU assessment in Table B.8.2-5 is based on the relaxation of 28.4 dB for FR2a and 30.5 dB for FR2b for 400MHz BW.

<< End of change 4>>

<< Start of change 5>>

# B.10	Frequency error

Following tables summarize the MU threshold for EIRP measurements for Frequency error. The origin MU values for different test setups can be found in following subclauses.

Table B.10-1: MU threshold for beam peak measurement for Frequency error

| Power Class | Frequency | MBW | Power | Threshold MU value for NTC and ETC (NOTE1) |
| --- | --- | --- | --- | --- |
| PC3 | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = Max Output Power | +/- 0.01 ppm |
|  | 32.125GHz < f <= 40.8GHz |  |  | +/- 0.01 ppm |
| PC1 | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = Max Output Power | FFS |
|  | 32.125GHz < f <= 40.8GHz |  |  | FFS |
| NOTE 1: Total Expanded MU for IFF for Quiet Zone size ≤ 30cm in section B.10.2 |  |  |  |  |

## B.10.1	Uncertainty budget format and assessment for DFF

+/- 0.01 ppm

-	The uncertainty assessment has been derived for the case of D = [5 cm], f = {23.45 GHz, 32.125 GHz, 40.8 GHz}, P = [Maximum output power].

## B.10.2	Uncertainty budget format and assessment for IFF

+/- 0.01 ppm

-	The uncertainty assessment has been derived for the case of Quiet zone size ≤ [30 cm], f = {23.45GHz, 32.125GHz, 40.8GHz}, P = [Maximum output power].

<< End of change 5>>

<< Start of change 6>>

# B.17	Adjacent Channel Leakage Ratio

Editor’s Note: MU value analysis for PC1, 2 and 4 are not complete.

Following tables summarize the MU threshold for EIRP measurements for Adjacent Channel Leakage Ratio. The origin MU values for different test setups can be found in following subclauses.

Table B.17-1: MU threshold for TRP measurement for Spectrum emission mask

| Power Class | Frequency | MBW | Power | Threshold MU value (NOTE 1) |
| --- | --- | --- | --- | --- |
| PC3 | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = Max Output Power | TBD |
|  | 32.125GHz < f <= 40.8GHz |  |  | TBD |
| PC1 | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = Max Output Power | TBD |
|  | 32.125GHz < f <= 40.8GHz |  |  | TBD |
| NOTE 1: Total Expanded MU for IFF for a Quiet Zone size ≤ 30 cm in Table B.17.2-2 for PC3 UEs and B.17.2-3 for PC1 UEs. |  |  |  |  |

Table B.17-1B: MU threshold for EIRP measurement for Adjacent Channel Leakage Ratio

| Frequency | CBW | Power | Threshold MU value for NTC and ETC (NOTE 1) |
| --- | --- | --- | --- |
| 23.45GHz <= f <= 32.125GHz | 50MHz | P = Max Output Power | 5.63 |
|  | 100MHz |  | 6.09 |
|  | 200MHz |  | 6.09(NOTE5) |
|  | 400MHz |  | 6.09 (NOTE2) |
| 32.125GHz < f <= 40.8GHz | 50MHz | P = Max Output Power | 6.09 |
|  | 100MHz |  | 6.09(NOTE6) |
|  | 200MHz |  | 6.09(NOTE3) |
|  | 400MHz |  | 6.09(NOTE4) |
| NOTE 1: Total Expanded MU for IFF for a Quiet Zone size ≤ 30 cm in Table B.17.2-2 NOTE 2: This value is based on the relaxation of (MPR – 3.0) dB for MPR > 3.0dB. NOTE 3: Not applicable for MPR > 3.5dB NOTE 4: Not applicable for MPR > 2.0dB NOTE 5: This value is based on the relaxation of (MPR – 5.0) dB for MPR > 5.0dB. NOTE 6: Not applicable for MPR > 5.0dB |  |  |  |

## B.17.1	Uncertainty budget format and assessment for DFF

The uncertainty contributions that may impact the overall MU value are listed in Table B.17.1-1.

Table B.17.1-1: Uncertainty contributions for EIRP measurement

| UID | Description of uncertainty contribution | Details in annex |
| --- | --- | --- |
| Stage 2: DUT measurement |  |  |
| 1 | Positioning misalignment | B.2.1.1 |
| 2 | Measure distance uncertainty | B.2.1.2 |
| 3 | Quality of quiet zone | B.2.1.3 |
| 4 | Mismatch | B.2.1.4 |
| 5 | Standing Wave Between the DUT and measurement antenna | B.2.1.5 |
| 6 | Uncertainty of the RF power measurement equipment | B.2.1.6 |
| 7 | Phase curvature | B.2.1.7 |
| 8 | Amplifier uncertainties | B.2.1.8 |
| 9 | Random uncertainty | B.2.1.9 |
| 10 | Influence of the XPD | B.2.1.10 |
| 11 | Insertion Loss Variation | B.2.1.11 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | B.2.1.12 |
| 13 | Influence of TRP measurement grid | B.2.1.22 |
| 14 | Influence of beam peak search grid | B.2.1.23 |
| 15 | Multiple measurement antenna uncertainty | B.2.1.25 |
| 16 | DUT repositioning | B.2.1.26 |
| Stage 1: Calibration measurement |  |  |
| 17 | Mismatch | B.2.1.4 |
| 18 | Amplifier uncertainties | B.2.1.8 |
| 19 | Misalignment of positioning System | B.2.1.13 |
| 20 | Uncertainty of the Network Analyzer | B.2.1.14 |
| 21 | Uncertainty of the absolute gain of the calibration antenna | B.2.1.15 |
| 22 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | B.2.1.16 |
| 23 | Phase centre offset of calibration antenna | B.2.1.18 |
| 24 | Quality of quiet zone for calibration process | B.2.1.19 |
| 25 | Standing wave between reference calibration antenna and measurement antenna | B.2.1.20 |
| 26 | Influence of the calibration antenna feed cable | B.2.1.21 |
| 27 | Insertion Loss Variation | B.2.1.11 |
| Systematic uncertainties |  |  |
| 28 | Systematic error due to TRP calculation/quadrature | B.2.1.24 |
| 29 | Influence of noise | B.2.1.27 |

The uncertainty assessment tables are organized as follows:

-	For the purpose of uncertainty assessment, the radiating antenna aperture of the DUT is denoted as D

-	The uncertainty assessment has been derived for the case of D = [5 cm], f = {23.45 GHz, 32.125 GHz, 40.8 GHz}, P = [Maximum output power].

-	The uncertainty assessment for TRP is provided in Table B.17.1-2.

Table B.17.1-2: Uncertainty assessment for EIRP measurement (f=TBD, D=TBD)

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment |  |  |  |  |
| 2 | Measure distance uncertainty |  |  |  |  |
| 3 | Quality of quiet zone (NOTE 2) |  |  |  |  |
| 4 | Mismatch (NOTE 3) |  |  |  |  |
| 5 | Standing Wave Between the DUT and measurement antenna |  |  |  |  |
| 6 | Uncertainty of the RF power measurement equipment (NOTE 4) |  |  |  |  |
| 7 | Phase curvature |  |  |  |  |
| 8 | Amplifier uncertainties |  |  |  |  |
| 9 | Random uncertainty |  |  |  |  |
| 10 | Influence of the XPD |  |  |  |  |
| 11 | Insertion Loss Variation |  |  |  |  |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) |  |  |  |  |
| 13 | Influence of TRP measurement grid (NOTE 5) |  |  |  |  |
| 14 | Influence of beam peak search grid (NOTE 6) |  |  |  |  |
| 15 | Multiple measurement antenna uncertainty |  |  |  |  |
| 16 | DUT repositioning |  |  |  |  |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 17 | Mismatch |  |  |  |  |
| 18 | Amplifier uncertainties |  |  |  |  |
| 19 | Misalignment of positioning System |  |  |  |  |
| 20 | Uncertainty of the Network Analyzer |  |  |  |  |
| 21 | Uncertainty of the absolute gain of the calibration antenna |  |  |  |  |
| 22 | Positioning and pointing misalignment between the reference antenna and the measurement antenna |  |  |  |  |
| 23 | Phase centre offset of calibration antenna |  |  |  |  |
| 24 | Quality of quiet zone for calibration process (NOTE 2) |  |  |  |  |
| 25 | Standing wave between reference calibration antenna and measurement antenna |  |  |  |  |
| 26 | Influence of the calibration antenna feed cable |  |  |  |  |
| 27 | Insertion Loss Variation |  |  |  |  |
| EIRP Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  |
| Systematic uncertainties (NOTE 7) |  |  |  |  | Value |
| 28 | Systematic error due to TRP calculation/quadrature (NOTE 5) |  |  |  |  |
| 29 | Influence of noise |  |  |  |  |
| Total measurement uncertainty |  |  |  |  |  |
| TRP total measurement uncertainty [dB] |  |  |  |  |  |
| NOTE 1: The impact of phase variation on EIRP is FFS. NOTE 2: The quality of quiet zone is different for EIRP and TRP. For TRP, the standard uncertainty is FFS; for EIRP, the standard uncertainty of quiet zone is FFS. NOTE 3: The analysis was done only for the case of operating at max output power, in-band, non-CA. NOTE 4: The assessment assumes maximum DUT output power. NOTE 5: This contributor shall only be considered for TRP measurements. NOTE 6: This contributor shall only be considered for EIRP measurements. NOTE 7: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 8: Void |  |  |  |  |  |

## B.17.2	Uncertainty budget format and assessment for IFF

The uncertainty contributions that may impact the overall MU value are listed in Table B.17.2-1.

Table B.17.2-1: Uncertainty contributions for EIRP measurement

| UID | Description of uncertainty contribution | Details in clause |
| --- | --- | --- |
| Stage 2: DUT measurement |  |  |
| 1 | Positioning misalignment | # B.2.2.1 |
| 2 | Measure distance uncertainty | B.2.2.2 |
| 3 | Quality of Quiet Zone | B.2.2.3 |
| 4 | Mismatch | B.2.2.4 |
| 5 | Standing wave between the DUT and measurement antenna | B.2.2.5 |
| 6 | Uncertainty of the RF power measurement equipment | B.2.2.6 |
| 7 | Phase curvature | B.2.2.7 |
| 8 | Amplifier uncertainties | B.2.2.8 |
| 9 | Random uncertainty | B.2.2.9 |
| 10 | Influence of the XPD | B.2.2.10 |
| 11 | Insertion Loss Variation | B.2.2.11 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | B.2.2.12 |
| 13 | Influence of TRP measurement grid | B.2.2.22 |
| 14 | Influence of beam peak search grid | B.2.2.23 |
| 15 | Multiple measurement antenna uncertainty | B.2.2.25 |
| 16 | DUT repositioning | B.2.2.26 |
| Stage 1: Calibration measurement |  |  |
| 17 | Mismatch | B.2.2.4 |
| 18 | Amplifier Uncertainties | B.2.2.8 |
| 19 | Misalignment of positioning System | B.2.2.13 |
| 20 | Uncertainty of the Network Analyzer | B.2.2.14 |
| 21 | Uncertainty of the absolute gain of the calibration antenna | B.2.2.15 |
| 22 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | B.2.2.16 |
| 23 | Phase centre offset of calibration antenna | B.2.2.18 |
| 24 | Quality of quiet zone for calibration process | B.2.2.19 |
| 25 | Standing wave between reference calibration antenna and measurement antenna | B.2.2.20 |
| 26 | Influence of the calibration antenna feed cable | B.2.2.21 |
| 27 | Insertion Loss Variation | B.2.2.11 |
| Systematic uncertainties |  |  |
| 28 | Systematic error due to TRP calculation/quadrature | B.2.2.24 |
| 29 | Influence of noise | B.2.2.27 |

The uncertainty assessment tables are organized as follows:

-	For the purpose of uncertainty assessment, the radiating antenna aperture of the DUT is denoted as D

-	The uncertainty assessment has been derived for the case of Quiet Zone size ≤ 30 cm, f = {23.45GHz, 32.125GHz, 40.8GHz}, P = Maximum output power - MPR – MBR(Multi-band relaxation).

-	The uncertainty assessment for EIRP is provided in Table B.17.2-2 for PC3 UEs and Table B.17.2-3 for PC1 UEs.

Table B.17.2-2: Uncertainty assessment for EIRP measurement (f=23.45GHz, 32.125GHz, 40.8GHz, Quiet Zone size ≤ 30 cm) for PC3 UEs and normal and extreme temperature condition

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment | 0.00 | Normal | 2.00 | 0.00 |
| 2 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 3 | Quality of Quiet Zone (NOTE 10) | 0.52 | Actual | 1.00 | 0.52 |
| 4 | Mismatch (NOTE 2) | 1.84 | Actual | 1.00 | 1.84 |
| 5 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | Uncertainty of the RF power measurement equipment (NOTE 3, 7) | 2.16 | Normal | 2.00 | 1.08 |
| 7 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 8 | Amplifier uncertainties | 2.1 | Normal | 2.00 | 1.05 |
| 9 | Random uncertainty | 0.50 | Normal | 2.00 | 0.25 |
| 10 | Influence of the XPD | 0.00 | U-shaped | 1.41 | 0.00 |
| 11 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 13 | Influence of TRP measurement grid (NOTE 4) | 0.0 | Actual | 1 | 0.0 |
| 14 | Influence of beam peak search grid (NOTE 5) | 0.00 | Actual | 1 | 0.00 |
| 15 | Multiple measurement antenna uncertainty (NOTE 9) | 0.0 | Actual | 1 | 0.0 |
| 16 | DUT repositioning (NOTE 4) | 0.00 | Rectangular | 1.73 | 0.00 |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 17 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 18 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 19 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 20 | Uncertainty of the Network Analyzer | 1.5 | Normal | 2.00 | 0.75 |
| 21 | Uncertainty of the absolute gain of the calibration antenna | 0.60 | Normal | 2.00 | 0.30 |
| 22 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 23 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 24 | Quality of quiet zone for calibration process (NOTE 10) | 0.32 | Actual | 1.00 | 0.32 |
| 25 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 26 | Influence of the calibration antenna feed cable | 0.00 | Normal | 2.00 | 0.00 |
| 27 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| EIRP Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 5.09 |
|  | Systematic uncertainties (NOTE 6) |  |  |  | Value |
| 28 | Systematic error due to TRP calculation/quadrature (NOTE 4) |  |  |  | 0.00 |
| 29 | Influence of noise |  |  |  | Table B.17.2-4 |
| 30 | Beam peak search |  |  |  | 0.00 |
| Total measurement uncertainty |  |  |  |  | Value |
| EIRP total measurement uncertainty [dB] |  |  |  |  | 5.09 + Influence of Noise |
| NOTE 1: Void NOTE 2: The analysis was done only for the case of operating at max output power – MPR – MBR(Multi-band relaxation)., in-band, non-CA. NOTE 3: The assessment assumes maximum DUT output power – MPR – MBR(Multi-band relaxation). NOTE 4: This contributor shall only be considered for TRP measurements. NOTE 5: Void NOTE 6: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 7: Void NOTE 8: Void NOTE 9: Applies to the system which has a structure of mechanical feed antenna positioning. NOTE 10: Defined as fixed value MU contributor. |  |  |  |  |  |

Table B.17.2-3: Uncertainty assessment for EIRP measurement (f=23.45GHz, 32.125GHz, 40.8GHz, Quiet Zone size ≤ 30 cm) for PC1 UEs

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment | 0.02 | Normal | 2.00 | 0.01 |
| 2 | Measure distance uncertainty | FFS | Rectangular | 1.73 | FFS |
| 3 | Quality of Quiet Zone (NOTE 10) | FFS | Actual | 1.00 | FFS |
| 4 | Mismatch (NOTE 2, NOTE 7) | FFS | Actual | 1.00 | FFS |
| 5 | Standing wave between the DUT and measurement antenna | FFS | U-shaped | 1.41 | FFS |
| 6 | Uncertainty of the RF power measurement equipment (NOTE 3, 7) | FFS | Normal | 2.00 | FFS |
| 7 | Phase curvature | FFS | U-shaped | 1.41 | FFS |
| 8 | Amplifier uncertainties | FFS | Normal | 2.00 | FFS |
| 9 | Random uncertainty | FFS | Normal | 2.00 | FFS |
| 10 | Influence of the XPD | FFS | U-shaped | 1.41 | FFS |
| 11 | Insertion Loss Variation | FFS | Rectangular | 1.73 | FFS |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | FFS | Actual | 1.00 | FFS |
| 13 | Influence of TRP measurement grid (NOTE 4) | 0.00 | Actual | 1 | 0.00 |
| 14 | Influence of beam peak search grid (NOTE 5) | 0.00 | Actual | 1 | 0.00 |
| 15 | Multiple measurement antenna uncertainty (NOTE 9) | FFS | Actual | 1 | FFS |
| 16 | DUT repositioning (NOTE 4) | 0.00 | Rectangular | 1.73 | 0.00 |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 17 | Mismatch | FFS | U-shaped | 1.41 | FFS |
| 18 | Amplifier Uncertainties | FFS | Normal | 2.00 | FFS |
| 19 | Misalignment of positioning System | FFS | Normal | 2.00 | FFS |
| 20 | Uncertainty of the Network Analyzer | FFS | Normal | 2.00 | FFS |
| 21 | Uncertainty of the absolute gain of the calibration antenna | FFS | Normal | 2.00 | FFS |
| 22 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | FFS | Rectangular | 1.73 | FFS |
| 23 | Phase centre offset of calibration antenna | FFS | Rectangular | 1.73 | FFS |
| 24 | Quality of quiet zone for calibration process (NOTE 10) | FFS | Actual | 1.00 | FFS |
| 25 | Standing wave between reference calibration antenna and measurement antenna | FFS | U-shaped | 1.41 | FFS |
| 26 | Influence of the calibration antenna feed cable | FFS | Normal | 2.00 | FFS |
| 27 | Insertion Loss Variation | FFS | Rectangular | 1.73 | FFS |
| TRP Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | FFS |
|  | Systematic uncertainties (NOTE 6) |  |  |  | Value |
| 28 | Systematic error due to TRP calculation/quadrature (NOTE 4) |  |  |  | 0.00 |
| 29 | Influence of noise |  |  |  | FFS |
| 30 | Beam peak search |  |  |  | FFS |
| Total measurement uncertainty |  |  |  |  | Value |
| TRP total measurement uncertainty [dB] |  |  |  |  | FFS |
| NOTE 1: Void NOTE 2: The analysis was done only for the case of operating at max output power, in-band, non-CA. NOTE 3: The assessment assumes maximum DUT output power. NOTE 4: This contributor shall only be considered for TRP measurements. NOTE 5: Void NOTE 6: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 7: Values extracted from TR 38.810 v2.6.1 in square brackets pending for further analysis. NOTE 8: Void. NOTE 9: Applies to the system which has a structure of mechanical feed antenna positioning. NOTE 10: Defined as fixed value MU contributor. |  |  |  |  |  |

Table B.17.2-4: Influence of noise measurement (f=23.45GHz, 32.125GHz, 40.8GHz, Quiet Zone size ≤ 30 cm) for PC3 UEs

|  | FR2a | FR2b |
| --- | --- | --- |
| ChBW (50MHz) | 0.54 | 1.0 |
| ChBW (100MHz) | 1.0 | 1.0 (NOTE5) |
| ChBW (200MHz) | 1.0 (NOTE 4) | 1.0 (NOTE2) |
| ChBW (400MHz) | 1.0 (NOTE1) | 1.0 (NOTE3) |
| NOTE 1: This value is based on the relaxation of (MPR – 3.0) dB for MPR > 3.0dB. NOTE 2: Not applicable for MPR > 3.5dB NOTE 3 : Not applicable for MPR > 2.0dB NOTE 4: This value is based on the relaxation of (MPR – 5.0) dB for MPR > 5.0dB. NOTE 5: Not applicable for MPR > 5.0dB |  |  |

<< End of change 6>>