3GPP TSG-RAN5 Meeting #101	R5-236070r1

# Chicago, US, 13th Nov 2023 – 17th Nov 2023

| CR-Form-v12.2 |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CHANGE REQUEST |  |  |  |  |  |  |  |  |
|  | 38.903 | CR | 0599 | rev | - | Current version: | 18.0.0 |  |
| For HELP on using this form: comprehensive instructions can be found at http://www.3gpp.org/Change-Requests. |  |  |  |  |  |  |  |  |

| Proposed change affects: | UICC apps |  | ME |  | Radio Access Network |  | Core Network |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |

| Title: | FR2c MU definition in 38.903 |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Source to WG: | Keysight Technologies UK Ltd |  |  |  |  |  |  |  |  |  |
| Source to TSG: | R5 |  |  |  |  |  |  |  |  |  |
| Work item code: | TEI15_Test, 5GS_NR_LTE-UEConTest |  |  |  |  |  | Date: |  |  | 2023-10-27 |
| Category: | F |  |  |  |  |  | Release: |  |  | Rel-18 |
|  | Use one of the following categories: F (correction) A (mirror corresponding to a change in an earlier release) B (addition of feature), C (functional modification of feature) D (editorial modification) Detailed explanations of the above categories can be found in 3GPP TR 21.900. |  |  |  |  |  |  |  | Use one of the following releases: Rel-8 (Release 8) Rel-9 (Release 9) Rel-10 (Release 10) Rel-11 (Release 11) … Rel-16 (Release 16) Rel-17 (Release 17) Rel-18 (Release 18) Rel-19 (Release 19) |  |
| Reason for change: |  | In R5-236063, progress has been made for MU analysis in FR2c. |  |  |  |  |  |  |  |  |
| Summary of change: |  | Defined QoQZ MU factors for EIRP/EIS in ETC for n259. Defined MU for MOP-EIRP, MOP-TRP and REFSENS-EIS in ETC. Defined MU for OBW. |  |  |  |  |  |  |  |  |
| Consequences if not approved: |  | Test specification will remain incomplete for FR2c. |  |  |  |  |  |  |  |  |
| Clauses affected: |  | B.2.2.3, B.2.2.19, B.3, B.15, B.16, B.19 |  |  |  |  |  |  |  |  |
|  |  | Y | N |  |  |  |  |  |  |  |
| Other specs |  |  | X | Other core specifications |  |  |  | TS/TR ... CR ... |  |  |
| affected: |  |  | X | Test specifications |  |  |  | TS/TR ... CR ... |  |  |
| (show related CRs) |  |  | X | O&M Specifications |  |  |  | TS/TR ... CR ... |  |  |
| Other comments: |  |  |  |  |  |  |  |  |  |  |
| This CR's revision history: |  | Revision 1: -r1: -Values revised taking into account the new proposals made int this meeting for the RF power measurement equipment uncertainty and the gNB emulator uncertainty. -Set to TBD the final values for SEM and OBW because agreement is to be postponed to RAN5#102 meeting. |  |  |  |  |  |  |  |  |

## <<< START OF CHANGES >>>

### B.2.2.3	Quality of Quiet Zone

See B.2.1.3.

The uncertainty value of quality of quiet zone is estimated as below table and used across clause B.

Table B.2.2.3-1: Uncertainty value for quality of quiet zone for IFF

| QZ size | Power class | Condition | Test case | Frequency range | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| All | All | All | ACLR (relative measurement) | FR2a, FR2b, FR2c | 0.52 | Actual | 1.00 | 0.52 |
| ≤30cm | PC1, PC3, PC5, PC6 | NC | NOTE1 | FR2a, FR2b | 0.6 | Actual | 1.00 | 0.6 |
|  |  |  |  | FR2c | 0.7 | Actual | 1.00 | 0.7 |
|  |  |  | SE | 6GHz to 12.75GHz | 0.7 | Actual | 1.00 | 0.7 |
|  |  |  |  | 12.75GHz to 23.45GHz | 0.6 | Actual | 1.00 | 0.6 |
|  |  |  |  | 23.45GHz to 40.8GHz | 0.6 | Actual | 1.00 | 0.6 |
|  |  |  |  | 40.8GHz to 66GHz | 0.6 | Actual | 1.00 | 0.6 |
|  |  |  |  | 66GHz to 80GHz | 0.6 | Actual | 1.00 | 0.6 |
|  |  | ETC | NOTE1 | FR2a, FR2b, FR2c (NOTE 3) | 0.9 | Actual | 1.00 | 0.9 |
| 40cm | PC1, PC3, PC5, PC6 | NC | NOTE1 | FR2a, FR2b | 0.7 (NOTE 2) | Actual | 1.00 | 0.7 |
|  |  |  |  | FR2c | 0.8 (NOTE 2) | Actual | 1.00 | 0.8 |
|  |  |  | SE | 6GHz to 12.75GHz | TBD | Actual | 1.00 | TBD |
|  |  |  |  | 12.75GHz to 23.45GHz | TBD | Actual | 1.00 | TBD |
|  |  |  |  | 23.45GHz to 40.8GHz | TBD | Actual | 1.00 | TBD |
|  |  |  |  | 40.8GHz to 66GHz | TBD | Actual | 1.00 | TBD |
|  |  |  |  | 66GHz to 80GHz | TBD | Actual | 1.00 | TBD |
|  |  | ETC | NOTE1 | FR2a, FR2b | TBD | Actual | 1.00 | TBD |
| NOTE 1: The uncertainty in current row applies to maximum output power with EIRP and TRP, EIRP spherical coverage, MPR, configured output power with power boost, minimum output power, transmit OFF power, spectrum emission mask, reference sensitivity, adjacent selectivity, in-band blocking. NOTE 2: For Uncertainty assessment in tables across clause B and MTSU calculation, QoQZ uncertainty value for QZ ≤ 30cm is used also for 40cm. NOTE 3: For FR2c in ETC, QoQZ uncertainty value only defined for EIRP/EIS. |  |  |  |  |  |  |  |  |

## <<< Skip unchanged sections >>>

### B.2.2.19	Quality of the Quiet Zone for Calibration Process

See B.2.1.19.

The uncertainty value of quality of quiet zone for calibration process is estimated as below table and used across clause B.

Table B.2.2.19-1: Uncertainty value for quiet zone for calibration process for IFF

| QZ size | Power class | Condition | Test case | Frequency range | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| All | All | All | ACLR (relative measurement) | FR2a, FR2b, FR2c | 0.32 | Actual | 1.00 | 0.32 |
| ≤ 40cm | PC1, PC3, PC5, PC6 | NC | NOTE1 | FR2a, FR2b | 0.4 | Actual | 1.00 | 0.4 |
|  |  |  |  | FR2c | 0.5 | Actual | 1.00 | 0.5 |
|  |  |  | SE | 6GHz to 12.75GHz | 0.7 | Actual | 1.00 | 0.7 |
|  |  |  |  | 12.75GHz to 23.45GHz | 0.6 | Actual | 1.00 | 0.6 |
|  |  |  |  | 23.45GHz to 40.8GHz | 0.6 | Actual | 1.00 | 0.6 |
|  |  |  |  | 40.8GHz to 66GHz | 0.6 | Actual | 1.00 | 0.6 |
|  |  |  |  | 66GHz to 80GHz | 0.6 | Actual | 1.00 | 0.6 |
|  |  | ETC | NOTE1 | FR2a, FR2b, FR2c (NOTE 2) | 0.6 | Actual | 1.00 | 0.6 |
| NOTE 1: The uncertainty in current row applies to maximum output power with EIRP and TRP, EIRP spherical coverage, MPR, configured output power with power boost, minimum output power, transmit OFF power, spectrum emission mask, reference sensitivity, adjacent selectivity, in-band blocking. NOTE 2: For FR2c in ETC, QoQZ uncertainty value only defined for EIRP/EIS. |  |  |  |  |  |  |  |  |

## <<< Skip unchanged sections >>>

# B.3	UE maximum output power

Following tables summarize the MU threshold for EIRP and TRP measurements for UE maximum output power. The origin MU values for different test setups with varies parameters can be found in following clauses.

Table B.3-1: MU threshold for EIRP measurement for UE maximum output power

| Power Class | Frequency | MBW | Power (NOTE2) | Threshold MU value for NTC [dB] (NOTE1) | Threshold MU value for ETC [dB] (NOTE1) |
| --- | --- | --- | --- | --- | --- |
| PC3 | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = Max Output Power | 5.08 | 5.35 |
|  | 32.125GHz < f <= 40.8GHz |  |  | 5.28 | 5.55 |
|  | 40.8GHz < f <= 44.3GHz |  |  | 6.64 | 6.78 |
| PC1 | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = Max Output Power | 5.33 | 5.60 |
|  | 32.125GHz < f <= 40.8GHz |  |  | 5.40 | 5.67 |
| PC5 | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = Max Output Power | 5.33 | 5.60 |
| NOTE 1: Total EIRP Expanded MU for IFF for Quiet Zone size ≤30cm in Table B.3.2-2 for PC3 UEs (NTC), in Table B.3.2-8 for PC3 UEs (ETC) and B.3.2-6 (NTC) and B.3.2-9 (ETC) for PC1 and PC5 UEs. NOTE 2: Max output power level for device with corresponding power class. |  |  |  |  |  |

Table B.3-2: MU threshold for TRP measurement for UE maximum output power

| Power Class | Frequency | MBW | Power (NOTE2) | Threshold MU value for NTC [dB] (NOTE 1) | Threshold MU value for ETC [dB] (NOTE 1) |
| --- | --- | --- | --- | --- | --- |
| PC3 | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = Max Output Power | 4.61 | 4.87 |
|  | 32.125GHz < f <= 40.8GHz |  |  | 4.81 | 5.07 |
|  | 40.8GHz < f <= 44.3GHz |  |  | 6.16 | TBD |
| PC1 | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = Max Output Power | 4.64 | 4.90 |
|  | 32.125GHz < f <= 40.8GHz |  |  | 4.78 | 5.04 |
| PC5 | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = Max Output Power | 4.64 | 4.90 |
| NOTE 1: Total TRP Expanded MU for IFF for Quiet Zone size ≤ 30cm in Table B.3.2-2 for PC3 UEs (NTC), in Table B.3.2-8 for PC3 UEs (ETC) and B.3.2-6 (NTC) and B.3.2-9 (ETC) for PC1 and PC5 UEs. NOTE 2: Max output power level for device with corresponding power class. |  |  |  |  |  |

Table B.3-3: MU threshold for Spherical coverage measurement for UE maximum output power

| Power Class | Frequency | MBW | Power | Threshold MU value (NOTE 1) |
| --- | --- | --- | --- | --- |
| PC3 | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = TBD | 4.78 |
|  | 32.125GHz < f <= 40.8GHz |  |  | 5.38 |
| PC1 | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = TBD | 4.69 |
|  | 32.125GHz < f <= 40.8GHz |  |  | 4.84 |
| PC5 | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = TBD | 4.69 |
| NOTE 1: Total Spherical coverage Expanded MU for IFF for Quiet Zone size ≤ 30cm in Tables B.3.2-4 for PC3 UEs and B.3.2-7 for PC1 and PC5 UEs |  |  |  |  |

## B.3.1	Uncertainty budget format and assessment for DFF

The uncertainty contributions that may impact the overall MU value are listed in Table B.3.1-1.

Table B.3.1-1: Uncertainty contributions for EIRP and TRP measurement

| UID | Description of uncertainty contribution | Details in annex |
| --- | --- | --- |
| Stage 2: DUT measurement |  |  |
| 1 | Positioning misalignment | B.2.1.1 |
| 2 | Measure distance uncertainty | B.2.1.2 |
| 3 | Quality of quiet zone | B.2.1.3 |
| 4 | Mismatch | B.2.1.4 |
| 5 | Standing Wave Between the DUT and measurement antenna | B.2.1.5 |
| 6 | Uncertainty of the RF power measurement equipment | B.2.1.6 |
| 7 | Phase curvature | B.2.1.7 |
| 8 | Amplifier uncertainties | B.2.1.8 |
| 9 | Random uncertainty | B.2.1.9 |
| 10 | Influence of the XPD | B.2.1.10 |
| 11 | Insertion Loss Variation | B.2.1.11 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | B.2.1.12 |
| 13 | Influence of TRP measurement grid | B.2.1.22 |
| 14 | Influence of beam peak search grid | B.2.1.23 |
| 15 | Multiple measurement antenna uncertainty | B.2.1.25 |
| 16 | DUT repositioning | B.2.1.26 |
| 17 | Influence of spherical coverage grid | B.2.1.29 |
| Stage 1: Calibration measurement |  |  |
| 18 | Mismatch | B.2.1.4 |
| 19 | Amplifier uncertainties | B.2.1.8 |
| 20 | Misalignment of positioning System | B.2.1.13 |
| 21 | Uncertainty of the Network Analyzer | B.2.1.14 |
| 22 | Uncertainty of the absolute gain of the calibration antenna | B.2.1.15 |
| 23 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | B.2.1.16 |
| 24 | Phase centre offset of calibration antenna | B.2.1.18 |
| 25 | Quality of quiet zone for calibration process | B.2.1.19 |
| 26 | Standing wave between reference calibration antenna and measurement antenna | B.2.1.20 |
| 27 | Influence of the calibration antenna feed cable | B.2.1.21 |
| 28 | Insertion Loss Variation | B.2.1.11 |
| Systematic uncertainties |  |  |
| 29 | Systematic error due to TRP calculation/quadrature | B.2.1.24 |
| 30 | Influence of noise | B.2.1.27 |
| 31 | Systematic error related to beam peak search | B.2.1.28 |

The uncertainty assessment tables are organized as follows:

-	For the purpose of uncertainty assessment, the radiating antenna aperture of the DUT is denoted as D

-	The uncertainty assessment has been derived for the case of D = [5 cm], f = {22.65GHz, 31.1GHz, 45.1GHz}, P = [maximum output power].

-	The uncertainty assessment for EIRP and TRP is provided in Table B.3.1-2.

Table B.3.1-2: Uncertainty assessment for EIRP and TRP measurement (f=TBD, D=TBD)

| UID |  | Uncertainty source |  | Uncertainty value |  | Distribution of the probability |  | Divisor |  | Standard uncertainty (σ) [dB] |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |  |  |  |  |
| 1 |  | Positioning misalignment |  |  |  |  |  |  |  |  |  |
| 2 |  | Measure distance uncertainty |  |  |  |  |  |  |  |  |  |
| 3 |  | Quality of quiet zone (NOTE 2) |  |  |  |  |  |  |  |  |  |
| 4 |  | Mismatch (NOTE 3) |  |  |  |  |  |  |  |  |  |
| 5 |  | Standing Wave Between the DUT and measurement antenna |  |  |  |  |  |  |  |  |  |
| 6 |  | Uncertainty of the RF power measurement equipment (NOTE 4) |  |  |  |  |  |  |  |  |  |
| 7 |  | Phase curvature |  |  |  |  |  |  |  |  |  |
| 8 |  | Amplifier uncertainties |  |  |  |  |  |  |  |  |  |
| 9 |  | Random uncertainty |  |  |  |  |  |  |  |  |  |
| 10 |  | Influence of the XPD |  |  |  |  |  |  |  |  |  |
| 11 |  | Insertion Loss Variation |  |  |  |  |  |  |  |  |  |
| 12 |  | RF leakage (from measurement antenna to the receiver/transmitter) |  |  |  |  |  |  |  |  |  |
| 13 |  | Influence of TRP measurement grid (NOTE 5) |  | 0.25 |  | Actual |  | 1 |  | 0.25 |  |
| 14 |  | Influence of beam peak search grid (NOTE 6) |  | 0.0 |  | Actual |  | 1 |  | 0.0 |  |
| 15 |  | Multiple measurement antenna uncertainty |  |  |  |  |  |  |  |  |  |
| 16 |  | DUT repositioning |  |  |  |  |  |  |  |  |  |
| 17 |  | Influence of spherical coverage grid (NOTE 8) |  | 0.12 |  | Actual |  | 1 |  | 0.12 |  |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |  |  |  |  |
| 18 |  | Mismatch |  |  |  |  |  |  |  |  |  |
| 19 |  | Amplifier uncertainties |  |  |  |  |  |  |  |  |  |
| 20 |  | Misalignment of positioning System |  |  |  |  |  |  |  |  |  |
| 21 |  | Uncertainty of the Network Analyzer |  |  |  |  |  |  |  |  |  |
| 22 |  | Uncertainty of the absolute gain of the calibration antenna |  |  |  |  |  |  |  |  |  |
| 23 |  | Positioning and pointing misalignment between the reference antenna and the measurement antenna |  |  |  |  |  |  |  |  |  |
| 24 |  | Phase centre offset of calibration antenna |  |  |  |  |  |  |  |  |  |
| 25 |  | Quality of quiet zone for calibration process (NOTE 2) |  |  |  |  |  |  |  |  |  |
| 26 |  | Standing wave between reference calibration antenna and measurement antenna |  |  |  |  |  |  |  |  |  |
| 27 |  | Influence of the calibration antenna feed cable |  |  |  |  |  |  |  |  |  |
| 28 |  | Insertion Loss Variation |  |  |  |  |  |  |  |  |  |
|  |  | Systematic uncertainties (NOTE 7) |  |  |  |  |  |  |  | Value |  |
| 29 |  | Systematic error due to TRP calculation/quadrature (NOTE 5) |  |  |  |  |  |  |  | 0.00 |  |
| 30 |  | Influence of noise |  |  |  |  |  |  |  |  |  |
| 31 |  | Systematic error related to beam peak search (NOTE 6) |  |  |  |  |  |  |  | 0.5 |  |
| Total measurement uncertainty |  |  |  |  |  |  |  |  |  | Value |  |
| EIRP Expanded uncertainty (1.96σ – confidence interval of 95 %) [dB] |  |  |  |  |  |  |  |  |  | TBD |  |
| TRP Expanded uncertainty (1.96σ – confidence interval of 95 %) [dB] |  |  |  |  |  |  |  |  |  | TBD |  |
| NOTE 1: The impact of phase variation on EIRP shall be taken into account during final MU definition for the test method. NOTE 2: The quality of quiet zone is different for EIRP and TRP in normal temperature conditions. For TRP, the standard uncertainty is FFS; for EIRP, the standard uncertainty of quiet zone is FFS. NOTE 3: The analysis was done only for the case of operating at max output power, in-band, non-CA. NOTE 4: The assessment assumes maximum DUT output power. NOTE 5: This contributor shall only be considered for TRP measurements. NOTE 6: This contributor shall only be considered for EIRP measurements. NOTE 7: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 8: This contributor shall only be considered for spherical EIRP measurements |  |  |  |  |  |  |  |  |  |  |  |

## B.3.2	Uncertainty budget format and assessment for IFF

The uncertainty contributions that may impact the overall MU value are listed in Table B.3.2-1.

Table B.3.2-1: Uncertainty contributions for EIRP and TRP measurement

| UID |  | Description of uncertainty contribution |  | Details in clause |  |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 |  | Positioning misalignment |  | B.2.2.1 |  |
| 2 |  | Measure distance uncertainty |  | B.2.2.2 |  |
| 3 |  | Quality of Quiet Zone |  | B.2.2.3 |  |
| 4 |  | Mismatch |  | B.2.2.4 |  |
| 5 |  | Standing wave between the DUT and measurement antenna |  | B.2.2.5 |  |
| 6 |  | Uncertainty of the RF power measurement equipment |  | B.2.2.6 |  |
| 7 |  | Phase curvature |  | B.2.2.7 |  |
| 8 |  | Amplifier uncertainties |  | B.2.2.8 |  |
| 9 |  | Random uncertainty |  | B.2.2.9 |  |
| 10 |  | Influence of the XPD |  | B.2.2.10 |  |
| 11 |  | Insertion Loss Variation |  | B.2.2.11 |  |
| 12 |  | RF leakage (from measurement antenna to the receiver/transmitter) |  | B.2.2.12 |  |
| 13 |  | Influence of TRP measurement grid |  | B.2.2.22 |  |
| 14 |  | Influence of beam peak search grid |  | B.2.2.23 |  |
| 15 |  | Multiple measurement antenna uncertainty |  | B.2.2.25 |  |
| 16 |  | DUT repositioning |  | B.2.2.26 |  |
| 17 |  | Influence of spherical coverage grid |  | B.2.2.29 |  |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 18 |  | Mismatch |  | B.2.2.4 |  |
| 19 |  | Amplifier Uncertainties |  | B.2.2.8 |  |
| 20 |  | Misalignment of positioning System |  | B.2.2.13 |  |
| 21 |  | Uncertainty of the Network Analyzer |  | B.2.2.14 |  |
| 22 |  | Uncertainty of the absolute gain of the calibration antenna |  | B.2.2.15 |  |
| 23 |  | Positioning and pointing misalignment between the reference antenna and the measurement antenna |  | B.2.2.16 |  |
| 24 |  | Phase centre offset of calibration antenna |  | B.2.2.18 |  |
| 25 |  | Quality of quiet zone for calibration process |  | B.2.2.19 |  |
| 26 |  | Standing wave between reference calibration antenna and measurement antenna |  | B.2.2.20 |  |
| 27 |  | Influence of the calibration antenna feed cable |  | B.2.2.21 |  |
| 28 |  | Insertion Loss Variation |  | B.2.1.11 |  |
| Systematic uncertainties |  |  |  |  |  |
| 29 |  | Systematic error due to TRP calculation/quadrature |  | B.2.2.24 |  |
| 30 |  | Influence of noise |  | B.2.1.27 |  |
| 31 |  | Systematic error related to beam peak search |  | B.2.2.28 |  |

The uncertainty assessment tables are organized as follows:

-	For the purpose of uncertainty assessment, the radiating antenna aperture of the DUT is denoted as D

-	The uncertainty assessment has been derived for the case of Quiet Zone size ≤ [30 cm], f = {23.45GHz, 32.125GHz, 40.8GHz, 44.3GHz}, [P = maximum output power].

-	The uncertainty assessment for EIRP and TRP in normal temperature conditions is provided in Table B.3.2-2 for PC3 UEs and in Table B.3.2-6 for PC1 and PC5 UEs.

-	The uncertainty assessment for EIRP and TRP in extreme temperature conditions is provided in Table B.3.2-8 for PC3 UEs and in Table B.3.2-9 for PC1 and PC5 UEs.

-	The uncertainty assessment for Spherical coverage is provided in Table B.3.2-4 for PC3 UEs in Table B.3.2-7 for PC1 and PC5 UEs.

Table B.3.2-2: Uncertainty assessment for EIRP and TRP measurement (f=23.45GHz, 32.125GHz, 40.8GHz, 44.3GHz, Quiet Zone size ≤ 30 cm) for PC3 UEs and normal temperature condition

| UID |  | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |
| 1 |  | Positioning misalignment | 0.00 | Normal | 2.00 | 0.00 |  |
| 2 |  | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |  |
| 3 |  | Quality of Quiet Zone (NOTE 1) (23.45GHz <= f <= 40.8GHz) | 0.6 | Actual | 1.00 | 0.6 |  |
| 3 |  | Quality of Quiet Zone (NOTE 1) (40.8GHz < f <= 44.3GHz) | 0.7 | Actual | 1.00 | 0.7 |  |
| 4 |  | Mismatch (23.45GHz <= f <= 40.8GHz) | 1.30 | Actual | 1.00 | 1.30 |  |
| 4 |  | Mismatch (40.8GHz < f <= 44.3GHz) | 1.70 | Actual | 1.00 | 1.70 |  |
| 5 |  | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |  |
| 6 |  | Uncertainty of the RF power measurement equipment (NOTE 3) (23.45GHz <= f <= 40.8GHz) | 2.16 | Normal | 2.00 | 1.08 |  |
| 6 |  | Uncertainty of the RF power measurement equipment (NOTE 3) (40.8GHz < f <= 44.3GHz) | 3.6 | Normal | 2.00 | 1.8 |  |
| 7 |  | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |  |
| 8 |  | Amplifier uncertainties | 2.10 | Normal | 2.00 | 1.05 |  |
| 9 |  | Random uncertainty | 0.50 | Normal | 2.00 | 0.25 |  |
| 10 |  | Influence of the XPD | 0.01 | U-shaped | 1.41 | 0.00 |  |
| 11 |  | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |  |
| 12 |  | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |  |
| 13 |  | Influence of TRP measurement grid (NOTE 4) | 0.25 | Actual | 1 | 0.25 |  |
| 14 |  | Influence of beam peak search grid (NOTE 5) | 0.00 | Actual | 1 | 0.00 |  |
| 15 |  | Multiple measurement antenna uncertainty (NOTE 9) | 0.15 | Actual | 1 | 0.15 |  |
| 16 |  | DUT repositioning | 0.00 (NOTE 4) 0.08 (NOTE 5) | Rectangular | 1.73 | 0.00 (NOTE 4) 0.05 (NOTE 5) |  |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |
| 17 |  | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |  |
| 18 |  | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |  |
| 19 |  | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |  |
| 20 |  | Uncertainty of the Network Analyzer (23.45GHz <= f <= 40.8GHz) | 1.50 | Normal | 2.00 | 0.75 |  |
| 20 |  | Uncertainty of the Network Analyzer (40.8GHz < f <= 44.3GHz) | 1.70 | Normal | 2.00 | 0.85 |  |
| 21 |  | Uncertainty of the absolute gain of the calibration antenna | 0.60 | Normal | 2.00 | 0.30 |  |
| 22 |  | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.01 | Rectangular | 1.73 | 0.00 |  |
| 23 |  | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |  |
| 24 |  | Quality of quiet zone for calibration process (NOTE 1) (23.45GHz <= f <= 40.8GHz) | 0.4 | Actual | 1.00 | 0.4 |  |
| 24 |  | Quality of quiet zone for calibration process (NOTE 1) (40.8GHz < f <= 44.3GHz) | 0.5 | Actual | 1.00 | 0.5 |  |
| 25 |  | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |  |
| 26 |  | Influence of the calibration antenna feed cable | 0.14 | Normal | 2.00 | 0.07 |  |
| 27 |  | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |  |
|  |  | Systematic uncertainties (NOTE 6) |  |  |  | Value |  |
| 28 |  | Systematic error due to TRP calculation/quadrature (NOTE 4) |  |  |  | 0.00 |  |
| 29 |  | Influence of noise (23.45GHz <= f <= 32.125GHz) |  |  |  | 0.1 |  |
| 29 |  | Influence of noise (32.125GHz < f <= 40.8GHz) |  |  |  | 0.3 |  |
| 29 |  | Influence of noise (40.8GHz < f <= 44.3GHz) |  |  |  | 0.3 |  |
| 30 |  | Systematic error related to beam peak search (NOTE 5) |  |  |  | 0.5 |  |
| Total measurement uncertainty |  |  |  |  |  | Value |  |
| EIRP Expanded uncertainty (23.45GHz <= f <= 32.125GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  | 5.08 |  |
| EIRP Expanded uncertainty (32.125GHz < f <= 40.8GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  | 5.28 |  |
| EIRP Expanded uncertainty (40.8GHz < f <= 44.3GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  | 6.64 |  |
| TRP Expanded uncertainty (23.45GHz <= f <= 32.125GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  | 4.61 |  |
| TRP Expanded uncertainty (32.125GHz < f <= 40.8GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  | 4.81 |  |
| TRP Expanded uncertainty (40.8GHz < f <= 44.3GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  | 6.16 |  |
| NOTE 1: The quality of quiet zone is the same for EIRP and TRP. Value based on procedure defined in clause D.2 of TR 38.810 for Quiet Zone size less or equal to 30 cm. NOTE 2: The analysis was done only for the case of operating at max output power, in-band, non-CA. NOTE 3: The assessment assumes maximum DUT output power. NOTE 4: This contributor shall only be considered for TRP measurements. NOTE 5: This contributor shall only be considered for EIRP measurements. NOTE 6: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 7: Void. NOTE 8: Void NOTE 9: Applies to the system which has a structure of mechanical feed antenna positioning. |  |  |  |  |  |  |  |

Table B.3.2-3: Void

Table B.3.2-4: Uncertainty assessment for Spherical coverage measurement (f=23.45GHz, 32.125GHz, 40.8GHz, Quiet Zone size ≤ 30 cm) for PC3 UEs and normal temperature condition

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment | 0.00 | Normal | 2.00 | 0.00 |
| 2 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 3 | Quality of Quiet Zone (NOTE 1) | 0.6 | Actual | 1.00 | 0.6 |
| 4 | Mismatch | 1.30 | Actual | 1.00 | 1.30 |
| 5 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | Uncertainty of the RF power measurement equipment (NOTE 3) | 2.16 | Normal | 2.00 | 1.08 |
| 7 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 8 | Amplifier uncertainties | 2.1 | Normal | 2.00 | 1.05 |
| 9 | Random uncertainty | 0.50 | Normal | 2.00 | 0.25 |
| 10 | Influence of the XPD | 0.01 | U-shaped | 1.41 | 0.00 |
| 11 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 13 | Multiple measurement antenna uncertainty (NOTE 5) | 0.15 | Actual | 1 | 0.15 |
| 14 | DUT repositioning | 0.00 | Rectangular | 1.73 | 0.00 |
| 15 | Influence of spherical coverage grid | 0.12 | Actual | 1 | 0.12 |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 16 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 17 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 18 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 19 | Uncertainty of the Network Analyzer | 1.50 | Normal | 2.00 | 0.75 |
| 20 | Uncertainty of the absolute gain of the calibration antenna | 0.60 | Normal | 2.00 | 0.30 |
| 21 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.01 | Rectangular | 1.73 | 0.00 |
| 22 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 23 | Quality of quiet zone for calibration process (NOTE 1) | 0.4 | Actual | 1.00 | 0.4 |
| 24 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 25 | Influence of the calibration antenna feed cable | 0.14 | Normal | 2.00 | 0.07 |
| 26 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
|  | Systematic uncertainties (NOTE 4) |  |  |  | Value |
| 27 | Influence of noise (23.45GHz <= f <= 32.125GHz) |  |  |  | 0.3 |
| 27 | Influence of noise (32.125GHz < f <= 40.8GHz) |  |  |  | 0.9 |
| Total measurement uncertainty |  |  |  |  | Value |
| Spherical coverage Expanded uncertainty (23.45GHz <= f <= 32.125GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 4.78 |
| Spherical coverage Expanded uncertainty (32.125GHz < f <= 40.8GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 5.38 |
| NOTE 1: The quality of quiet zone is the same for EIRP and TRP. Value based on procedure defined in clause D.2 of TR 38.810 for Quiet Zone size less or equal to 30 cm. NOTE 2: The analysis was done only for the case of operating at max output power, in-band, non-CA. NOTE 3: The assessment assumes maximum DUT output power. NOTE 4: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 5: Applies to the system which has a structure of mechanical feed antenna positioning. |  |  |  |  |  |

Table B.3.2-5: Void

Table B.3.2-6: Uncertainty assessment for EIRP and TRP measurement (f=23.45GHz, 32.125GHz, 40.8GHz, Quiet Zone size ≤ 30 cm) for PC1 and PC5 UEs and normal temperature condition

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment | 0.02 | Normal | 2.00 | 0.01 |
| 2 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 3 | Quality of Quiet Zone (NOTE 1) | 0.6 | Actual | 1.00 | 0.6 |
| 4 | Mismatch | 1.30 | Actual | 1.00 | 1.30 |
| 5 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | Uncertainty of the RF power measurement equipment (NOTE 3) | 2.16 | Normal | 2.00 | 1.08 |
| 7 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 8 | Amplifier uncertainties | 2.10 | Normal | 2.00 | 1.05 |
| 9 | Random uncertainty | 0.50 | Normal | 2.00 | 0.25 |
| 10 | Influence of the XPD | 0.01 | U-shaped | 1.41 | 0.00 |
| 11 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 13 | Influence of TRP measurement grid (NOTE 4) | 0.25 | Actual | 1 | 0.25 |
| 14 | Influence of beam peak search grid (NOTE 5) | 0.00 | Actual | 1 | 0.00 |
| 15 | Multiple measurement antenna uncertainty (NOTE 9) | 0.15 | Actual | 1 | 0.15 |
| 16 | DUT repositioning | 0.00 (NOTE 4) 0.35 (NOTE 5) | Rectangular | 1.73 | 0.00 (NOTE 4) 0.20 (NOTE 5) |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 17 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 18 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 19 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 20 | Uncertainty of the Network Analyzer | 1.50 | Normal | 2.00 | 0.75 |
| 21 | Uncertainty of the absolute gain of the calibration antenna | 0.60 | Normal | 2.00 | 0.30 |
| 22 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.01 | Rectangular | 1.73 | 0.00 |
| 23 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 24 | Quality of quiet zone for calibration process (NOTE 1) | 0.4 | Actual | 1.00 | 0.4 |
| 25 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 26 | Influence of the calibration antenna feed cable | 0.14 | Normal | 2.00 | 0.07 |
| 27 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
|  | Systematic uncertainties (NOTE 6) |  |  |  | Value |
| 28 | Systematic error due to TRP calculation/quadrature (NOTE 4) |  |  |  | 0.00 |
| 29 | Influence of noise (23.45GHz <= f <= 32.125GHz) |  |  |  | 0.13 |
| 29 | Influence of noise (32.125GHz < f <= 40.8GHz) |  |  |  | 0.27 (NOTE 4) 0.20 (NOTE 5) |
| 30 | Systematic error related to beam peak search (NOTE 5) |  |  |  | 0.7 |
| Total measurement uncertainty |  |  |  |  | Value |
| EIRP Expanded uncertainty (23.45GHz <= f <= 32.125GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 5.33 |
| EIRP Expanded uncertainty (32.125GHz < f <= 40.8GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 5.40 |
| TRP Expanded uncertainty (23.45GHz <= f <= 32.125GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 4.64 |
| TRP Expanded uncertainty (32.125GHz < f <= 40.8GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 4.78 |
| NOTE 1: The quality of quiet zone is the same for EIRP and TRP. Value based on procedure defined in clause D.2 of TR 38.810 for Quiet Zone size less or equal to 30 cm. NOTE 2: The analysis was done only for the case of operating at max output power, in-band, non-CA. NOTE 3: The assessment assumes maximum DUT output power. NOTE 4: This contributor shall only be considered for TRP measurements. NOTE 5: This contributor shall only be considered for EIRP measurements. NOTE 6: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 7: Void. NOTE 8: Void NOTE 9: Applies to the system which has a structure of mechanical feed antenna positioning. |  |  |  |  |  |

Table B.3.2-7: Uncertainty assessment for Spherical coverage measurement (f=23.45GHz, 32.125GHz, 40.8GHz, Quiet Zone size ≤ 30 cm) for PC1 and PC5 UEs and normal temperature condition

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment | 0.02 | Normal | 2.00 | 0.01 |
| 2 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 3 | Quality of Quiet Zone (NOTE 1) | 0.60 | Actual | 1.00 | 0.60 |
| 4 | Mismatch | 1.30 | Actual | 1.00 | 1.30 |
| 5 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | Uncertainty of the RF power measurement equipment (NOTE 3) | 2.16 | Normal | 2.00 | 1.08 |
| 7 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 8 | Amplifier uncertainties | 2.10 | Normal | 2.00 | 1.05 |
| 9 | Random uncertainty | 0.50 | Normal | 2.00 | 0.25 |
| 10 | Influence of the XPD | 0.01 | U-shaped | 1.41 | 0.00 |
| 11 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 13 | Multiple measurement antenna uncertainty (NOTE 5) | 0.15 | Actual | 1 | 0.15 |
| 14 | DUT repositioning | 0.00 | Rectangular | 1.73 | 0.00 |
| 15 | Influence of spherical coverage grid | 0.13 | Actual | 1 | 0.13 |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 16 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 17 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 18 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 19 | Uncertainty of the Network Analyzer | 1.50 | Normal | 2.00 | 0.75 |
| 20 | Uncertainty of the absolute gain of the calibration antenna | 0.60 | Normal | 2.00 | 0.30 |
| 21 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.01 | Rectangular | 1.73 | 0.00 |
| 22 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 23 | Quality of quiet zone for calibration process (NOTE 1) | 0.40 | Actual | 1.00 | 0.40 |
| 24 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 25 | Influence of the calibration antenna feed cable | 0.14 | Normal | 2.00 | 0.07 |
| 26 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
|  | Systematic uncertainties (NOTE 4) |  |  |  | Value |
| 27 | Influence of noise (23.45GHz <= f <= 32.125GHz) |  |  |  | 0.20 |
| 27 | Influence of noise (32.125GHz < f <= 40.8GHz) |  |  |  | 0.35 |
| Total measurement uncertainty |  |  |  |  | Value |
| Spherical coverage Expanded uncertainty (23.45GHz <= f <= 32.125GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 4.69 |
| Spherical coverage Expanded uncertainty (32.125GHz < f <= 40.8GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 4.84 |
| NOTE 1: The quality of quiet zone is the same for EIRP and TRP. Value based on procedure defined in clause D.2 of TR 38.810 for Quiet Zone size less or equal to 30 cm. NOTE 2: The analysis was done only for the case of operating at max output power, in-band, non-CA. NOTE 3: The assessment assumes maximum DUT output power. NOTE 4: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 5: Applies to the system which has a structure of mechanical feed antenna positioning. |  |  |  |  |  |

Table B.3.2-8: Uncertainty assessment for EIRP and TRP measurement (f=23.45GHz, 32.125GHz, 40.8GHz, 44.3GHz, Quiet Zone size ≤ 30 cm) for PC3 UEs and extreme temperature condition

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment | 0.00 | Normal | 2.00 | 0.00 |
| 2 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 3 | Quality of Quiet Zone (NOTE 1) (NOTE 8) | 0.9 | Actual | 1.00 | 0.9 |
| 4 | Mismatch (23.45GHz <= f <= 40.8GHz) | 1.30 | Actual | 1.00 | 1.30 |
| 4 | Mismatch (40.8GHz < f <= 44.3GHz) | 1.70 | Actual | 1.00 | 1.70 |
| 5 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | Uncertainty of the RF power measurement equipment (NOTE 3) (23.45GHz <= f <= 40.8GHz) | 2.16 | Normal | 2.00 | 1.08 |
| 6 | Uncertainty of the RF power measurement equipment (NOTE 3) (40.8GHz < f <= 44.3GHz) | 3.6 | Normal | 2.00 | 1.8 |
| 7 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 8 | Amplifier uncertainties | 2.10 | Normal | 2.00 | 1.05 |
| 9 | Random uncertainty | 0.50 | Normal | 2.00 | 0.25 |
| 10 | Influence of the XPD | 0.01 | U-shaped | 1.41 | 0.00 |
| 11 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 13 | Influence of TRP measurement grid (NOTE 4) | 0.25 | Actual | 1 | 0.25 |
| 14 | Influence of beam peak search grid (NOTE 5) | 0.00 | Actual | 1 | 0.00 |
| 15 | Multiple measurement antenna uncertainty (NOTE 7) | 0.15 | Actual | 1 | 0.15 |
| 16 | DUT repositioning | 0.00 (NOTE 4) 0.08 (NOTE 5) | Rectangular | 1.73 | 0.00 (NOTE 4) 0.05 (NOTE 5) |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 17 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 18 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 19 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 20 | Uncertainty of the Network Analyzer (23.45GHz <= f <= 40.8GHz) | 1.50 | Normal | 2.00 | 0.75 |
| 20 | Uncertainty of the Network Analyzer (40.8GHz < f <= 44.3GHz) | 1.70 | Normal | 2.00 | 0.85 |
| 21 | Uncertainty of the absolute gain of the calibration antenna | 0.60 | Normal | 2.00 | 0.30 |
| 22 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.01 | Rectangular | 1.73 | 0.00 |
| 23 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 24 | Quality of quiet zone for calibration process (NOTE 1) (NOTE 8) | 0.6 | Actual | 1.00 | 0.6 |
| 25 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 26 | Influence of the calibration antenna feed cable | 0.14 | Normal | 2.00 | 0.07 |
| 27 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
|  | Systematic uncertainties (NOTE 6) |  |  |  | Value |
| 28 | Systematic error due to TRP calculation/quadrature (NOTE 4) |  |  |  | 0.00 |
| 29 | Influence of noise (23.45GHz <= f <= 32.125GHz) |  |  |  | 0.1 |
| 29 | Influence of noise (32.125GHz < f <= 40.8GHz) |  |  |  | 0.3 |
| 30 | Systematic error related to beam peak search (NOTE 5) |  |  |  | 0.5 |
| Total measurement uncertainty |  |  |  |  | Value |
| EIRP Expanded uncertainty (23.45GHz <= f <= 32.125GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 5.35 |
| EIRP Expanded uncertainty (32.125GHz < f <= 40.8GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 5.55 |
| EIRP Expanded uncertainty (40.8GHz < f <= 44.3GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 6.78 |
| TRP Expanded uncertainty (23.45GHz <= f <= 32.125GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 4.87 |
| TRP Expanded uncertainty (32.125GHz < f <= 40.8GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 5.07 |
| TRP Expanded uncertainty (40.8GHz < f <= 44.3GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | TBD |
| NOTE 1: The quality of quiet zone is the same for EIRP and TRP. Value based on procedure defined in clause D.2 of TR 38.810 for Quiet Zone size less or equal to 30 cm. The ETC QoQZ MU and ETC calibration path losses shall be applied to the NTC test cases if the ETC environment is used for NTC test cases. NOTE 2: The analysis was done only for the case of operating at max output power, in-band, non-CA. NOTE 3: The assessment assumes maximum DUT output power. NOTE 4: This contributor shall only be considered for TRP measurements. NOTE 5: This contributor shall only be considered for EIRP measurements. NOTE 6: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 7: Applies to the system which has a structure of mechanical feed antenna positioning. NOTE 8: For n259 in ETC, QoQZ MU factor only defined for EIRP. |  |  |  |  |  |

Table B.3.2-9: Uncertainty assessment for EIRP and TRP measurement (f=23.45GHz, 32.125GHz, 40.8GHz, Quiet Zone size ≤ 30 cm) for PC1 and PC5 UEs and extreme temperature condition

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment | 0.02 | Normal | 2.00 | 0.01 |
| 2 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 3 | Quality of Quiet Zone (NOTE 1) | 0.90 | Actual | 1.00 | 0.90 |
| 4 | Mismatch | 1.30 | Actual | 1.00 | 1.30 |
| 5 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | Uncertainty of the RF power measurement equipment (NOTE 3) | 2.16 | Normal | 2.00 | 1.08 |
| 7 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 8 | Amplifier uncertainties | 2.10 | Normal | 2.00 | 1.05 |
| 9 | Random uncertainty | 0.50 | Normal | 2.00 | 0.25 |
| 10 | Influence of the XPD | 0.01 | U-shaped | 1.41 | 0.00 |
| 11 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 13 | Influence of TRP measurement grid (NOTE 4) | 0.25 | Actual | 1 | 0.25 |
| 14 | Influence of beam peak search grid (NOTE 5) | 0.00 | Actual | 1 | 0.00 |
| 15 | Multiple measurement antenna uncertainty (NOTE 7) | 0.15 | Actual | 1 | 0.15 |
| 16 | DUT repositioning | 0.00 (NOTE 4) 0.35 (NOTE 5) | Rectangular | 1.73 | 0.00 (NOTE 4) 0.20 (NOTE 5) |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 17 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 18 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 19 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 20 | Uncertainty of the Network Analyzer | 1.50 | Normal | 2.00 | 0.75 |
| 21 | Uncertainty of the absolute gain of the calibration antenna | 0.60 | Normal | 2.00 | 0.30 |
| 22 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.01 | Rectangular | 1.73 | 0.00 |
| 23 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 24 | Quality of quiet zone for calibration process (NOTE 1) | 0.60 | Actual | 1.00 | 0.60 |
| 25 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 26 | Influence of the calibration antenna feed cable | 0.14 | Normal | 2.00 | 0.07 |
| 27 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
|  | Systematic uncertainties (NOTE 6) |  |  |  | Value |
| 28 | Systematic error due to TRP calculation/quadrature (NOTE 4) |  |  |  | 0.00 |
| 29 | Influence of noise (23.45GHz <= f <= 32.125GHz) |  |  |  | 0.13 |
| 29 | Influence of noise (32.125GHz < f <= 40.8GHz) |  |  |  | 0.27 (NOTE 4) 0.20 (NOTE 5) |
| 30 | Systematic error related to beam peak search (NOTE 5) |  |  |  | 0.7 |
| Total measurement uncertainty |  |  |  |  | Value |
| EIRP Expanded uncertainty (23.45GHz <= f <= 32.125GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 5.60 |
| EIRP Expanded uncertainty (32.125GHz < f <= 40.8GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 5.67 |
| TRP Expanded uncertainty (23.45GHz <= f <= 32.125GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 4.90 |
| TRP Expanded uncertainty (32.125GHz < f <= 40.8GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 5.04 |
| NOTE 1: The quality of quiet zone is the same for EIRP and TRP. Value based on procedure defined in clause D.2 of TR 38.810 for Quiet Zone size less or equal to 30 cm. The ETC QoQZ MU and ETC calibration path losses shall be applied to the NTC test cases if the ETC environment is used for NTC test cases. NOTE 2: The analysis was done only for the case of operating at max output power, in-band, non-CA. NOTE 3: The assessment assumes maximum DUT output power. NOTE 4: This contributor shall only be considered for TRP measurements. NOTE 5: This contributor shall only be considered for EIRP measurements. NOTE 6: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 7: Applies to the system which has a structure of mechanical feed antenna positioning. |  |  |  |  |  |

## B.3.3	Uncertainty budget format and assessment for NFTF

The uncertainty contributions that may impact the overall MU value are listed in Table B.3.3-1.

Table B.3.3-1: Uncertainty contributions for EIRP and TRP measurement

| UID | Description of uncertainty contribution | Details in paragraph |
| --- | --- | --- |
|  | Stage 2: EIRP Near Field Radiation Pattern Measurement and EIRP Near Field DUT power measurement |  |
| 1 | Axis Alignment | # B.2.3.1 |
| 2 | Measurement Distance Uncertainty | B.2.3.2 |
| 3 | Quality of the Quiet Zone | B.2.3.3 |
| 4 | Mismatch | B.2.3.4 |
| 5 | Multiple Reflections: Coupling between Measurement Antenna and DUT | B.2.3.5 |
| 6 | Uncertainty of the RF power measurement equipment | B.2.3.6 |
| 7 | Phase curvature | B.2.3.7 |
| 8 | Amplifier uncertainties | B.2.3.8 |
| 9 | Random uncertainty | B.2.3.9 |
| 10 | Influence of the XPD | B.2.3.10 |
| 11 | NF to FF truncation | B.2.3.11 |
| 12 | Probe Polarization Amplitude and Phase | B.2.3.12 |
| 13 | Probe Array Uniformity (for multi-probe systems only) | B.2.3.13 |
| 14 | Phase Recovery Non-Linearity over signal bandwidth | B.2.3.16 |
| 15 | Probe Pattern Effect | B.2.3.17 |
| 16 | Phase Drift and Noise | B.2.3.20 |
| 17 | Leakage and Crosstalk | B.2.3.25 |
|  | Stage 1: Calibration measurement |  |
| 18 | Mismatch | B.2.3.4 |
| 19 | Amplifier uncertainties | B.2.3.8 |
| 20 | Uncertainty of the Network Analyzer | B.2.3.14 |
| 21 | Uncertainty of the absolute gain of the calibration antenna | B.2.3.15 |
| 22 | Phase centre offset of calibration | B.2.3.18 |
| 23 | Quality of the Quiet Zone for Calibration Process | B.2.3.19 |
| 24 | Mismatch in the connection of the calibration antenna | B.2.3.21 |

The uncertainty assessment table is organized as follows:

-	For the purpose of uncertainty assessment, the radiating antenna aperture of the DUT is denoted as D

-	The uncertainty assessment has been derived for the case of D = [5 cm], f = {22.65GHz, 31.1GHz, 45.1GHz}, P = [maximum output power].

-	The uncertainty assessment for EIRP and TRP is provided in Table B.3.1-2.

Table B.3.3-2: Uncertainty assessment for EIRP and TRP measurement (f=TBD, D=TBD)

| UID | Description of uncertainty contribution | Uncertainty Value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
|  | Stage 2: EIRP Near Field Radiation Pattern Measurement and EIRP Near Field DUT power measurement |  |  |  |  |
| 1 | Axis Alignment |  |  |  |  |
| 2 | Measurement Distance Uncertainty |  |  |  |  |
| 3 | Quality of the Quiet Zone |  |  |  |  |
| 4 | Mismatch |  |  |  |  |
| 5 | Multiple Reflections: Coupling between Measurement Antenna and DUT |  |  |  |  |
| 6 | Uncertainty of the RF power measurement equipment |  |  |  |  |
| 7 | Phase curvature |  |  |  |  |
| 8 | Amplifier uncertainties |  |  |  |  |
| 9 | Random uncertainty |  |  |  |  |
| 10 | Influence of the XPD |  |  |  |  |
| 11 | NF to FF truncation |  |  |  |  |
| 12 | Probe Polarization Amplitude and Phase |  |  |  |  |
| 13 | Probe Array Uniformity (for multi-probe systems only) |  |  |  |  |
| 14 | Phase Recovery Non-Linearity over signal bandwidth |  |  |  |  |
| 15 | Probe Pattern Effect |  |  |  |  |
| 16 | Phase Drift and Noise |  |  |  |  |
| 17 | Leakage and Crosstalk |  |  |  |  |
|  | Stage 1: Calibration measurement |  |  |  |  |
| 18 | Mismatch |  |  |  |  |
| 19 | Amplifier uncertainties |  |  |  |  |
| 20 | Uncertainty of the Network Analyzer |  |  |  |  |
| 21 | Uncertainty of the absolute gain of the calibration antenna |  |  |  |  |
| 22 | Phase centre offset of calibration |  |  |  |  |
| 23 | Quality of the Quiet Zone for Calibration Process |  |  |  |  |
| 24 | Mismatch in the connection of the calibration antenna |  |  |  |  |
| EIRP Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  |
| TRP Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  |
| NOTE 1: The impact of phase variation on EIRP shall be taken into account during final MU definition for the test method. NOTE 2: The quality of quiet zone is different for EIRP and TRP. For TRP, the standard uncertainty is FFS; for EIRP FFS. NOTE 3: The analysis was done only for the case of operating at max output power, in-band, non-CA NOTE 4: The assessment assumes maximum DUT output power. NOTE 5: The Phase Recovery Non-Linearity over signal bandwidth is shall be taken into account during final MU definition for the test method. |  |  |  |  |  |

## <<< Skip unchanged sections >>>

# B.15	Occupied bandwidth

Following tables summarize the MU threshold for EIRP measurements for Occupied bandwidth. The origin MU values for different test setups can be found in following subclauses.

Table B.15-1: MU threshold for beam peak measurement for Occupied bandwidth

| Power Class | Frequency | Power | BW | Threshold MU value (NOTE1) [%CBW] |
| --- | --- | --- | --- | --- |
| PC3, PC1 | 23.45 GHz ≤ f ≤ 32.125 GHz | P = Max Output Power | 50 MHz | ±0.4 |
|  |  |  | 100 MHz | ±0.4 |
|  |  |  | 200 MHz | ±1.2 |
|  |  |  | 400 MHz | ±1.2 |
|  | 32.125 GHz < f ≤ 40.8 GHz | P = Max Output Power | 50 MHz | ±0.4 |
|  |  |  | 100 MHz | ±0.4 |
|  |  |  | 200 MHz | ±1.3 |
|  |  |  | 400 MHz | ±1.3 |
| PC3 | 40.8GHz < f <= 44.3GHz | P = Max Output Power | 50 MHz | TBD |
|  |  |  | 100 MHz | TBD |
|  |  |  | 200 MHz | TBD |
|  |  |  | 400 MHz | TBD |
| NOTE 1: Total Expanded MU for IFF for Quiet Zone size ≤ 30cm. |  |  |  |  |

## <<< Skip unchanged sections >>>

# B.16	Spectrum emission mask

Following tables summarize the MU threshold for TRP measurements for Spectrum emission mask. The origin MU values for different test setups can be found in following subclauses.

Table B.16-1: MU threshold for TRP measurement for Spectrum emission mask

| Power Class | Frequency | MBW | Power | Threshold MU value (NOTE 1) |
| --- | --- | --- | --- | --- |
| PC3 | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = Max Output Power | 5.13 |
|  | 32.125GHz < f <= 40.8GHz |  |  | 5.51 |
|  | 40.8GHz < f <= 44.3GHz |  |  | TBD |
| PC1 | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = Max Output Power | 6.32 |
|  | 32.125GHz < f <= 40.8GHz |  |  | FFS |
| NOTE 1: Total Expanded MU for IFF for Quiet Zone size ≤ 30cm in Table B.16.2-2 for PC3 UEs and in Table B.16.2-4 for PC1 UEs |  |  |  |  |

## B.16.1	Uncertainty budget format and assessment for DFF

The uncertainty contributions that may impact the overall MU value are listed in Table B.16.1-1.

Table B.16.1-1: Uncertainty contributions for TRP measurement

| UID | Description of uncertainty contribution | Details in annex |
| --- | --- | --- |
| Stage 2: DUT measurement |  |  |
| 1 | Positioning misalignment | B.2.1.1 |
| 2 | Measure distance uncertainty | B.2.1.2 |
| 3 | Quality of quiet zone | B.2.1.3 |
| 4 | Mismatch | B.2.1.4 |
| 5 | Standing Wave Between the DUT and measurement antenna | B.2.1.5 |
| 6 | Uncertainty of the RF power measurement equipment | B.2.1.6 |
| 7 | Phase curvature | B.2.1.7 |
| 8 | Amplifier uncertainties | B.2.1.8 |
| 9 | Random uncertainty | B.2.1.9 |
| 10 | Influence of the XPD | B.2.1.10 |
| 11 | Insertion Loss Variation | B.2.1.11 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | B.2.1.12 |
| 13 | Influence of TRP measurement grid | B.2.1.22 |
| 14 | Influence of beam peak search grid | B.2.1.23 |
| 15 | Multiple measurement antenna uncertainty | B.2.1.25 |
| 16 | DUT repositioning | B.2.1.26 |
| Stage 1: Calibration measurement |  |  |
| 17 | Mismatch | B.2.1.4 |
| 18 | Amplifier uncertainties | B.2.1.8 |
| 19 | Misalignment of positioning System | B.2.1.13 |
| 20 | Uncertainty of the Network Analyzer | B.2.1.14 |
| 21 | Uncertainty of the absolute gain of the calibration antenna | B.2.1.15 |
| 22 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | B.2.1.16 |
| 23 | Phase centre offset of calibration antenna | B.2.1.18 |
| 24 | Quality of quiet zone for calibration process | B.2.1.19 |
| 25 | Standing wave between reference calibration antenna and measurement antenna | B.2.1.20 |
| 26 | Influence of the calibration antenna feed cable | B.2.1.21 |
| 27 | Insertion Loss Variation | B.2.1.11 |
| Systematic uncertainties |  |  |
| 28 | Systematic error due to TRP calculation/quadrature | B.2.1.24 |
| 29 | Influence of noise | B.2.1.27 |

The uncertainty assessment tables are organized as follows:

-	For the purpose of uncertainty assessment, the radiating antenna aperture of the DUT is denoted as D

-	The uncertainty assessment has been derived for the case of D = [5 cm], f = {23.45 GHz, 32.125 GHz, 40.8 GHz}, P = [Maximum output power].

-	The uncertainty assessment for TRP is provided in Table B.16.1-2.

Table B.16.1-2: Uncertainty assessment for TRP measurement (f=TBD, D=TBD)

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment |  |  |  |  |
| 2 | Measure distance uncertainty |  |  |  |  |
| 3 | Quality of quiet zone (NOTE 2) |  |  |  |  |
| 4 | Mismatch (NOTE 3) |  |  |  |  |
| 5 | Standing Wave Between the DUT and measurement antenna |  |  |  |  |
| 6 | Uncertainty of the RF power measurement equipment (NOTE 4) |  |  |  |  |
| 7 | Phase curvature |  |  |  |  |
| 8 | Amplifier uncertainties |  |  |  |  |
| 9 | Random uncertainty |  |  |  |  |
| 10 | Influence of the XPD |  |  |  |  |
| 11 | Insertion Loss Variation |  |  |  |  |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) |  |  |  |  |
| 13 | Influence of TRP measurement grid (NOTE 5) |  |  |  |  |
| 14 | Influence of beam peak search grid (NOTE 6) |  |  |  |  |
| 15 | Multiple measurement antenna uncertainty |  |  |  |  |
| 16 | DUT repositioning |  |  |  |  |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 17 | Mismatch |  |  |  |  |
| 18 | Amplifier uncertainties |  |  |  |  |
| 19 | Misalignment of positioning System |  |  |  |  |
| 20 | Uncertainty of the Network Analyzer |  |  |  |  |
| 21 | Uncertainty of the absolute gain of the calibration antenna |  |  |  |  |
| 22 | Positioning and pointing misalignment between the reference antenna and the measurement antenna |  |  |  |  |
| 23 | Phase centre offset of calibration antenna |  |  |  |  |
| 24 | Quality of quiet zone for calibration process (NOTE 2) |  |  |  |  |
| 25 | Standing wave between reference calibration antenna and measurement antenna |  |  |  |  |
| 26 | Influence of the calibration antenna feed cable |  |  |  |  |
| 27 | Insertion Loss Variation |  |  |  |  |
| TRP Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  |
| Systematic uncertainties (NOTE 7) |  |  |  |  | Value |
| 28 | Systematic error due to TRP calculation/quadrature (NOTE 5) |  |  |  |  |
| 29 | Influence of noise |  |  |  |  |
| Total measurement uncertainty |  |  |  |  |  |
| TRP total measurement uncertainty [dB] |  |  |  |  |  |
| NOTE 1: The impact of phase variation on EIRP is FFS. NOTE 2: The quality of quiet zone is different for EIRP and TRP. For TRP, the standard uncertainty is FFS; for EIRP, the standard uncertainty of quiet zone is FFS. NOTE 3: The analysis was done only for the case of operating at max output power, in-band, non-CA. NOTE 4: The assessment assumes maximum DUT output power. NOTE 5: This contributor shall only be considered for TRP measurements. NOTE 6: This contributor shall only be considered for EIRP measurements. NOTE 7: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. |  |  |  |  |  |

## B.16.2	Uncertainty budget format and assessment for IFF

The uncertainty contributions that may impact the overall MU value are listed in Table B.16.2-1.

Table B.16.2-1: Uncertainty contributions for TRP measurement

| UID | Description of uncertainty contribution | Details in clause |
| --- | --- | --- |
| Stage 2: DUT measurement |  |  |
| 1 | Positioning misalignment | # B.2.2.1 |
| 2 | Measure distance uncertainty | B.2.2.2 |
| 3 | Quality of Quiet Zone | B.2.2.3 |
| 4 | Mismatch | B.2.2.4 |
| 5 | Standing wave between the DUT and measurement antenna | B.2.2.5 |
| 6 | Uncertainty of the RF power measurement equipment | B.2.2.6 |
| 7 | Phase curvature | B.2.2.7 |
| 8 | Amplifier uncertainties | B.2.2.8 |
| 9 | Random uncertainty | B.2.2.9 |
| 10 | Influence of the XPD | B.2.2.10 |
| 11 | Insertion Loss Variation | B.2.2.11 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | B.2.2.12 |
| 13 | Influence of TRP measurement grid | B.2.2.22 |
| 14 | Influence of beam peak search grid | B.2.2.23 |
| 15 | Multiple measurement antenna uncertainty | B.2.2.25 |
| 16 | DUT repositioning | B.2.2.26 |
| Stage 1: Calibration measurement |  |  |
| 17 | Mismatch | B.2.2.4 |
| 18 | Amplifier Uncertainties | B.2.2.8 |
| 19 | Misalignment of positioning System | B.2.2.13 |
| 20 | Uncertainty of the Network Analyzer | B.2.2.14 |
| 21 | Uncertainty of the absolute gain of the calibration antenna | B.2.2.15 |
| 22 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | B.2.2.16 |
| 23 | Phase centre offset of calibration antenna | B.2.2.18 |
| 24 | Quality of quiet zone for calibration process | B.2.2.19 |
| 25 | Standing wave between reference calibration antenna and measurement antenna | B.2.2.20 |
| 26 | Influence of the calibration antenna feed cable | B.2.2.21 |
| 27 | Insertion Loss Variation | B.2.2.11 |
| Systematic uncertainties |  |  |
| 28 | Systematic error due to TRP calculation/quadrature | B.2.2.24 |
| 29 | Influence of noise | B.2.2.27 |

The uncertainty assessment tables are organized as follows:

-	For the purpose of uncertainty assessment, the radiating antenna aperture of the DUT is denoted as D

-	The uncertainty assessment has been derived for the case of Quiet Zone size ≤ [30 cm], f = {23.45GHz, 32.125GHz, 40.8GHz}, P = [Maximum output power].

-	The uncertainty assessment for TRP is provided in Table B.16.2-2 for PC3 UEs and in Table B.16.2-4 for PC1 UEs.

Table B.16.2-2: Uncertainty assessment for TRP measurement (f=23.45GHz, 32.125GHz, 40.8GHz, , 44.3GHz Quiet Zone size ≤ 30 cm) for PC3 UEs

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment | 0.00 | Normal | 2.00 | 0.00 |
| 2 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 3 | Quality of Quiet Zone (NOTE 9) (23.45GHz <= f <= 40.8GHz) | 0.6 | Actual | 1.00 | 0.6 |
| 3 | Quality of Quiet Zone (NOTE 1) (40.8GHz < f <= 44.3GHz) | 0.7 | Actual | 1.00 | 0.7 |
| 4 | Mismatch (NOTE 1) (23.45GHz <= f <= 40.8GHz) | 1.30 | Actual | 1.00 | 1.30 |
| 4 | Mismatch (40.8GHz < f <= 44.3GHz) | 1.70 | Actual | 1.00 | 1.70 |
| 5 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | Uncertainty of the RF power measurement equipment (NOTE 2) (23.45GHz <= f <= 40.8GHz) | 2.16 | Normal | 2.00 | 1.08 |
| 6 | Uncertainty of the RF power measurement equipment (NOTE 3) (40.8GHz < f <= 44.3GHz) | 3.6 | Normal | 2.00 | 1.8 |
| 7 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 8 | Amplifier uncertainties | 2.1 | Normal | 2.00 | 1.05 |
| 9 | Random uncertainty | 0.50 | Normal | 2.00 | 0.25 |
| 10 | Influence of the XPD | 0.01 | U-shaped | 1.41 | 0.00 |
| 11 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 13 | Influence of TRP measurement grid (NOTE 3) | 0.25 | Actual | 1 | 0.25 |
| 14 | Influence of beam peak search grid (NOTE 4) | 0.00 | Actual | 1 | 0.00 |
| 15 | Multiple measurement antenna uncertainty (NOTE 8) | 0.15 | Actual | 1 | 0.15 |
| 16 | DUT repositioning (NOTE 3) | 0.00 | Rectangular | 1.73 | 0.00 |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 17 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 18 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 19 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 20 | Uncertainty of the Network Analyzer (23.45GHz <= f <= 40.8GHz) | 1.50 | Normal | 2.00 | 0.75 |
| 20 | Uncertainty of the Network Analyzer (40.8GHz < f <= 44.3GHz) | 1.70 | Normal | 2.00 | 0.85 |
| 21 | Uncertainty of the absolute gain of the calibration antenna | 0.60 | Normal | 2.00 | 0.30 |
| 22 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.01 | Rectangular | 1.73 | 0.00 |
| 23 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 24 | Quality of quiet zone for calibration process (NOTE 9) | 0.4 | Actual | 1.00 | 0.4 |
| 25 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 26 | Influence of the calibration antenna feed cable | 0.14 | Normal | 2.00 | 0.07 |
| 27 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
|  | Systematic uncertainties (NOTE 5) |  |  |  | Value |
| 28 | Systematic error due to TRP calculation/quadrature (NOTE 3) |  |  |  | 0.00 |
| 29 | Influence of noise (23.45GHz <= f <= 32.125GHz) |  |  |  | 0.62 |
| 29 | Influence of noise (32.125GHz < f <= 40.8GHz) |  |  |  | 1.00 |
| 29 | Influence of noise (40.8GHz < f <= 44.3GHz) |  |  |  | TBD |
| Total measurement uncertainty |  |  |  |  | Value |
| TRP total measurement uncertainty (23.45GHz <= f <= 32.125GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 5.13 |
| TRP total measurement uncertainty (32.125GHz < f <= 40.8GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 5.51 |
| TRP Expanded uncertainty (40.8GHz < f <= 44.3GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | TBD |
| NOTE 1: The analysis was done only for the case of operating at max output power, in-band, non-CA. NOTE 2: The assessment assumes maximum DUT output power. NOTE 3: This contributor shall only be considered for TRP measurements. NOTE 4: This contributor shall only be considered for EIRP measurements. NOTE 5: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 6: Values extracted from TR 38.810 v2.6.1 in square brackets pending for further analysis. NOTE 7: Void. NOTE 8: Applies to the system which has a structure of mechanical feed antenna positioning. NOTE 9: Value based on procedure defined in clause D.2 of TR 38.810 for Quiet Zone size less or equal to 30 cm. |  |  |  |  |  |

Table B.16.2-3: Void

Table B.16.2-4: Uncertainty assessment for TRP measurement (f=23.45GHz, 32.125GHz, 40.8GHz, Quiet Zone size ≤ 30 cm) for PC1 UEs

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment | 0.02 | Normal | 2.00 | 0.02 |
| 2 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 3 | Quality of Quiet Zone (NOTE 9) | 0.60 | Actual | 1.00 | 0.60 |
| 4 | Mismatch (NOTE 1) | 1.30 | Actual | 1.00 | 1.30 |
| 5 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | Uncertainty of the RF power measurement equipment (NOTE 2) | 2.16 | Normal | 2.00 | 1.08 |
| 7 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 8 | Amplifier uncertainties | 2.10 | Normal | 2.00 | 1.05 |
| 9 | Random uncertainty | 0.50 | Normal | 2.00 | 0.25 |
| 10 | Influence of the XPD | 0.01 | U-shaped | 1.41 | 0.00 |
| 11 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 13 | Influence of TRP measurement grid (NOTE 3) | 0.25 | Actual | 1 | 0.25 |
| 14 | Influence of beam peak search grid (NOTE 4) | 0.00 | Actual | 1 | 0.00 |
| 15 | Multiple measurement antenna uncertainty (NOTE 8) | 0.15 | Actual | 1 | 0.15 |
| 16 | DUT repositioning (NOTE 3) | 0.00 | Rectangular | 1.73 | 0.00 |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 17 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 18 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 19 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 20 | Uncertainty of the Network Analyzer | 1.50 | Normal | 2.00 | 0.75 |
| 21 | Uncertainty of the absolute gain of the calibration antenna | 0.60 | Normal | 2.00 | 0.30 |
| 22 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.01 | Rectangular | 1.73 | 0.00 |
| 23 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 24 | Quality of quiet zone for calibration process (NOTE 9) | 0.40 | Actual | 1.00 | 0.40 |
| 25 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 26 | Influence of the calibration antenna feed cable | 0.14 | Normal | 2.00 | 0.07 |
| 27 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
|  | Systematic uncertainties (NOTE 5) |  |  |  | Value |
| 28 | Systematic error due to TRP calculation/quadrature (NOTE 3) |  |  |  | 0.00 |
| 29 | Influence of noise (23.45GHz <= f <= 32.125GHz) |  |  |  | 1.81 |
| 29 | Influence of noise (32.125GHz < f <= 40.8GHz) |  |  |  | FFS |
| Total measurement uncertainty |  |  |  |  | Value |
| TRP total measurement uncertainty (23.45GHz <= f <= 32.125GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 6.32 |
| TRP total measurement uncertainty (32.125GHz < f <= 40.8GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | FFS |
| NOTE 1: The analysis was done only for the case of operating at max output power, in-band, non-CA. NOTE 2: The assessment assumes maximum DUT output power. NOTE 3: This contributor shall only be considered for TRP measurements. NOTE 4: This contributor shall only be considered for EIRP measurements. NOTE 5: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 6: Values extracted from TR 38.810 v2.6.1 in square brackets pending for further analysis. NOTE 7: Void. NOTE 8: Applies to the system which has a structure of mechanical feed antenna positioning. NOTE 9: Value based on procedure defined in clause D.2 of TR 38.810 for Quiet Zone size less or equal to 30 cm. |  |  |  |  |  |

## <<< Skip unchanged sections >>>

# B.19	Reference Sensitivity

Following tables summarize the MU threshold for EIS measurements for Reference Sensitivity. The origin MU values for different test setups with varies parameters can be found in following subclauses.

Table B.19-1: MU threshold for EIS for Reference Sensitivity

| Power Class | Frequency | MBW | Power | Threshold MU value for NTC (NOTE 1) | Threshold MU value for ETC (NOTE 1) |
| --- | --- | --- | --- | --- | --- |
| PC3 | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = Max Output Power | 5.36 | 5.61 |
|  | 32.125GHz < f <= 40.8GHz |  |  | 5.36 | 5.61 |
|  | 40.8GHz < f <= 44.3GHz |  |  | 6.34 | 6.48 |
| PC1 | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = Max Output Power | 5.58 | 5.83 |
|  | 32.125GHz < f <= 40.8GHz |  |  | 5.58 | 5.83 |
| PC5 | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = Max Output Power | 5.58 | 5.83 |
| NOTE 1: Total Expanded MU for IFF for Quiet Zone size ≤ 30cm in Table B.19.2-2 for PC3 UEs (NTC), in Table B.19.2-4 for PC3 UEs (ETC), and Table B.19.2-3 for PC1 UEs |  |  |  |  |  |

Table B.19-2: MU threshold for Spherical coverage measurement for Reference Sensitivity

| Power Class | Frequency | MBW | Power | Threshold MU value (NOTE 1) |
| --- | --- | --- | --- | --- |
| PC3 | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = Max Output Power | 5.07 |
|  | 32.125GHz < f <= 40.8GHz |  |  | 5.07 |
|  | 40.8GHz < f <= 44.3GHz |  |  | 6.04 |
| PC1 | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = Max Output Power | 5.07 |
|  | 32.125GHz < f <= 40.8GHz |  |  | 5.07 |
| PC5 | 23.45GHz <= f <= 32.125GHz | BW <= 400MHz | P = Max Output Power | 5.07 |
| NOTE 1: Total Expanded MU for IFF for Quiet Zone size ≤ 30cm in Table B.19.2-2 for PC3 UEs and Table B.19.2-3 for PC1 UEs |  |  |  |  |

## B.19.1	Uncertainty budget format and assessment for DFF

FFS

## B.19.2	Uncertainty budget format and assessment for IFF

The uncertainty contributions that may impact the overall MU value are listed in Table B.19.2-1.

Table B.19.2-1: Uncertainty contributions for EIS measurement

| UID |  | Description of uncertainty contribution |  | Details in clause |  |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 |  | Positioning misalignment |  | B.2.2.1 |  |
| 2 |  | Measure distance uncertainty |  | B.2.2.2 |  |
| 3 |  | Quality of Quiet Zone |  | B.2.2.3 |  |
| 4 |  | Mismatch |  | B.2.2.4 |  |
| 5 |  | Standing wave between the DUT and measurement antenna |  | B.2.2.5 |  |
| 6 |  | gNB emulator uncertainty |  | B.2.2.17 |  |
| 7 |  | Phase curvature |  | B.2.2.7 |  |
| 8 |  | Amplifier uncertainties |  | B.2.2.8 |  |
| 9 |  | Random uncertainty |  | B.2.2.9 |  |
| 10 |  | Influence of the XPD |  | B.2.2.10 |  |
| 11 |  | Insertion Loss Variation |  | B.2.2.11 |  |
| 12 |  | RF leakage (from measurement antenna to the receiver/transmitter) |  | B.2.2.12 |  |
| 13 |  | Multiple measurement antenna uncertainty |  | B.2.2.25 |  |
| 14 |  | DUT repositioning |  | B.2.2.26 |  |
| 15 |  | Influence of spherical coverage grid |  | B.2.2.29 |  |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 16 |  | Mismatch |  | B.2.2.4 |  |
| 17 |  | Amplifier Uncertainties |  | B.2.2.8 |  |
| 18 |  | Misalignment of positioning System |  | B.2.2.13 |  |
| 19 |  | Uncertainty of the Network Analyzer |  | B.2.2.14 |  |
| 20 |  | Uncertainty of the absolute gain of the calibration antenna |  | B.2.2.15 |  |
| 21 |  | Positioning and pointing misalignment between the reference antenna and the measurement antenna |  | B.2.2.16 |  |
| 22 |  | Phase centre offset of calibration antenna |  | B.2.2.18 |  |
| 23 |  | Quality of quiet zone for calibration process |  | B.2.2.19 |  |
| 24 |  | Standing wave between reference calibration antenna and measurement antenna |  | B.2.2.20 |  |
| 25 |  | Influence of the calibration antenna feed cable |  | B.2.2.21 |  |
| 26 |  | Insertion Loss Variation |  | B.2.2.11 |  |
| Systematic uncertainties |  |  |  |  |  |
| 27 |  | Systematic error related to beam peak search |  | B.2.2.28 |  |
| 28 |  | Systematic error related to EIS spherical coverage |  | B.2.2.30 |  |

The uncertainty assessment tables are organized as follows:

-	For the purpose of uncertainty assessment, the radiating antenna aperture of the DUT is denoted as D

-	The uncertainty assessment has been derived for the case of Quiet Zone size ≤ [30 cm], f = {23.45GHz, 32.125GHz, 40.8GHz, 44.3GHz}, [P = maximum output power].

-	The uncertainty assessment for EIS is provided in Table B.19.2-2 for PC3 UEs and Table B.19.2-3 for PC1 UEs.

Table B.19.2-2: Uncertainty assessment for EIS measurement (f=23.45GHz, 32.125GHz, 40.8GHz, 44.3GHz, Quiet Zone size ≤ 30 cm) for PC3 UEs and normal temperature condition

| UID |  | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |  |  |
| 1 |  | Positioning misalignment | 0.00 | Normal | 2.00 | 0.00 |  |
| 2 |  | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |  |
| 3 |  | Quality of Quiet Zone (NOTE 7) (23.45GHz <= f <= 40.8GHz) | 0.6 | Actual | 1.00 | 0.6 |  |
| 3 |  | Quality of Quiet Zone (NOTE 7) (40.8GHz < f <= 44.3GHz) | 0.7 | Actual | 1.00 | 0.7 |  |
| 4 |  | Mismatch (23.45GHz <= f <= 40.8GHz) | 1.30 | Actual | 1.00 | 1.30 |  |
| 4 |  | Mismatch (40.8GHz < f <= 44.3GHz) | 1.70 | Actual | 1.00 | 1.70 |  |
| 5 |  | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |  |
| 6 |  | gNB uncertainty on absolute level (23.45GHz <= f <= 40.8GHz) | 2.9 | Normal | 2.00 | 1.45 |  |
| 6 |  | gNB uncertainty on absolute level (40.8GHz < f <= 44.3GHz) | 3.6 | Normal | 2.00 | 1.8 |  |
| 7 |  | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |  |
| 8 |  | Amplifier uncertainties | 2.1 | Normal | 2.00 | 1.05 |  |
| 9 |  | Random uncertainty | 0.50 | Normal | 2.00 | 0.25 |  |
| 10 |  | Influence of the XPD | 0.01 | U-shaped | 1.41 | 0.00 |  |
| 11 |  | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |  |
| 12 |  | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |  |
| 13 |  | Multiple measurement antenna uncertainty (NOTE 6) | 0.15 | Actual | 1.00 | 0.15 |  |
| 14 |  | DUT repositioning | 0.00 (NOTE 4) 0.08 (NOTE 5) | Rectangular | 1.73 | 0.00 (NOTE 4) 0.05 (NOTE 5) |  |
| 15 |  | Influence of spherical coverage grid (NOTE 4) | 0.12 | Actual | 1 | 0.12 |  |
| Stage 1: Calibration measurement |  |  |  |  |  |  |  |
| 16 |  | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |  |
| 17 |  | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |  |
| 18 |  | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |  |
| 19 |  | Uncertainty of the Network Analyzer (23.45GHz <= f <= 40.8GHz) | 1.50 | Normal | 2.00 | 0.75 |  |
| 19 |  | Uncertainty of the Network Analyzer (40.8GHz < f <= 44.3GHz) | 1.70 | Normal | 2.00 | 0.85 |  |
| 20 |  | Uncertainty of the absolute gain of the calibration antenna | 0.60 | Normal | 2.00 | 0.30 |  |
| 21 |  | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.01 | Rectangular | 1.73 | 0.00 |  |
| 22 |  | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |  |
| 23 |  | Quality of quiet zone for calibration process (NOTE 7) (23.45GHz <= f <= 40.8GHz) | 0.4 | Actual | 1.00 | 0.4 |  |
| 23 |  | Quality of quiet zone for calibration process (NOTE 7) (40.8GHz < f <= 44.3GHz) | 0.5 | Actual | 1.00 | 0.5 |  |
| 24 |  | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |  |
| 25 |  | Influence of the calibration antenna feed cable | 0.14 | Normal | 2.00 | 0.07 |  |
| 26 |  | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |  |
|  |  | Systematic uncertainties (NOTE 3) |  |  |  | Value |  |
| 27 |  | Systematic error related to beam peak search (NOTE 5) |  |  |  | 0.5 |  |
| 28 |  | Systematic error related to EIS spherical coverage (NOTE 4) |  |  |  | DL power step size, 0.2 |  |
| Total measurement uncertainty |  |  |  |  |  | Value |  |
| EIS Expanded uncertainty (23.45GHz <= f <= 40.8GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  | 5.36 |  |
| EIS Expanded uncertainty (40.8GHz < f <= 44.3GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  | 6.34 |  |
| EIS Spherical coverage Expanded uncertainty (23.45GHz <= f <= 40.8GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  | 5.07 |  |
| EIS Spherical coverage Expanded uncertainty (40.8GHz < f <= 44.3GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  |  | 6.04 |  |
| NOTE 1: The analysis was done only for the case of operating at max output power, in-band, non-CA. NOTE 2: Void. NOTE 3: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 4: This contributor shall only be considered for spherical EIS measurements. NOTE 5: This contributor shall only be considered for EIS measurements. NOTE 6: Applies to the system which has a structure of mechanical feed antenna positioning. NOTE 7: Value based on procedure defined in clause D.2 of TR 38.810 for Quiet Zone size less or equal to 30 cm. |  |  |  |  |  |  |  |

Table B.19.2-3: Uncertainty assessment for EIS measurement (f=23.45GHz, 32.125GHz, 40.8GHz, Quiet Zone size ≤ 30 cm) for PC1 and PC5 UEs and normal temperature condition

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment | 0.02 | Normal | 2.00 | 0.01 |
| 2 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 3 | Quality of Quiet Zone (NOTE 7) | 0.6 | Actual | 1.00 | 0.6 |
| 4 | Mismatch | 1.30 | Actual | 1.00 | 1.30 |
| 5 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | gNB uncertainty on absolute level | 2.9 | Normal | 2.00 | 1.45 |
| 7 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 8 | Amplifier uncertainties | 2.1 | Normal | 2.00 | 1.05 |
| 9 | Random uncertainty | 0.50 | Normal | 2.00 | 0.25 |
| 10 | Influence of the XPD | 0.01 | U-shaped | 1.41 | 0.00 |
| 11 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 13 | Multiple measurement antenna uncertainty (NOTE 6) | 0.15 | Actual | 1.00 | 0.15 |
| 14 | DUT repositioning | 0.00 (NOTE 4) 0.35 (NOTE 5) | Rectangular | 1.73 | 0.00 (NOTE 4) 0.20 (NOTE 5) |
| 15 | Influence of spherical coverage grid (NOTE 4) | 0.13 | Actual | 1 | 0.13 |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 16 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 17 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 18 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 19 | Uncertainty of the Network Analyzer | 1.50 | Normal | 2.00 | 0.75 |
| 20 | Uncertainty of the absolute gain of the calibration antenna | 0.60 | Normal | 2.00 | 0.30 |
| 21 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.01 | Rectangular | 1.73 | 0.00 |
| 22 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 23 | Quality of quiet zone for calibration process (NOTE 7) | 0.4 | Actual | 1.00 | 0.4 |
| 24 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 25 | Influence of the calibration antenna feed cable | 0.14 | Normal | 2.00 | 0.07 |
| 26 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
|  | Systematic uncertainties (NOTE 3) |  |  |  | Value |
| 27 | Systematic error related to beam peak search (NOTE 5) |  |  |  | 0.7 |
| 28 | Systematic error related to EIS spherical coverage (NOTE 4) |  |  |  | DL power step size, 0.2 |
| Total measurement uncertainty |  |  |  |  | Value |
| EIS Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 5.58 |
| EIS Spherical coverage Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 5.07 |
| NOTE 1: The analysis was done only for the case of operating at max output power, in-band, non-CA. NOTE 2: Void. NOTE 3: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 4: This contributor shall only be considered for spherical EIS measurements. NOTE 5: This contributor shall only be considered for EIS measurements. NOTE 6: Applies to the system which has a structure of mechanical feed antenna positioning. NOTE 7: Value based on procedure defined in clause D.2 of TR 38.810 for Quiet Zone size less or equal to 30 cm. |  |  |  |  |  |

Table B.19.2-4: Uncertainty assessment for EIS measurement (f=23.45GHz, 32.125GHz, 40.8GHz, 44.3GHz, Quiet Zone size ≤ 30 cm) for PC3 UEs and extreme temperature condition

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment | 0.00 | Normal | 2.00 | 0.00 |
| 2 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 3 | Quality of Quiet Zone (NOTE 7) | 0.9 | Actual | 1.00 | 0.9 |
| 4 | Mismatch (23.45GHz <= f <= 40.8GHz) | 1.30 | Actual | 1.00 | 1.30 |
| 4 | Mismatch (40.8GHz < f <= 44.3GHz) | 1.70 | Actual | 1.00 | 1.70 |
| 5 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | gNB uncertainty on absolute level (23.45GHz <= f <= 40.8GHz) | 2.9 | Normal | 2.00 | 1.45 |
| 6 | gNB uncertainty on absolute level (40.8GHz < f <= 44.3GHz) | 3.6 | Normal | 2.00 | 1.8 |
| 7 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 8 | Amplifier uncertainties | 2.1 | Normal | 2.00 | 1.05 |
| 9 | Random uncertainty | 0.50 | Normal | 2.00 | 0.25 |
| 10 | Influence of the XPD | 0.01 | U-shaped | 1.41 | 0.00 |
| 11 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 13 | Multiple measurement antenna uncertainty (NOTE 6) | 0.15 | Actual | 1.00 | 0.15 |
| 14 | DUT repositioning | 0.00 (NOTE 4) 0.08 (NOTE 5) | Rectangular | 1.73 | 0.00 (NOTE 4) 0.05 (NOTE 5) |
| 15 | Influence of spherical coverage grid (NOTE 4) | 0.12 | Actual | 1 | 0.12 |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 16 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 17 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 18 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 19 | Uncertainty of the Network Analyzer (23.45GHz <= f <= 40.8GHz) | 1.50 | Normal | 2.00 | 0.75 |
| 19 | Uncertainty of the Network Analyzer (40.8GHz < f <= 44.3GHz) | 1.70 | Normal | 2.00 | 0.85 |
| 20 | Uncertainty of the absolute gain of the calibration antenna | 0.60 | Normal | 2.00 | 0.30 |
| 21 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.01 | Rectangular | 1.73 | 0.00 |
| 22 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 23 | Quality of quiet zone for calibration process (NOTE 7) | 0.6 | Actual | 1.00 | 0.6 |
| 24 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 25 | Influence of the calibration antenna feed cable | 0.14 | Normal | 2.00 | 0.07 |
| 26 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
|  | Systematic uncertainties (NOTE 3) |  |  |  | Value |
| 27 | Systematic error related to beam peak search (NOTE 5) |  |  |  | 0.5 |
| 28 | Systematic error related to EIS spherical coverage (NOTE 4) |  |  |  | DL power step size, 0.2 |
| Total measurement uncertainty |  |  |  |  | Value |
| EIS Expanded uncertainty (23.45GHz <= f <= 40.8GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 5.61 |
| EIS Expanded uncertainty (40.8GHz < f <= 44.3GHz) (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 6.48 |
| NOTE 1: The analysis was done only for the case of operating at max output power, in-band, non-CA. NOTE 2: Void. NOTE 3: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 4: This contributor shall only be considered for spherical EIS measurements. NOTE 5: This contributor shall only be considered for EIS measurements. NOTE 6: Applies to the system which has a structure of mechanical feed antenna positioning. NOTE 7: Value based on procedure defined in clause D.2 of TR 38.810 for Quiet Zone size less or equal to 30 cm. The ETC QoQZ MU and ETC calibration path losses shall be applied to the NTC test cases if the ETC environment is used for NTC test cases. |  |  |  |  |  |

Table B.19.2-5: Uncertainty assessment for EIS measurement (f=23.45GHz, 32.125GHz, 40.8GHz, Quiet Zone size ≤ 30 cm) for PC1 and PC5 UEs and extreme temperature condition

| UID | Uncertainty source | Uncertainty value | Distribution of the probability | Divisor | Standard uncertainty (σ) [dB] |
| --- | --- | --- | --- | --- | --- |
| Stage 2: DUT measurement |  |  |  |  |  |
| 1 | Positioning misalignment | 0.02 | Normal | 2.00 | 0.01 |
| 2 | Measure distance uncertainty | 0.00 | Rectangular | 1.73 | 0.00 |
| 3 | Quality of Quiet Zone (NOTE 7) | 0.90 | Actual | 1.00 | 0.90 |
| 4 | Mismatch | 1.30 | Actual | 1.00 | 1.30 |
| 5 | Standing wave between the DUT and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 6 | gNB uncertainty on absolute level | 2.90 | Normal | 2.00 | 1.45 |
| 7 | Phase curvature | 0.00 | U-shaped | 1.41 | 0.00 |
| 8 | Amplifier uncertainties | 2.10 | Normal | 2.00 | 1.05 |
| 9 | Random uncertainty | 0.50 | Normal | 2.00 | 0.25 |
| 10 | Influence of the XPD | 0.01 | U-shaped | 1.41 | 0.00 |
| 11 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
| 12 | RF leakage (from measurement antenna to the receiver/transmitter) | 0.00 | Actual | 1.00 | 0.00 |
| 13 | Multiple measurement antenna uncertainty (NOTE 6) | 0.15 | Actual | 1.00 | 0.15 |
| 14 | DUT repositioning | 0.00 (NOTE 4) 0.35(NOTE 5) | Rectangular | 1.73 | 0.00 (NOTE 4) 0.20 (NOTE 5) |
| 15 | Influence of spherical coverage grid (NOTE 4) | 0.13 | Actual | 1 | 0.13 |
| Stage 1: Calibration measurement |  |  |  |  |  |
| 16 | Mismatch | 0.00 | U-shaped | 1.41 | 0.00 |
| 17 | Amplifier Uncertainties | 0.00 | Normal | 2.00 | 0.00 |
| 18 | Misalignment of positioning System | 0.00 | Normal | 2.00 | 0.00 |
| 19 | Uncertainty of the Network Analyzer | 1.50 | Normal | 2.00 | 0.75 |
| 20 | Uncertainty of the absolute gain of the calibration antenna | 0.60 | Normal | 2.00 | 0.30 |
| 21 | Positioning and pointing misalignment between the reference antenna and the measurement antenna | 0.01 | Rectangular | 1.73 | 0.00 |
| 22 | Phase centre offset of calibration antenna | 0.00 | Rectangular | 1.73 | 0.00 |
| 23 | Quality of quiet zone for calibration process (NOTE 7) | 0.60 | Actual | 1.00 | 0.60 |
| 24 | Standing wave between reference calibration antenna and measurement antenna | 0.00 | U-shaped | 1.41 | 0.00 |
| 25 | Influence of the calibration antenna feed cable | 0.14 | Normal | 2.00 | 0.07 |
| 26 | Insertion Loss Variation | 0.00 | Rectangular | 1.73 | 0.00 |
|  | Systematic uncertainties (NOTE 3) |  |  |  | Value |
| 27 | Systematic error related to beam peak search (NOTE 5) |  |  |  | 0.7 |
| 28 | Systematic error related to EIS spherical coverage (NOTE 4) |  |  |  | DL power step size, 0.2 |
| Total measurement uncertainty |  |  |  |  | Value |
| EIS Expanded uncertainty (1.96σ - confidence interval of 95 %) [dB] |  |  |  |  | 5.83 |
| NOTE 1: The analysis was done only for the case of operating at max output power, in-band, non-CA. NOTE 2: Void. NOTE 3: In order to obtain the total measurement uncertainty, systematic uncertainties have to be added to the expanded root sum square of the standard deviations of the Stage 1 and Stage 2 contributors. NOTE 4: This contributor shall only be considered for spherical EIS measurements. NOTE 5: This contributor shall only be considered for EIS measurements. NOTE 6: Applies to the system which has a structure of mechanical feed antenna positioning. NOTE 7: Value based on procedure defined in clause D.2 of TR 38.810 for Quiet Zone size less or equal to 30 cm. The ETC QoQZ MU and ETC calibration path losses shall be applied to the NTC test cases if the ETC environment is used for NTC test cases. |  |  |  |  |  |

## <<< END OF CHANGES >>>